## CH9
TID (unique) | KEYWORD(unique) | TAG(Separate with comma) | CATEGORY | PID(Separate with comma if more than one) | QUESTION | ANSWER | Documentation link | YouTube solution link | RSID
GVR81 | Before Power On | power supply | power | VENR860 | procedures to be done before powering on the system ? | 1. Connect the Inspiratory Safety Guard (ISG) to the inspiratory port. 2. Connect the Exhalation Valve Assembly (EVA) to the exhalation port. 3. Familiarize yourself with the hard keys on the display. a. 2-minute audio pause b. Increase O2 c. Snapshot d. Lock/Unlock e. Home f. Trim Knob 4. Attach the breathing circuit and all accessories that will be used to ventilate the patient. | https://drive.google.com/file/d/1Gub9cUvd1MbbtEV6c5c1B4PTVqf510ed/view?usp=sharing
GVR82 | Power on | Turn on | power | VENR860 | How to turn on the system ? | 1. Connect the air and oxygen supply hoses to the gas supply source. 2. Plug power cord into an approved electrical outlet. 3. Press the power switch to turn on the ventilator. | https://drive.google.com/file/d/1vhVgeBiqC-KXBhWBB-tWE4tHDtspOb1r/view?usp=sharing
GVR83 | Patient setup | patient care | interface | VENR860 | How to setup the system for new patient ? | 1. Select NEW PATIENT. 2. Select Adult, Pediatric or Neonatal patient type. 3. Select Patient ID (Identification), if necessary, and enter 10 characters and select Confirm. 4. Select Gender (male or female). 5. Select Height. 6. Select Weight. 7. Select Tube Type (Endotrach, Trach, or ---. 8. Select Tube Diameter. 9. Verify and confirm settings by pressing Confirm. | https://drive.google.com/file/d/1aCSe_Rg8cx_0lJegShdeOev22xT3YM5z/view?usp=sharing
GVR84 | System Check | system check,system troubleshoot,circuit check up | Interface | VENR860 | How to perform system check ?
GVR85 | RUNNING_A_SYSTEM_CHECK | Run, troubleshooting | Interface | VENR860 | How to check the system in running | 1. From Standy, select system check. 2. Attach the breathing circuit and all accessories that will be used to ventilate the patient. 3. Occlude the patient wye using the occlusion port. 4. Select start 5. Select the information icon to see the system check details menu | https://drive.google.com/file/d/10lDZamRT-fhPTgs4zdNGXqwMxfH_ChTu/view?usp=drivesdk
GVR86 | SETTING_A_VENTILATION_MODE | setup, setting, ventilation | Interface | VENR860 | How to setting the ventilation mode | 1. Select current mode. 2. Select the desired ventilation mode. 3. Select Assist Control, Leak Comp or Trigger Comp, if desired. 4. Set the desired settings for the Ventilation mode and Confirm. | https://drive.google.com/file/d/10n2XhqFRIIcx9Cf18ncoRuz1zEP59ZKr/view?usp=drivesdk
GVR87 | PATIENT_VENTILATION | Ventilation, start | Interface | VENR860 | How to start the patient ventilation | 1. From Standby, select Start Ventilation. 2. After ventilation has started, connect the breathing circuit to the patient | https://drive.google.com/file/d/10p0H_eUxR5Gu36jUv9hHHfNRcUyY9X-m/view?usp=drivesdk
GVR88 | ALARM_PANEL | Alarm, indicator, panel | Indicator | VENR860 | How to use the Alarm panel | 1. Audio pause. 2. Alarm Status. 3. Alarm setup. 4. Audio Pause Timer. 5. Active Alarm count | https://drive.google.com/file/d/10r-ICQLMuIZTAkhXB9vk7nQ3H25qTCGh/view?usp=drivesdk
GVR89 | ALARM_SETUP | Setup, Alarm, settings | Indicator | VENR860 | How to set the Alarm | 1. Apnea Time: Set apnea time up to 60 seconds. 2. Alarm Light Brightness: Set the brightness level of the alarm light. The range is 1 (Low) to 5 (High). 3. Alarm Volume: Set the alarm volume of the alarm tone. The range is 1 (Low) to 5 (High). 4. High Alert Audio: Set the delay time in which a high-priority alarm must be resolved before the audio tone pitch and volume increases to the maximum level. The range is 0 to 30 seconds or Off. A general message shows when High Alert Audio is set to Off. 5. Alarm Limits: Select the check box to show alarm limits adjacent to the measured data in the patient data workspace. The alarm limit always shows when an alarm occurs for the measured data, even if it is set to Off. 6. Auto Limits: Select to set auto alarm limits based on current measured data for: a. MVexp-Low and High b. VTexp-Low and High c. Rate-Low and High d. EtCO2-Low and High e. PEEPe-Low and High | https://drive.google.com/file/d/10um-GV2sIbFX0vncHXUwqLv-E65nl8JS/view?usp=drivesdk
GLOGIQ91 | POWERING_THE_SYSTEM | power supply,power,system | power | LOGIQ9 | Procedures to be done before powering on the system ? | 1. Ensure that the wall outlet is of the appropriate type. 2. Ensure that the power switch is turned off. 3. Unwrap the power cable. Make sure to allow sufficient slack in the cable so that the plug is not pulled out of the wall if the system is moved slightly. 4. Attach the power plug to the system and secure it in place by using the retaining clamp
GLOGIQ92 | PASSWORD_PROTECTION | Password,Protection | Power | LOGIQ9 | How to set Password to the system? | 1. Operator: Select the Operator. 2. Password: Enter Operator’s password (optional). 3. Select type of Logon or Cancel. • Emergency: Data stored only for the duration of the current examination. • OK: Standard logon • Cancel: Cancel logon If the correct ID and Password are entered, the system displays the Patient screen. Logoff To logoff, press the Power On/Off switch momentarily and a SYSTEM-EXIT window appears. | https://drive.google.com/file/d/1CQUDHicV0E1Mq1qrgxu8ViTqoDxibjWA/view?usp=sharing
GLOGIQ93 | STANDBY_MODE | Power,Time | Power | LOGIQ9 | How to Standby Mode? | 1. Press the Power On/Off switch at the front of the system once. 2. The System-Exit window displays. 3. Using the Trackball, select Standby. It takes approximately 20 seconds to go into Standby Mode. 4. Transport the system as needed and plug it back in. See ‘Moving the System’ for more information. 5. Press and immediately release the Power On/Off switch to come out of Standby Mode. This takes about 45 seconds.
GLOGIQ94 | POWER_OFF | Power | Power | LOGIQ9 | How to turn off the System | 1. Press the Power On/Off switch at the front of the system once. 2. The System-Exit window is displayed. 3. Using the Trackball, select Shutdown. The shutdown process takes a few seconds and is completed when the Power On/Off switch illumination is turned amber. 4. Disconnect the probes.Clean or disinfect all probes as necessary. Store them in their shipping cases to avoid damage. | https://drive.google.com/file/d/1FsMYoALX9AQ_moC52ROvCTR71Xc5aaJl/view?usp=sharing
GLOGIQ95 | CIRCUIT_BREAKER | Power | Power | LOGIQ9 | How to manage Power overloads? | If a power overload occurs: 1. Turn off all peripheral devices. 2. Reactivate the Circuit Breaker switch. The Circuit Breaker switch should stay in the On position; DO NOT hold the switch in the On position. If the Circuit Breaker switch remains On, follow the Power On procedure previously . NOTE: If the Circuit Breaker switch does not remain in the On position or trips again: 1. Disconnect the Power Cable. 2. Call Service immediately. DO NOT attempt to use the system. | https://drive.google.com/file/d/1TEpDX-XQzMa53QQGpI08p3xUw89fMMkW/view?usp=sharing
GLOGIQ96 | TRANSPORTING_THE_SYSTEM | Transpoting,Positioning | Transport | LOGIQ9 | Precautions taken before Transporting or Moving the System? | 1. Press the Power On/Off switch to power off. See ‘Power Off’ for more information. 2. Unplug the power cord. 3. All cables from off-board peripheral devices (IIE camera, external printer, VCR, etc.) must be disconnected from the console. 4. Disconnect the footswitch from the console. 5. Ensure that no loose items are left on the console. 6. Wind the power cable around the cable hook below the rear handle or wrap the cord in a bundle and store it behind the rear storage panel. 7. Connect all probes to be used while off site. Ensure that probe cables are out of the way from the wheels and not protruding beyond the console. NOTE: If more than four (4) probes are intended to be used, store the additional probes securely. 8. Store all other probes in their original cases or in soft cloth or foam to prevent damage. 9. Store sufficient gel and other essential accessories in the provided space. 10. Adjust the monitor and control panel to its lowest position possible by using the release lever located underneath the control panel. Make sure the keyboard is locked in place. 11. Unlock the wheels
GLOGIQ97 | PROBE_CONNECTION | Probe,Connection | Interface | LOGIQ9 | How to Connect the Probe safely? | 1. Place the probe's carrying case on a stable surface and open the case. 2. Carefully remove the probe and unwrap the probe cord. 3. DO NOT allow the probe head to hang free. Impact to the probe head could result in irreparable damage. Use the integrated cable management hook to wrap the cord. 4. Turn the connector locking handle counterclockwise. 5. Align the connector with the probe port and carefully push into place. 6. Turn the connector locking handle clockwise to secure the probe connector. 7. Carefully position the probe cord so it is free to move and is not resting on the floor.
GLOGIQ98 | PROBE_CABLE | Probe,Cable | Interface | LOGIQ9 | How to handle the Probe cable? | • Keep free from wheels • Do not bend the cable acutely • Avoid crossing cables between probes.
GLOGIQ99 | PROBE_ACTIVATION | Probe,Activation,Deactivation | interface | LOGIQ9 | How to manage probe? | To activate the probe, select the appropriate probe from the probe indicators on the Touch Panel. 1. Probe Indicators The probe's default settings for the mode and selected exam are used automatically. Deactivating the Probe When deactivating the probe, the probe is automatically placed in standby mode. Disconnecting the Probe Probes can be disconnected at any time. However, the probe should not be active when disconnecting the probe. • Ensure the probe is deactivated. Deactivate by selecting another probe or pressing Freeze. • Move the probe locking handle counterclockwise. • Pull the probe and connector straight out of the probe port. • Carefully slide the probe and connector away from the probe port and around the right side of the keyboard. • Ensure the cable is free. • Be sure that the probe head is clean before placing the probe in its storage box or a wall hanging unit. Freeze Deactivate the probe: 1. Press the Freeze key. 2. Gently wipe the excess gel from the face of the probe. 3. Carefully slide the probe around the right side of the keyboard, toward the probe holder. Ensure that the probe is placed gently in the probe holder. | https://drive.google.com/file/d/13mNkOWXPbrziZLT503m6JvHHFkgsVUlV/view?usp=sharing
GLOGIQ910 | CONTROL_PANEL_MAP | Control Panel | Control | LOGIQ9 | How to use Control panel? | 1. Touch panel and controls 2. Touch Panel Brightness Control 3. Video 4. Audio On/Off and Volume 5. TGC 6. Reverse 7. Additional Feature Keys 8. Keyboard 9. Mode/Gain Keys 10. Imaging/Measurement Keys 11. Depth 12. Imaging Feature Keys 13. Freeze and Print Keys 14. Probe and Cord Holder 15. Gel Holder | https://drive.google.com/file/d/1bvGNuuxhL67AmDIOXN4xzVCrpaAQ0zBa/view?usp=sharing
GLOGIQ911 | EXAM_FUNCTION_CONTROLS | Function,controls | Control | LOGIQ9 | How to manage exam function? | 1. Patient: Enters Patient screen 2. Scan: Enters scanning mode screen 3. Reports: Activates default report and Touch Panel of report choices. 4. End Exam: Activates Image Management and Touch Panel with end of exam options. 5. Utility: Activates system configuration menus. 6. Preset: Selects the application to use. 7. Probe Indicator: Indicates and selects the probes. | https://drive.google.com/file/d/1TSp6aIohRjjK9JT_q_qM8muQN0oqrg5E/view?usp=sharing
GLOGIQ912 | PATIENT_SEARCH | patient care | Exam | LOGIQ9 | How to search for an exsiting Patient? | 1. Press Patient to display the Patient Screen. The Patient Search Screen is displayed. 2. Select the search key and Location (From). Enter the search string. 3. An appropriate patient is displayed in the patient list. 4. Select the patient and press Review. or Click ‘+’ mark on the left side of Patient ID, then all exams of the patient is displayed in the patient list. Select the required exam and press Active Images. The image list of the exam is displayed. 5. The selected patient exam screen is displayed. | https://drive.google.com/file/d/1LGoVwktPdyhJtcCTwl8DqOTIJLyNcFAo/view?usp=sharing
GLOGIQ913 | IMAGE_REVIEWING | Image | Interface | LOGIQ9 | How to review the Image? | 1. Select the exam which includes the image to review. 2. Press Active Images. 3. Select the image and press Review. 4. The image is displayed. | https://drive.google.com/file/d/1LGoVwktPdyhJtcCTwl8DqOTIJLyNcFAo/view?usp=sharing
GLOGIQ914 | INSPECTING_THE_SYSTEM | System Care,Maintenance | Maintenance | LOGIQ9 | How to Examine the System? | • Connectors on cables for any mechanical defects. • Entire length of electrical and power cables for cuts or abrasions. • Equipment for loose or missing hardware. • Control panel and keyboard for defects. • Casters for proper locking operation. To avoid electrical shock hazard, do not remove panels or covers from console. This servicing must be performed by qualified service personnel. Failure to do so could cause serious injury.
GLOGIQ915 | CLEANING_AIR_FILTER | Air filter,Filter | Maintenance | LOGIQ9 | How to Clean the Air Filter? | Cleaning To clean the filter: 1. Be sure to shake the filter in an area away from the system. 2. Wash the filter in a mild soapy solution, rinse and air dry or dry with a cloth. NOTE: Allow wet filter to dry thoroughly before installing. Installing 1. Slide the filter back into the system. 2. Secure the filter with the Velcro tab.
GLOGIQ916 | SCANNING_A_PATIENT | Patient,Scanning | Exam | LOGIQ9 | How to Scan a New patient? | 1. Press Patient. 2. Press New Patient on the Patient menu. 3. A pop-up menu appears. Specify whether you want to store images permanently, delete images, or go to active images. 4. Choose the exam category. 5. Verify the dataflow. 6. Fill in patient information. 7. Press Register. Enter Past OB Exam information, if desired. 8. Select the probe from the Touch Panel to start scanning (or select Exit, Escape, Scan, or Freeze). 9. Perform the exam. 10. When you have completed the study, press End Exam, End Current Patient, then Store All Images Permanently. | https://drive.google.com/file/d/1_9PiSQH2HuSY-ll3v8_r9YAD1sClk2rJ/view?usp=sharing
GLOGIQ917 | PATIENT_LIST | Patient,List | Exam | LOGIQ9 | How to enter a Patient List? | 1. Press Patient to display Patient Screen. 2. Press New Patient to erase the current patient data. 3. Enter the Patient ID. 4. Enter the patient and exam information. 5. Press Register. 6. Repeat above steps as required. Select the patient from the Patient List and press New Exam to start an exam.
GLOGIQ918 | ERASE_EXISTING_PATIENT | Patient,Delete | Exam | LOGIQ9 | How to delete the Existing Patient? | 1. Search and select the patient in the patient list. 2. Press Delete. 3. The confirmation dialog box is displayed. 4. Press OK to delete or Cancel.
GLOGIQ919 | ERASE_EXISTING_EXAM | Exam,Delete | Exam | LOGIQ9 | How to Delete the Existing Exam List? | 1. Search and select the patient in the patient list. 2. Press Review. 3. The patient exam screen is displayed. Select the exam which you want to delete. 4. The confirmation dialog box is displayed. 5. Press OK to delete or Cancel.
GLOGIQ920 | ERASE_EXISTING_IMAGE | Image,Delete | Exam | LOGIQ9 | How to Delete the Existing Image list? | 1. Search and select the patient in the patient list. 2. Press Review. 3. The patient exam screen is displayed. Select the exam which include the image to delete. 4. Press Active Images to display the image list. 5. Select the image to delete and press Delete. 6. The confirmation dialog box is displayed. 7. Press Yes to delete or No to cancel.
GLOGIQ921 | DEAPTH_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Adjust Depth? | Each adjustment cycles you to the next Depth setting. Imaging and display parameters adjust automatically. To increase/decrease, adjust Depth.
GLOGIQ922 | GAIN_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Adjust Gain? | Gain values vary depending on the probe; they are not associated with a particular position of the knob. NOTE: TruAccess allows you to change the gain on a frozen image. B-Mode gain is independent of M-Mode and Doppler and Color Flow Gain. Changing the gain while in M-Mode does not affect the B-Mode image gain. To decrease/increase, rotate Gain.
GLOGIQ923 | FOCUS_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Adjust Focus? | To increase/decrease the number of focal zones, adjust Focus Number. NOTE: Focus Position and Focus Number share the same control. Press the control to toggle between Focus Position and Focus Number. To move the focal zone to the near/far field, adjust Focus Position. You can reverse how the Focus control functions on the Utility -> System -> General page.
GLOGIQ924 | B_FLOW_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Adjust B flow? | To activate/deactivate B Flow, press B Flow. Doppler Mode is available while in B Flow; however, M-Mode and Color Flow/PDI Modes are not available.
GLOGIQ925 | SRI_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Ajust SRI? | To activate SRI, adjust the SRI Touch Panel.
GLOGIQ926 | TIC_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Analyse TIC? | 1. Scan the patient after injecting the contrast agent. 2. Watch the agent flow through the anatomy of interest. 3. When the desired contrast effect has been visualized, freeze the image and select a range of images for analysis. 4. Position an ROI (region of interest) on one of those images where the contrast effect is visible. 5. The system then calculates the mean pixel intensity within that ROI for all frames in the user designated loop and plots the resulting data as a function of time.
GLOGIQ927 | ACTIVATING_TIC | Activation,Adjust | optimizing the image | LOGIQ9 | How to Activate TIC? | 1. Scan the patient in Contrast mode or select a desired cine loop from the stored images. 2. Press Freeze. 3. TIC Analysis displays on the Cine Touch Panel. 4. Select TIC Analysis. The TIC Analysis screen and Touch Panel displays. To toggle the trackball function between TIC and Scroll, press Scan Area. | https://drive.google.com/file/d/1fX4epc288CSrbIcGYe2wn1j7f1nY0Jwu/view?usp=sharing
GLOGIQ928 | EXAMINATION_M-MODE | Exam, | optimizing the image | LOGIQ9 | How to Examine using M-mode? | 1. Get a good B-Mode image. Survey the anatomy and place the area of interest near the center of the B-Mode image. 2. Press M/D Cursor. 3. Trackball to position the mode cursor over the area that you want to display in M-Mode. 4. Press M-Mode. 5. Adjust the Sweep Speed, TGC, Gain, Power Output, and Focus Position, as needed. 6. Press Freeze to stop the M trace. 7. Record the trace to disk or to the hard copy device. 8. Press Freeze to continue imaging. 9. To exit, press M-Mode. | https://drive.google.com/file/d/1mLsTVZBGHU_V--hOkXdmr0A-h1Nc44Ky/view?usp=sharing
GLOGIQ929 | EXAMINATION_DOPPLERMODE | Exam | optimizing the image | LOGIQ9 | How to Examine Using Doppler mode? | 1. Connect the appropriate probe, leaving the probes in their respective holders. 2. Position the patient for the examination. 3. Press Patient. Enter the appropriate patient data using the appropriate exam category and worksheet. 4. Select the preset, application and probe to be used. 5. Locate the anatomy to be examined. Get a good B Mode image. Refer to B-Mode section of this chapter for more information. 6. Press M/D Cursor to display the sample volume cursor and gate. or Press PW. The PW Doppler spectrum appears and the system operates in combined B+Doppler Mode. Adjust Volume to adjust Doppler audio. The Doppler signal is heard through the speakers. 7. Position the sample volume cursor by moving the Trackball left and right. Position the sample volume gate by moving the Trackball up and down. Size the gate by clicking SV Length. 8. Optimize the PW Doppler spectrum, as necessary. Refer to the Doppler Optimization section of this chapter for more information. 9. Press B Pause to toggle between real time B-Mode with sample volume and spectral display with audio. 10. Sample along the whole length of the vessel. Make sure that the probe is parallel to flow. Listen, then look, when positioning the sample volume cursor. 11. Press Freeze to hold the trace in memory and stop imaging. Activate CINE Timeline, as necessary. See ‘Activating CINE’ for more information. 12. Perform measurements and calculations, as necessary. 13. Record results by pressing the appropriate print key, depending on the setup of your recording devices. 14. Press Freeze to resume imaging. 15. Repeat the above procedure until all relevant flow sites have been examined. 16. Replace the probe in its respective holder.
GLOGIQ930 | DOPPLERMODE_ACTIVATION | Activation | optimizing the image | LOGIQ9 | How to Activate Doppler Mode? | To activate PW Doppler Mode, press PW. The Doppler spectrum displays along with the B-Mode image. The cursor changes to a Doppler cursor. You can now position and size the sample volume gate to get a velocity. Use Doppler Audio to listen for when the sample volume gate is positioned over an area of flow. | https://drive.google.com/file/d/1OFlU1k-endQLjOf_HZrbpYIz1H29vggt/view?usp=sharing
GCT52040 | FAST_CAL_PROCEDURE | Procedure,Maintain | Checking | OPTIMA CT520 | How to Maintain image quality? | 1. Display the Scan Monitor screen. – Clear the gantry opening. – Raise the table above the patient loading level. 2. Select [Daily Prep] and [Fast Cal] to display the screen. – Before the start of every scan day – Before Cal check – Before system calibration 3. The system automatically selects the Auto Scan function. – The system automatically selects the following sequence of scans: a) Mylar Window Check b) Cold Warmup c) Warmup 1 d) Warmup 2 e) Collimator Cal f) FPA Check Scans g) Clever Gain Calibration h) Fast Calibration
GCT52041 | SYSTEM_PREPRATION | System | Checking | OPTIMA CT520 | How to Prepare the System? | • Clean the Accessories and check for damage. • Check and remove dried contrast agent from: – Mylar ring (around the gantry opening) – Detector window – Table extension and cradle surfaces — especially the Patient Restraint plastic channels on the table – Accessories (Head holders, pads and cushions, etc.) • Check supplies.
GCT52042 | CHECK_IMAGE_SPACE | System,Image | Checking | OPTIMA CT520 | How to Check image space? | • Check the daily schedule, and multiply the list of patients by the estimated number of images each study requires. • Compare your estimate to the remaining 5122 images listed in the Feature Status Date and Time area. • If your estimate exceeds the available Image Space: – Film any previously unfilmed studies. (Optional) – Transfer designated images to another suite or console. – Archive and remove the oldest images from the system disk
GCT52043 | SYSTEM_SHUTDOWN | System,Shutdown | Power | OPTIMA CT520 | How to Shutdown the system? | • On Display Monitor, select the [Shutdown] icon. Dialog is posted • Select Shutdown and click OK. Dialog will be posted – The system is shutting down please wait... • When the system is down a message is posted – System Halted • Press the POWER button on computer panel. • To start the system from the System Halted prompt, power cycle the console. • Dialog is posted on the display monitor – The system is coming up please wait... • A dialog box will be posted – Attention OC is initializing please wait... • When this dialog disappears the system is ready to use. • If your system has the HIPAA login enabled, you are required to login to the system. • To turn off main disconnect control (MDC) or A1. • Shutdown the system.
GCT52044 | LASER_LIGHT_ACCURACY | Light,Laser,Accuracy | Test | OPTIMA CT520 | what are Procedure test for internal or external laser light Accuracy? | 1. Remove cradle pad. 2. Put the wire with the box on the cradle and use the level to position the box as horizontal as possible in both z and x directions. 3. Align the wire with the internal laser light field and make it parallel to the scan plane, and use coronal Laser light to center the wire in the up/down direction. 4. Landmark the wire using internal landmark. 5. Scan the wire using the scan protocol as listed in Table 13-1. 6. For the external laser light accuracy, move the cradle out and align the wire with the external laser light field, and use coronal laser light to center the wire in the up/down direction. 7. Landmark the wire using external landmark. 8. Scan the wire using the protocol as listed in the table
GCT52045 | TOMOGRAPHIC_ACCURACY | Accuracy | Test | OPTIMA CT520 | What are Procedure test for preview Tomographic Accuracy | 1. Take a scout scan as listed in Table 13-2. 2. Scan the wire using the helical protocol as listed in Table 13-2.
GCT52046 | SAGITTAL_AND_CORONAL_LIGHT_ACCURACY | Accuracy | Test | OPTIMA CT520 | What are procedure test for sagittal and coronal light accuracy? | 1. Remove the cradle pad. 2. Put the wire with the box flat on the cradle and use the level to position the box as horizontal as possible in both z and x direction. 3. Position the wire along the iso center using both sagittal (left/right) and coronal (up/down) laser light, and for this test, the wire should be perpendicular to the scan plane. 4. Landmark the wire using internal landmark. 5. Scan the wire using the protocol as listed in Table 13-3.
GCT52047 | SET_DOSE_ADMINISTRATOR_ROLE | Dose,Administrator,SEt | Check | OPTIMA CT520 | How to set Dose Administrator role? | 1. Click Dose Check Management from Protocol Management. 2. Click User Admin Tool . 3. From the EA3 Administration screen, type your Username and Password. Consult your service engineer or Application Representative for user name and password. 4. From the Local Users tab, click Add Local User. 5. From the pop-up panel, type information for each of the following: – A unique user ID – Full Name – Password – Confirm Password 6. Click Add User. 7. From the Groups tab, click Add Local Group. 8. From the Add Local Group window, type and enter a unique group name.9. Click Add Group. – The group is highlighted in the Local Groups list box. All information and buttons in the center panel refer to the highlighted group. 10. To change a group's roles, select the Dose Admin Role option box and click Apply Roles. – A green label confirms the applied roles. – An error message box displays if it is unsuccessful.
GCT52048 | SET_AV_EXCEEDING_USER_ROLE | User,Set | Check | OPTIMA CT520 | How to set AV exceeding user role? | 1. Click Dose Check Management from Protocol Management. 2. Click User Admin Tool . 3. From the EA3 Administration screen, type your Username and Password. Consult your service engineer or Application Representative for user name and password. 4. From the Local Users tab, click Add Local User. 5. From the pop-up panel, type information for each of the following: – A unique user ID – Full Name – Password – Confirm Password 6. Click Add User. 7. From the Groups tab, click Add Local Group. 8. From the Add Local Group window, type and enter a unique group name. 9.Click Add Group
GCT52049 | PROTOCOL_CHANGE_CONTROL | Change,Control | Interface | OPTIMA CT520 | How to change Protocol using Protocol Change Control? | 1. Click Dose Check Management from Protocol Management. 2. Click Logon Name. 3. Type the user name. 4. Click Password. 5. Type the User password.
GCT52050 | PROTOCOLS_WITH_NOTIFICATION_VALUES | Notification | Interface | OPTIMA CT520 | How to build protocols with Notification Values? | 1. From the scan monitor, click the Protocol Management icon . 2. On the Auto Voice and Protocol Management screen, click Protocol Management. 3. On the Protocol Selection screen, click User. 4. On the User tab click an Anatomical area. 5. In the Anatomical Area click on a protocol. 6. Click Next Series to go to the first non-scout series in the protocol. 7. From the Dose Information area click Set Up. 8. Type the desired CTDIvol or DLP NV value by group. – Only one CTDIvol or DLP value can be set for SmartPrep scan groups. 9. Click Info to return to the Dose Information display. 10. Click Accept to save the protocol parameters.
GCT52051 | ENERGY_SAVING_MODE | Energy | Power | OPTIMA CT520 | How to turn on energy saving mode? | 1. Disconnect any USB devices from the USB drive before performing Energy Saving mode. 2. From the display monitor, click Shutdown icon. A pop-up attention box appears. 3. From the Attention Shutdown pop-up screen, click Energy Saving Mode Getting Started 4. Click OK. Energy Saving Mode pop up appears: 5. From the Energy Saving Mode pop up, click to change the start up date if needed. 6. Click to change the start up hour if needed. Use up/down arrows to set values. 7. Click to change the start up minutes if needed. Use up/down arrows to set values. 8. Click OK to turn on energy saving mode. 9. The power off prompt is displayed. Then depress the power switch on the front of the operator console to the off position.
GCT52052 | SET_UP_ENERGY_SAVING_SCHEDULE | Energy | Power | OPTIMA CT520 | How to set up an energy saving schedule? | 1. From the display monitor, click Shutdown icon. 2. From the Attention Shutdown pop-up screen, click Energy Saving Mode. 3. Click OK. 4. Click Schedule to display the schedule set up screen. Energy Saving Mode Schedule Setup pop up appears: 5. Set the working days: 6. Click Select All to select all days. or Click each individual day to be desired. 7. Click to change the start up hour. 8. Click to change the start up minutes. 9. Click Accept to set the schedule.
GCT52053 | ADD_A_LOCAL_USER | Add | Interface | OPTIMA CT520 | How to Add a local User? | 1. Click [Add Local User].  When this button is clicked a popup panel is displayed. 2. Enter the new User ID (which must be unique). 3. Enter a Full Name. 4. Enter a Password. 5. Enter the Confirmed Password.  If any errors are encountered, you will receive an error messagebox. If you receive the error messagebox, changes were not committed to the database, and you can correct your errors and try again. Possible errors that can be encountered when adding a user are: – User ID already exists in the local user database (Choose a different unique username) – Password does not meet the minimum length requirements (Choose a longer password) – Password and Confirm Password box do not match (Make sure the passwords match) 6. Click [Add User].
GCT52054 | CHANGE_A_USER_PASSWORD | Change | Interface | OPTIMA CT520 | How to change a user password? | 1. Select User. 2. Click [Change Password].  This brings up a popup panel with two textboxes for the password. 3. Make changes to the password. 4. Click [Confirm Change].  If you do not want to make the change, simply click Cancel.  If the password doesn't meet the minimum length requirements, you will receive an error messagebox. If this occurs, your changes were not saved. Simply make the necessary corrections, and click Confirm Change again.
GCT52055 | REMOVE_LOCAL_USER | Change,User | Interface | OPTIMA CT520 | How to remove a local user? | 1. Select User. 2. Click [Remove User]. 3. Click [Confirm Removal].  If you do not want to make the change, simply click the Cancel button.
GCT52056 | ADD_A_LOCAL_GROUP | Add,Group | Interface | OPTIMA CT520 | How to Add a local Group? | 1. Click [Add Local Group].  When this button is clicked a popup panel is displayed. 2. Enter the new group's name (which must be unique).  If any errors are encountered, you will receive an error message box. If you receive the error message box, changes were not saved to the database, and you can correct your errors and try again. Possible errors that can be encountered when adding a group are: – Group name already exists in the database – Application session timeout  Once a local group is added, it is automatically highlighted in the Local Groups list box on the left-hand site, and it is highlighted. Once a group is highlighted, all information and buttons in the center panel (i.e. Group Name, Remove Group Button, Roles checkboxes, Apply Roles button, Group Members list box, Add Membership button, and Remove Membership button) refer to that group. 3. Click [Add Group].
GCT52057 | PERFORM_AN_EXAM_SPLIT_AFTER_SCAN | Exam | Exam | OPTIMA CT520 | How to Perform an exam split after scan is complete? | 1. Complete an exam with multiple Requested Procedure ID or Accession Numbers have been selected. 2. Select the ImageWorks desktop. 3. Click on the exam you wish to split. 4. Select the series to split. 5. Click [Exam Split] from the side of the Browser.  The Exam Split menu will appear with the selected images displayed in the upper viewports. The system may display every image or skip some images depending on the total number of images selected.  If the exam you selected does not have Multiple Procedures, a dialog will be posted and Exam Split will exit.  You must select all procedures you wish to split up front when scanning is done. You cannot add them after the fact. 6. Select the images to be grouped together for a procedure.  Click to select first image and shift and click on the last image for the desired range of images per procedure code. 7. Select desired Requested Procedure ID or Accession Number from the available list which reflects the procedures selected from Patient Schedule for New Patient.  This is from the list of exams registered to this patient from ConnectPro.. 8. Select the receiving host from the Host Config pull down.  Host Selection will show GSPS if configured in VES mode and HES if configured in HES mode. 9. Adjust the Window Level if needed.  Set the window/level for the images using the mouse or W/L preference buttons. 10. Click [Add].  The images will be added to the procedure list. 11. Hilight the procedure or procedures you wish to send and click [Send].  Repeat until you have sent all procedures you have images added to.  The images are transferred to the desired host.  If the system is configured for VES, a GSPS object is sent to the remote station. If HES mode is configured, a new series of images will be sent to the remote hosts.  You can select more than one procedure if you are sending them to the same destination.  You can select the Browser to exit Exam Split before the images have completed transferring. A pop up message will display after the transfer is complete.
GCT52058 | BIOPSY_MODE | Biopsy | Interface | OPTIMA CT520 | How to use the Biopsy mode? | 1. From the view/edit screen, click [Biopsy Rx]. 2. Set the Biopsy Reference. 3. Enter Biopsy Location or click [Get Alignment Light Location]. 4. Enter the Number of Images to be acquired for each pass. 5. Enter Gantry Tilt, if necessary. 6. Click [Thickness]. 7. Enter the Image Interval. 8. Click [Confirm Biopsy Rx] or [Cancel].
GCT52059 | ADD_A_FILTER | Filter,Add | Image manage | OPTIMA CT520 | How to Add filter to the available filter? | 1. From the display monitor, click the ImageWorks icon . 2. Select a remote host from the Source menu. 3. From the Patient List controls, click the Filter icon to open the Filter Data screen. 4. Enter all desired fields. In the Filter Name field, enter a descriptive name for the filter. 5. Click Save. 6. Click OK to save the filter.
GCT52060 | FILTER_THE_LIST | List,Filter | Image manage | OPTIMA CT520 | How to filter the list? | 1. From the display monitor, click the ImageWorks icon. 2. Select a remote host from the Source menu. 3. Select a filter from the Available Filters. 4. Click OK.
GCT52061 | REMOVE_A_PATIENT_LIST_FILTER | Filter,LIst | Image manage | OPTIMA CT520 | How to remove patient list filter? | 1. From the display monitor, click the ImageWorks icon . 2. From the Source menu, select a remote host. 3. From the Patient List controls, click the Filter icon to open the Filter Data screen. 4. From the Available filters menu, select a filter. 5. Click Remove. 6. Click Yes to remove the filter.
GCT52062 | UPDATE_PATIENT_LIST | Patient,List | Image manage | OPTIMA CT520 | How to Update the patient list? | 1. From the display monitor, click the ImageWorks icon . 2. From the Patient List controls, click the Refresh icon .
GCT52063 | FIND_AN_EXAM | Exam | Image manage | OPTIMA CT520 | How to find an exam? | 1. From the display monitor, click the ImageWorks icon . 2. In the Exam area, click Find and select a category. For example, search an exam by Exam. 3. In the Find text box, type the exam number you want to locate. For example, 121.
GCT52064 | ZOOM_THE_IMAGES | Images | Image manage | OPTIMA CT520 | How to zoom the images? | 1. From the display monitor, click the ImageWorks icon . 2. Click an image in the Image List area. 3. Adjust the zoom.
GCT52065 | SCROLL_THE_IMAGES | Images | Image manage | OPTIMA CT520 | How to scroll the images? | 1. From the display monitor, click the ImageWorks icon . 2. Click an image in the Image List area. 3. Adjust the scroll.
GCT52066 | LOCAL_DB_RETRIEVE_IMAGES | Images | Image manage | OPTIMA CT520 | How to Restore the images to our system from a remote host? | 1. From the display monitor, click the ImageWorks icon . 2. Click Source and select a remote Archive host. 3. Optional: Refine the remote Patient List. 4. From the remote Patient List, select the exam, series, or images you want to retrieve. 5. In the Destinations area of the Patient List, click Local DB to begin the restore process. 6. View the archive or network status in the Feature Status area to view the state of the images as they are transferred to the local database.
GCT52067 | ARCHIEVE_A_GSPS_SERIES | Images | Image manage | OPTIMA CT520 | How to archive or network a GSPS1 series when you have saved images with the Save State function.? | 1. Press Ctrl and simultaneously select the source series and the associated GSPS series from the Patient List. 2. To archive or network, continue with the Manual archive/network procedure.
GCT52068 | VIEW_THE_BACKLOG | Images | Image manage | OPTIMA CT520 | How to view a backlog of archive or network activity or to delete an archive or network job? | 1. From the display monitor, click the ImageWorks icon . 2. Click the Job Management icon . 3. From the Job Management screen, click Show Jobs and select an option to refine the lists. 4. From the Registered Jobs list, select the jobs you would like to change. 5. Choose the desired job action: 6. Click Close.
GCT52069 | DETACH_MEDIA | Images,CD,DVD | Image manage | OPTIMA CT520 | How to detach a CD, DVD, or USB media device? | 1. Make sure a CD, DVD, or USB is inserted in the appropriate drive. 2. Check the Archive/Network backlog to verify that all image transfer from the device is complete. 3. From the display monitor, click the ImageWorks icon . 4. Click Tools > Detach Media. 5. From the Detach Media screen, select the media to detach from the menu. 6. Click Detach and wait for the "Media Detached" message before ejecting the device from the drive. 7. Click Close to exit.
GCT52070 | CHECK_NETWORK_HISTORY | Image | Image manage | OPTIMA CT520 | How to Check the network history? | 1. From the display monitor, click the ImageWorks icon . 2. Click the Job Management icon . 3. From the Job Management screen, click the Network History tab. 4. Look through the log to find the exam number of the network history you wish to see. 5. Type a specific exam/series/image number into the Search ESI or Host field. 6. Click the Refresh icon to update the network history log. 7. Click Close.
GCT52071 | SAVE_IMAGES | Image | Image manage | OPTIMA CT520 | How to save images from your system to a CD/DVD or USB device? | 1. From the Data Apps list, click CD/DVD/USB. 2. From the Media Creator screen , select the media type from the Media menu. 3. To save a Viewer with the images, select Include Image Viewer. 4. Select the desired exams or series from the Patient List. 5. From the Media Creator screen, click Add. 6. Optional: Clear exams or series from the Media Creator list. 7. Click the Copy icon to start the recording process. 8. Once the recording is finished, click Close icon . 9. To view the Media Creator screen again, click CD/DVD/USB from the Data Apps list. 10. When the contents have been burned to the media, a message posts in the Message area. It is now safe to remove the CD/DVD or USB. 11. Click the Eject icon to remove the media from the drive.
GCT52072 | RESTORE_IMAGES | Image | Image manage | OPTIMA CT520 | How to restore images from a CD/DVD or USB device to your system? | 1. Place a CD-R or DVD-R in the CD/DVD external drive or place a USB device into one of the USB drives. Wait until the drive light turns off, which is an indication that the CD-R or DVD-R is spinning up to speed. 2. From the display monitor, click the ImageWorks icon . 3. From the header area, click Source and select the drive that has the CD/DVD or USB device in it. 4. From the Patient List, select the exam/series/images you want to restore. 5. In the Destinations area of the Patient List, click Local DB to begin the restore process. 6. To view the restore process status, see the View the backlog/queue procedure. 7. When the restore process is finished, indicated by a messa
GCT52073 | VIEW_IMAGES_ON_PC | Image | Image manage | OPTIMA CT520 | How to view images from a CD/DVD inserted in your PC? | 1. Load a recorded CD-R or DVD-R into the drive of your PC or laptop running Windows XP, Windows Vista, or Windows 7 operating system and containing Java 1.5 or higher versions. 2. Click Accept to the license agreement. 3. From the DataSelector, select an exam and wait until the data is loaded. 4. Select a series to view. 5. Use the Media Viewer image controls, located at the top of the Media Viewer, to manipulate the images.
GCT52074 | COMPOSE_A_REPORT | Image | Image manage | OPTIMA CT520 | How to compose a report? | 1. From the display monitor, click the ImageWorks icon . 2. From the Patient List, select the series you want to export. 3. From the Data Apps list, click Data Export. 4. Review the images in the Compose image viewport. 5. Click the Conversion Format menu and select an image format for the currently selected data set that is compatible with the movie player on your PC or laptop. 6. In the Image Range Selection area, choose your desired image range. 7. Move the Compression Factor slider to select a value. 8. In the Annotation area, choose your desired annotation display format. 9. Once you are satisfied with the image appearance (W/L, zoom, scroll), select Propagate Image Operations. 10. Type a name for both the report and the folder (use no spaces or characters other than alpha numeric). 11. Click Add to Report. 12. To add another data set to the report, repeat these steps. 13. Click the Export tab to export the report (for details see the Export a report procedure) or click Quit to exit the Data Export application.
GCT52075 | VIEW_A_REPORT_ON_PC | Image | Image manage | OPTIMA CT520 | How to view a report on PC? | 1. Place the CD in the CD drive of a PC or laptop running Windows 2000 or XP. 2. The CD launches automatically. 3. When finished viewing the report, select File > Close to close the Internet browser. 4. Remove the CD from the drive and store it.
GCT52076 | EDIT_PATIENT_DATA | Image | Image manage | OPTIMA CT520 | How to edit patient data? | 1. From the display monitor, click the ImageWorks icon. 2. From the Patient List, select the exam to be edited. 3. From the Data Apps list, click Edit Patient. 4. From the Edit Patient Applications screen, click Edit Patient Data. 5. Read the cautions and then, click Accept to continue. 6. From the Edit Patient Data screen, complete the text fields. 7. Click Accept when you are satisfied with the changes. 8. Click Accept to the confirmation prompt.
GCT52077 | REORDER_THE_DATA_APP_LIST | Image | Image manage | OPTIMA CT520 | How to change the order of the applications on the Data Apps list so that the most frequently used applications are at the top of the list? | 1. From the display monitor, click the ImageWorks icon . 2. From the ImageWorks desktop , click Tools > Applications Reorder. 3. On the Application Selection screen, click the item that you want to move in Data Apps list. 4. Click the up/down arrows to move the item to the new location. 5. Once the items in each list are in the desired order, click OK. 6. Click OK to the acknowledgement prompt.
GCT52078 | EXPORT_PROTOCOLS | Image | Image manage | OPTIMA CT520 | How to export user protocols to a CD so that you can load the CD on a PC and print. The format is excel? | 1. From the display monitor, click ImageWorks icon. 2. From the Tool Chest screen, click Export Protocols. 3. The CD/DVD tray opens and the system prompts you to place a blank CD media in the CD/DVD drive. 4. Place a CD inserted in the tray. 5. Click OK to close the drive and initiate the save process. – The User adult and pediatric protocols are saved. 6. The CD tray opens automatically when the save is complete. Remove the CD and click OK to the message prompt, which closes the tray.
GCT52079 | SET_SCREENSAVER_TIMEOUT | Image,Screensaver | Image manage | OPTIMA CT520 | How to set the time period for the scan and display monitor screen saver to start? | 1. From the display monitor, click the Service icon. 2. From the Service browser, click the Service browser title bar to move it so you can see the service Tool Chest. 3. From the Tool Chest, click Screensaver Time. 4. From the Set Screensaver Timeout screen, click and drag the slider to adjust the time that the screensaver mode starts to between 5 to 60 minutes. – Click on either side of the slider grab bar to increment in steps of 1 minute. – The default is 30 minutes. 5. Click Save.
GCT52080 | CHANGE_THE_ORIENTATION_OF_IMAGE | Image | Image manage | OPTIMA CT520 | How to flip an image left to right, top to bottom, or both at the same time. You can also rotate to the right or to the left 90 degrees? | 1. Put the image of interest in primary focus. 2. Click [Routine Display]. 3. Click [Flip/Rotate]. 4. Click [FLR] to flip image left to right, click [FTB] to flip image top to bottom or click [FTB/FLR] to flip image both top to bottom and left to right. 5. Click [RR] to rotate image to the right (clockwise) or click [RL] to rotate image to the left (counter-clockwise). 6. Click [Cancel] to cancel out of flip/rotate pop up window. 7. Click [Display Normal] to return images to their original orientation.
GCT52081 | MAKE_IMAGES_SHARPER | Image | Image manage | OPTIMA CT520 | How to make Images Sharper or Smoother? | 1. Put the image of interest in primary focus. 2. Click [Routine Display]. 3. Click [Proview]. 4. Click [E1, E2, E21, E22, E23, or E3] for edge enhancement. 5. Click [Lung] for lung enhancement. 6. Click [S1,S11, S2, S21, or S3] for smoothing. 7. Click [Off] to turn enhancement off. 8. Click [Cancel] to cancel out of Proview pop up window.
GCT52082 | IMPROVE_THE_BRAIN_INTERFACE | Image | Image manage | OPTIMA CT520 | How to improve the Brain/Bone Interface? | 1. Put the image of interest in primary focus. 2. Click [Routine Display]. 3. Click [GSE]. 4. Click [G1, G2 or G3]. 5. Click [Off] to turn enhancement off. 6. Click [Cancel] to cancel out of GSE pop up window.
GCT52083 | RESTORE_AN_IMAGE_TO_ITS_ORIGINAL_STATE | Image | Image manage | OPTIMA CT520 | How to Restore an Image to its Original State? | 1. Put the image of interest in primary focus. 2. Click [Routine Display]. 3. Click [Display Normal].
GCT52084 | REVERSE_THE_VIDEO | Video | Image manage | OPTIMA CT520 | How to Reverse the Video? | 1. Put the image of interest in primary focus. 2. Click [Display Preferences]. 3. Click [Inverse Video]. 4. Click [Inverse Video] again to turn feature off.
GCT52085 | APPLY_MATTES_TO_AN_IMAGE | Image | Image manage | OPTIMA CT520 | How to Apply Mattes to an Image? | 1. Put the image of interest in primary focus. 2. On the Accelerator Line, type ematte for an elliptical matte or rmatte for a rectangular matte and click Enter.
GCT52086 | CREATE_A_GSPS | Image | Image manage | OPTIMA CT520 | How to create GSPS? | 1. Put the image of interest in primary focus. 2. Adjust the image to the desired presentation. a) Set WW and WL b) Image Roam c) Image Zoom d) Image Flip e) Set rotate f) Add graphics such as ROI, Measure distance and Image annotation 3. On the Accelerator Line, type ss and click Enter.
GCT52087 | VIEW_A_GSPS | Image | Image manage | OPTIMA CT520 | How to View a Gray Scale Presentation State Object (GSPS)? | 1. Hilight the GSPS object in the Browser 2. Select [Viewer]. 3. Use the pull down to change the viewer format. 4. Select the ± CPI object buttons to view additional GSPS objects in the series.
GCT52088 | NEURO_3D_IMAGES | Image | Image manage | OPTIMA CT520 | How to Process Neuro 3D Images? | 1. Click [Neuro 3D]. 2. Click [Low], [Med], or [High]. 3. Click [Go]. 4. Click the door icon to close the window when the filtering is complete.
GCT52089 | TYPE_TEXT_ON_THE_IMAGE | Text | Image manage | OPTIMA CT520 | How to Type Text on the Image? | 1. Put the image that you want to type annotation on into primary focus. 2. Click [Routine Display]. 3. Click [User Annotation]. 4. Place mouse cursor over the image you are typing text onto. 5. Type the desired text and it appears in the blue box. 6. Adjust the box and attached arrow as needed. 7. Click [User Annotation] again to add additional annotation. 8. Click [Erase] to erase the primary (blue) annotation.
GCT52090 | DISPLAY_THE_EXAM_INFORMATION | Test,Exam | Image manage | OPTIMA CT520 | How to Display the Exam Information? | 1. Click [Routine Display]. 2. Click [Exam Page]. 3. Click [Manual Film] or [Auto Film] to select the film composer you would like the exam page to go to. 4. Click [Quit].
GCT52091 | SAVE_AN_IMAGE_SCREEN | Screen,Image | Image manage | OPTIMA CT520 | How to Save an Image Screen? | 1. Click [Routine Display]. 2. Click [Screen Save].
GLOGIQ492 | MONITOR_ASSY | Assy,Disassembly | Renewal Parts | LOGIQ 400 | How to Monitor Assy Disassembly? | 1. Turn OFF the system. 2. Remove the CRT Filter. Refer to 6–2–6 on page 6–79. 3. Unscrew four hexagonal screws ( 1 – 4 ) located beneath the horizontal part of the Neck Assy. 4. Remove the cover located at the bottom of the Neck Assy. 5. Unscrew two screws ( 5 and 6 ) and remove the plate fixing the cables. 6. Disconnect the four connectors ( 7 – 10 ). 7. Remove the screw cap located at position ( 11 ). 8. Loosen one screw ( 11 ) and pull the Monitor Assy upwards.
GLOGIQ493 | MONITOR_COVER_SET | Assy,Disassembly | Renewal Parts | LOGIQ 400 | How to Monitor cover set Disassembly? | 1. Turn OFF the system. 2. Unscrew three screws ( 1 – 3 ). 3. Remove the Monitor Cover L Assy (FRU 104). 4. Unscrew three screws ( 4 – 6 ). 5. Remove the Monitor Cover R Assy (FRU 105). 6. Remove the Monitor Cover Top Assy (FRU 106). 7. Push and open the Escutcheon Front Door (FRU 107). 8. Pull the two caps ( 7 , FRU 108: CRT Cap Set) out. 9. Loosen four screws ( 8 – 11 ). 10. Disconnect one connector of task lamp switch. 11. Remove the Escutcheon Assy (FRU 102).
GLOGIQ494 | ESCUTUHEON_FRONT_DOOR | Disassembly | Renewal Parts | LOGIQ 400 | How to Escutcheon Front Door removal? | 1. Turn OFF the system. 2. Pull out the Escutcheon Front Door (FRU 107) from the Escutcheon Assy (FRU 102)
GLOGIQ495 | CRT_CAP_SET | Disassembly | Renewal Parts | LOGIQ 400 | How to CRT Cap set removal? | 1. Turn OFF the system. 2. Remove the Escutcheon Front Door (FRU 107). 3. Pull out two CRT Caps from the front side of the Monitor Assy.
GLOGIQ496 | CRT_FILTER_REMOVAL | Disassembly | Renewal Parts | LOGIQ 400 | How to CRT filter Removal? | 1. Turn OFF the system. 2. Slide the left side filter clamp leftwards and pull it out. 3. Slide the right side filter clamp rightwards and pull it out. 4. Remove the CRT Filter.
GLOGIQ497 | ACCESSORY_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to Accessory Assy Disassembly? | 1. Turn OFF the system. 2. Remove the Monitor Cover Set (FRU 101). 3. Unscrew five screws ( 1 – 5 ). 4. Disconnect two connectors ( 6 and 7 ). 5. Remove the Accessory Assy.
GLOGIQ498 | SPEAKER_ASSY_DISASSEMBLY | Disassembly | Renewal Parts | LOGIQ 400 | How to Speaker assy Disassembly? | 1. Turn OFF the system. 2. Remove the Monitor Cover L Assy (FRU 104) and Monitor Cover R Assy (FRU 105). Refer to 6–2–2 on page 6–72. 3. Cut off the tie wrap fixing the CRT Cable Set (FRU 116) to the speaker. 4. Unscrew two screws ( 1 and 2 ). 5. Loosen two screws ( 3 and 4 ). 6. Disconnect one connector ( 5 ) from the back side of monitor. 7. Remove the speaker at the left. 8. Unscrew two screws ( 6 and 7 ). 9. Loosen two screws ( 8 and 9 ). 10. Disconnect one connector ( 10 ) from the back side of monitor. 11. Remove the speaker at the right.
GLOGIQ499 | CRT_LAMP_ASSY_DISASSEMBLY | Disassembly | Renewal Parts | LOGIQ 400 | How to CRT lamp assy Disassembly? | 1. Turn OFF the system. 2. Remove the Monitor Cover L Assy (FRU 104) and Monitor Cover R Assy (FRU 105). Refer to 6–2–2 on page 6–72. 3. Turn the CRT lamp counterclockwise and remove it. 4. Unscrew two screws ( 1 and 2 ). 5. Disconnect one connector ( 3 ). 6. Remove the CRT Lamp Assy.
GLOGIQ4100 | CRT_ASSY_DISASSEMBLY | Disassembly | Renewal Parts | LOGIQ 400 | How to CRT Assy Disassembly? | 1. Turn OFF the system. 2. Remove the Monitor Assy (FRU 100). Refer to 6–2–1 on page 6–70. 3. Remove the Monitor Cover Set ( 1 : FRU 101). Refer to 6–2–2 on page 6–72. 4. Remove the Speaker Assy ( 2 : FRU 112). Refer to 6–2–8 on page 6–82. 5. Remove the tilt mechanism.
GLOGIQ4101 | MONITOR_ASSY_NTSC_DISASSEMBLY | Disassembly | Renewal Parts | LOGIQ 400 | How to Monitor Assy NTSC Disassembly? | 1. Turn OFF the system. 2. Unscrew four hexagonal screws ( 1 – 4 ) located beneath the horizontal part of the Neck Assy. 3. Remove the cover located at the bottom of the Neck Assy. 4. Unscrew two screws ( 5 and 6 ) and remove the plate fixing the cables. 5. Disconnect the three connectors ( 7 – 9 ). 6. Remove two screw caps located at position ( 10 and 11 ). 7. Loosen two screws ( 10 and 11 ) and pull the Monitor Assy upwards.
GLOGIQ4102 | MONITOR_COVER_DISASSEMBLY | Disassembly | Renewal Parts | LOGIQ 400 | How to Monitor Cover Disassembly? | 1. Turn OFF the system. 2. Remove the Monitor Bottom Assy. Refer to 6–2–17 on page 6–98. 3. Unscrew four screws ( 1 – 4 ). 4. Remove the Monitor Cover (FRU 152).
GLOGIQ4103 | NECK_ASSY_DISASSEMBLY | Disassembly | Renewal Parts | LOGIQ 400 | How to Neck Assy Disassembly? | 1. Turn OFF the system and unplug the unit. 2. Push down the Up/Down Release Button located on the Neck Assy, lift the monitor up and set the monitor to the highest position. 3. Remove the Monitor Assy (FRU No. 100), Monitor Assy NTSC (FRU No. 150) or Monitor Assy PAL (FRU No. 151). Refer to 6–2–1 on page 6–70 or 6–2–12 on page 6–90. 4. Remove the Neck Space Plate. 5. Remove the Front Base Cover (FRU 305). Refer to 6–2–35 on page 6–131. 6. Remove the Side Cover Right (FRU 302). Refer to 6–2–32 on page 6–128.. 7. Remove the Side Cover Left (FRU 301). Refer to 6–2–31 on page 6–127. 8. Remove the Rear Cover (FRU 303). Refer to 6–2–33 on page 6–129. 9. Remove the Top Cover (FRU 308). Refer to 6–2–38 on page 6–135. 10. Remove the Rear CONN Panel Assy (FRU 501). Refer to 6–2–61 on page 6–166. 11. Disconnect one connector ( 1 ). 12. Disconnect two connectors ( 2 and 3 ) from the P25A and P26A connectors on Mother Board. 13. Disconnect one flat cable ( 4 ). 14. Cut off the tie wraps to bind cables. 15. Remove thirteen hexagonal bolts ( 5 – 17 ). 16. Remove the Neck Assy from console.
GLOGIQ4104 | NECK_SPACE_PLATE | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassembly Neck Space Plate? | 1. Turn OFF the system. 2. Remove the Monitor Assy (FRU No. 100), Monitor Assy NTSC (FRU No. 150) or Monitor Assy PAL (FRU No. 151). Refer to 6–2–1 on page 6–70 or 6–2–12 on page 6–90. 3. Remove the Neck Space Plate 2. Refer to 6–2–29 on page 6–122. The Neck Space Plate 2 is used only for the B/W monitor console. In case of color monitor console, skip this step. 4. Remove the Rotation Spacer. Refer to 6–2–30 on page 6–124. The Rotation Spacer is used only for the B/W monitor console. In case of color monitor console, skip this step. 5. Remove the Neck Space Plate.
GLOGIQ4105 | ROTATION_SPACER | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Rotation spacer? | 1. Turn OFF the system. 2. Remove the Monitor Assy NTSC (FRU No. 150) or Monitor Assy PAL (FRU No. 151). Refer to 6–2–12 on page 6–90. 3. Unscrew one screw ( 1 ). 4. Remove the Rotation Spacer.
GLOGIQ4106 | SIDE_COVER_LEFT | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Side Cover left? | 1. Turn OFF the system. 2. Remove two screw caps (located at 3 and 4 ) and unscrew four screws ( 1 – 4 ). 3. Remove the Side Cover Left.
GLOGIQ4107 | SIDE_COVER_RIGHT | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Side cover Right? | 1. Turn OFF the system. 2. Remove the Front Base Cover (FRU No. 305). Refer to 6–2–35 on page 6–131. 3. Remove the two screw caps (located at 3 and 4 ) and unscrew five screws ( 1 – 55 ). 4. Remove the Rear Door Assy. Refer to 6–2–34 on page 6–130. 5. Remove the Side Cover Right.
GLOGIQ4108 | REAR_COVER | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Rear Cover? | 1. Turn OFF the system. 2. Pull the Air Filter out. 3. Remove the Rear Door Assy. Refer to 6–2–34 on page 6–130. 4. Unscrew eight screws ( 1 – 8 ). 5. Remove the Rear Cover.
GLOGIQ4109 | REAR_DOOR_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble rear door assy? | 1. Turn OFF the system. 2. Turn the knob on the Rear Door Assy counterclockwise. 3. Remove the Rear Door Assy.
GLOGIQ4110 | KEYBOARD_BOTTOM_COVER_DISASSEMBLE | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Keyboard Bottom cover? | 1. Turn OFF the system. 2. Remove the Front Base Cover (FRU 305). Refer to 6–2–35 on page 6–131. 3. Remove the Front Cover (FRU 306). Refer to 6–2–36 on page 6–132. 4. Remove the Side Cover Right (FRU 302). Refer to 6–2–32 on page 6–128. 5. Unscrew six screws ( 1 – 6 ). 6. Remove the Keyboard Bottom Cover.
GLOGIQ4111 | TRACK_BALL | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Trackball? | 1. Turn OFF the system. 2. Remove the Keyboard Panel Assy. Refer to 6–2–49 on page 6–147. 3. Unscrew four screws ( 1 – 4 ). 4. Disconnect one connector ( 5 ). 5. Unscrew four screws ( 6 – 9 ). 6. Remove The Holding Metal Plate.
GLOGIQ4112 | GAIN_ENCODER | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Gain Encoder? | 1. Turn OFF the system. 2. Remove the Keyboard Panel Assy. Refer to 6–2–49 on page 6–147. 3. Unscrew four screws ( 1 – 4 ). 4. Disconnect the gain encoder connector ( 5 ).
GLOGIQ4113 | KEYBOARD_KNOBE_SET | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Keyboard knobe set? | 1. Turn OFF the system. 2. Loosen two hexagonal bolts for each encoder knob. 3. Remove six Encoder Knobs ( 1 – 6 ). 4. Pull the eight TGC Pot Knobs ( 7 ) out
GLOGIQ4114 | TGC_DISASSEMBLY | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble TGC? | 1. Turn OFF the system. 2. Remove the Keyboard Panel Assy. Refer to ILLUSTRATION 6–77 on page 6–155. 3. Unscrew four screws ( 1 – 4 ). 4. Disconnect one connector ( 5 ).
GLOGIQ4115 | ROTARY_ENCODER | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Rotary Encoder? | 1. Turn OFF the system. 2. Remove the Keyboard Panel Assy. Refer to ILLUSTRATION 6–77 on page 6–155. 3. Unscrew four screws ( 1 – 2 ). 4. Disconnect the gain encoder connector ( 3 ).
GLOGIQ4116 | REAR_CONN_SPACER | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Rear conn spacer? | 1. Turn OFF the system. 2. Remove the Rear Door Assy (FRU 304). Refer to 6–2–34 on page 6–130. 3. Remove the Rear Cover Assy (FRU 303). Refer to 6–2–33 on page 6–129.
GLOGIQ4117 | CIRCUIT_PROTECTER | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble circuit Protector? | 1. Turn OFF the system. 2. Remove the Rear CONN Panel Assy (FRU 501). Refer to 6–2–61 on page 6–166. 3. Remove the soldered cord ( 1 and 2 ). 4. Remove the Circuit Protector toward the direction indicated by the arrow.
GLOGIQ4118 | NEST_FAN_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Nest Fan Assy? | 1. Turn OFF the system. 2. Remove the Side Cover Left (FRU 301). Refer to 6–2–31 on page 6–127. 3. Disconnect two connectors ( 1 and 2 ). 4. Unscrew two screws ( 3 and 4 ). 5. Pull out the Nest Fan Assy leftwards.
GLOGIQ4119 | CONNECTOR_COVER | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Connector cover? | 1. Turn OFF the system. 2. Unscrew four screws ( 1 – 4 ). 3. Remove the Connector Cover.
GLOGIQ4120 | SHIELD_FINGER_LONG | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble Shield Finger long? | 1. Turn OFF the system. 2. Remove the Front Cover (FRU 306). Refer to 6–2–36 on page 6–132. 3. Unscrew eight screws ( 1 – 8 ). 4. Remove the connector frame. 5. Remove the Shield Finger Long (FRU 515) or Shield Finger Short (FRU 516).
GLOGIQ4121 | PROBE CONN 1 ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble conn 1? | 1. Turn OFF the system. 2. Remove the Probe CONN Set Assy (FRU 512/513). Refer to 6–2–68 on page 6–176. 3. Remove the PRAG Assy (FRU 519). Refer to 6–2–72 on page 6–182. 4. Unscrew eight screws ( 1 – 8 ). 5. Remove the connector frame. 6. Unscrew twelve screws ( 9 – 20 ). 7. Unscrew three screws (21 – 23 ). 8. Unscrew two studs (24 and 25 ). 9. Remove the Probe CONN 1 Assy or DCON Assy.
GLOGIQ4122 | PRAG_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to Disassemble PRAG Assy? | 1. Turn OFF the system. 2. Remove the Probe CONN Set Assy (FRU 512/513). Refer to 6–2–68 on page 6–176. 3. Unscrew seven screws ( 1 – 7 ). 4. Unscrew three screws ( 8 – 10 ) to remove the metal bracket. 5. Disconnect two connectors ( 11 and 12 ). 6. Remove the PRAG Assy.
GLOGIQ4123 | MODD_ASSY_REMOVAL | Disassembly | Renewal Parts | LOGIQ 400 | How to remove MODD assemble? | 1. Turn OFF the system. 2. Remove the Keyboard Panel Assy (FRU 400). Refer to 6–2–49 on page 6–147. 3. Disconnect the MODD Fan Power Cable. 4. Unscrew two screws ( 1 and 2 ). 5. Remove the MODD Fan Assy 6. Remove the HDDB Assy (FRU 525). Refer to 6–2–75 on page 6–195. 7. Unscrew four screws ( 3 – 6 ). 8. Remove the upper metal plate. 9. Disconnect the SCSI Cable and MODD Power Cable. 10. Remove the MODD Holder Assy. 11. Unscrew four screws ( 7 – 10 ). 12. Remove the MODD.
GLOGIQ4124 | HDD_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to remove HDD assemble? | 1. Make sure the customer data is saved on a MO disk. Refer to 4–3–3 (D), USER DATA BACK UP on page 4–22. 2. Turn OFF the system. 3. Remove the Keyboard Panel Assy (FRU 400). Refer to 6–2–49 on page 6–147. 4. Remove the HDDB Assy (FRU 522). Refer to 6–2–75 on page 6–195. 5. Unscrew two screws ( 1 and 2 ). 6. Remove the HDD Holder Assy. 7. Unscrew four screws ( 3 – 6 ). 8. Remove the HDD (FRU 523) out of the HDD Holder Assy (FRU 524).
GLOGIQ4125 | HDDB_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to remove HDDB Assemble? | 1. Turn OFF the system. 2. Remove the Keyboard Panel Assy (FRU 400). Refer to 6–2–49 on page 6–147. 3. Disconnect three connectors ( 1 – 3 ). 4. Unscrew four screws ( 4 – 7 ). 5. Remove the HDDB Assy.
GLOGIQ4126 | HDD_LED_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to remove HDD LED assemble? | 1. Turn OFF the system. 2. Remove the Front Base Cover (FRU 305). Refer to 6–2–35 on page 6–131. 3. Remove the Front Cover (FRU 306). Refer to 6–2–36 on page 6–132. 4. Unscrew two screws ( 1 and 2 ). 5. Disconnect one connector ( 3 ). 6. Remove the HDD LED Assy.
GLOGIQ4127 | POWER_SWITCH_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to remove power switch assemble? | 1. Turn OFF the system. 2. Remove the Front Base Cover (FRU 305). Refer to 6–2–35 on page 6–131. 3. Remove the Front Cover (FRU 306). Refer to 6–2–36 on page 6–132. 4. Unscrew two screws ( 1 and 2 ). 5. Disconnect one connector ( 3 ). 6. Remove the Power Switch Assy.
GLOGIQ4128 | TIME_KEEPER_RAM | Disassembly | Renewal Parts | LOGIQ 400 | How to remove Time Keeper Ram? | 1. Turn OFF the system. 2. Remove the Side Cover Left (FRU 301). Refer to 6–2–31 on page 6–127. 3. Unscrew five Nest Cover screws. 4. Remove the Nest Cover. 5. Remove the MPU Assy (FRU 612). Refer to 6–2–78 on page 6–198. 6. Remove the Time Keeper RAM mounted on the CPU Assy.
GLOGIQ4129 | HIGH_VOLTAGE_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to remove high voltage Assy? | 1. Turn OFF the system and unplug the main power cord. 2. Remove the Front Base Cover (FRU 305). Refer to 6–2–35 on page 6–131. 3. Remove the Side Cover Right (FRU 302). Refer to 6–2–32 on page 6–128. 4. Unscrew four screws ( 1 – 4 ). 5. Remove the PS Cover Right. 6. Unscrew two screws ( 5 and 6 ). 7. Disconnect five connectors ( 7 – 11 ). 8. Pull the High Voltage Assy out. 9. Disconnect the HV Fan Cable ( 12 ). 10. Remove the High Voltage Assy.
GLOGIQ4130 | HV_FAN_ASSY | Disassembly | Renewal Parts | LOGIQ 400 | How to remove HV fan Assy? | 1. Turn OFF the system and unplug the main power cord. 2. Remove the High Voltage Assy (FRU 701). Refer to 6–2–85 on page 6–214. 3. Unscrew four screws ( 1 – 4 ). 4. Remove the HV Fan.
GLOGIQ4131 | LOW_VOLTAGE_UNIT_DISASSEMBLE | Disassembly | Renewal Parts | LOGIQ 400 | How to remove Low voltage Assemble? | 1. Turn OFF the system and unplug the main power cord. 2. Remove the Side Cover Left (FRU 301). Refer to 6–2–31 on page 6–127. 3. Unscrew four screws ( 1 – 4 ). 4. Remove the Left PS Cover. 5. Unscrew two screws ( 5 and 6 ). 6. Pull the Low Voltage Unit outwards. 7. Unscrew fifteen screws ( 7 – 21 ) to remove wires. 8. Pull the Low Voltage Unit out. 9. Remove the Plastic Cover as shown in ILLUSTRATION 6–122 on page 6–219. 10. Disconnect one connector ( 22 ). 11. Unscrew two screws ( 23 and 24 ) to remove the FG wires. 12. Remove the Low Voltage Unit.
GLOGIQ4132 | TRIAC_UNIT | Disassembly | Renewal Parts | LOGIQ 400 | How to remove TRIAC Assemble? | 1. Turn OFF the system and unplug the main power cord. 2. Remove the Side Cover left. Refer to 6–2–31 on page 6–127. 3. Remove the Left PS Cover. Refer to 6–2–95 on page 6–234. 4. Unscrew two screws ( 1 and 2 ). 5. Disconnect eleven (11) connectors ( 3 – 13 ) from the component side of the TRIAC Unit. 6. Pull the TRIAC Unit outwards. 7. Disconnect five connectors (14 – 18 ) from the NF2 located on the back side of the TRIAC Unit. 8. Disconnect two connectors (19 and 20 ) from the rear of fuse socket. 9. Remove the TRIAC Unit.
GLOGIQ4133 | VISUAL_INSPECTION | Visual,Maintenance | Maitenance | LOGIQ 400 | What are the Procedures for visual inspection? | a. Check cables for cracks nicks and abrasions of their insulation. b. Check that cables sheath are securely clamped in cable glands/connector clamps where applicable. c. Check that cables are correctly routed and secured. d. Check that all screws are in place and not missing. e. Check that all detachable cable connectors are fully home in their sockets. f. Check that all equipment covers and fixing are in place and undamaged. g. Check that all plug–in boards are fully seated in cardcage connector. h. Check operation of the cooling fans and that the airflow paths are clear. LOGIQ 400 has five Fans : three for PC Boards, one for Power Supplies, one for MO drive. i. Check that the keyboard keys, some volumes and Track Ball have correct function and no mechanical damage. j. Check that three Air Filters indicated in the ILLUSTRATION 7–2 are clean and undamaged.
GOPXR646134 | SYSTEM_START_UP | Start,System | Power | OPTIMA XR646 | How to Start the system? | 1. Press the Power On button on the RCIM [Radiology Control Interface Module]. 2. Wait a few minutes until the entire system is powered.  The system powers up automatically.  If enabled, the Login screen appears on the monitor when the system is ready.  If Login is not enabled, the Worklist appears on the monitor when the system is ready.. 3. Wait for the detector to boot. This takes approximately 30 seconds.  A message in the system status area will appear: “Detector Boot is in progress. Please wait 29 seconds for the detector boot to complete.”  The system is ready when the detector boot message disappears.
GOPXR646135 | SHUTDOWN | Shutdown,System | Power | OPTIMA XR646 | How to Shutdown the system? | 1. Close all current exams. 2. Click the [UTILITY] button at the top of the Worklist screen. 3. Select System on the Utilities screen. 4. Click [SHUTDOWN].  A message appears: “The system will be shut down. NOTE: Please allow for the tube fan speed to slow (if running high) before shutting down system.” 5. Click [YES] to proceed with shut down.  The system powers off and the monitors go blank.  [CANCEL] stops the system from shutting down and returns you to the Utilities screen.
GOPXR646136 | STANDARD_LOGIN | Login | Interface | OPTIMA XR646 | How to login the system? | 1. Start up the system or log off the previous user.  The Login screen appears. 2. Enter your Login Name. 3. Enter your Password. 4. Click [LOGIN].
GOPXR646137 | LOG_OFF | Log off | Interface | OPTIMA XR646 | How to log off? | 1. Close, suspend, or discontinue any open exams, if necessary. 2. Close the Image Viewer, if necessary. 3. Click [LOGOFF] at the top of the Worklist or Image Management screen.  Or open the Utility screen, go to System and click [LOGOFF]  A message appears: “Do you really want to log off?” 4. Click [OK].  The Login screen appears.  [CANCEL] closes the screen and returns you to the last screen.
GOPXR646138 | EMERGENCY_STOP_BUTTON | Stop | Interface | OPTIMA XR646 | How to stop the system in Emergency condition? | 1. In an emergency situation, press the Emergency Stop button in with force. 2. Resolve the emergency situation. 3. When normal conditions are confirmed, turn (RCIM) or pull (table) the e-stop button to reset.  The system will power up automatically.
GOPXR646139 | POSITION_THE_TABLE | Position | Hardware | OPTIMA XR646 | How to Position the Table Longitudinally and Transversely? | 1. Release the table lock, if necessary. 2. Press the table top positioning foot pedal two consecutive times (“double-tap”). This activates the foot pedal. 3. Hold the foot pedal down and position the table top.  You can float the table top in all directions while the pedal remains held down. 4. Manually move the tabletop in a longitudinal or transverse direction to the desired position. 5. Release the foot pedal to lock the tabletop.
GOPXR646140 | ADJUST_OTS | Adjust | Hardware | OPTIMA XR646 | How to Adjust the Overhead Tube Suspension (OTS) Position? | 1. Use the Longitudinal Lock Release button to move the OTS along the bridge of the overhead rail system. a) Press and hold the Longitudinal Lock Release button on the User Interface. b) Move the OTS to the desired position. c) Release the Longitudinal Lock Release button. 2. Use the Vertical Lock Release button to move the telescopic column up and down. a) Press and hold the Vertical Lock Release button on the User Interface. b) Move the telescopic to the desired vertical position. c) Release the Vertical Lock Release button. 3. Use the Lateral Lock Release button to move the OTS carriage from side to side on the bridge. a) Press and hold the Lateral Lock Release button on the User Interface. b) Move the OTS carriage to the desired position. c) Release the Lateral Lock Release button. 4. Use the All-Lock, Lock Release button to simultaneously move the OTS in vertical, lateral and longitudinal directions. If detents are on and the OTS is in lateral detent the system will maintain the lateral position but move freely in vertical an longitudinal. a) Press and hold the All-Lock, Lock Release button on the User Interface. b) Move the OTS to the desired position. c) Release the All-Lock, Lock Release button.
GOPXR646141 | ADJUST_THE_TUBE_POSITION | Adjust | Hardware | OPTIMA XR646 | How to Adjust the Tube Position? | 1. Use the Tube Angulation Lock Release button to rotate the tube about the short axis (cranial to caudal). a) Press and hold the Tube Angulation Lock Release button on the User Interface. b) Move the tube unit to the desired angle. c) Release the Tube Angulation Lock Release button. 2. Use the OTS Rotation Detent Release lever to rotate the tube about the vertical axis of the telescopic column. a) Press the Rotation Detent Release lever on the right side of the OTS. b) Rotate the tube unit. c) Release the lever. The OTS locks in the next 30 degree position.
GOPXR646142 | ROTATE_THE_COLLIMATOR | Rotate | Hardware | OPTIMA XR646 | How to Rotate the Multi-Leaf Collimator? | 1. Move the locking lever on the multi-leaf collimator toward the front panel, i.e. toward you. This releases the collimator from the 0 lock-in position. 2. Grasp the multi-leaf collimator with both hands. 3. Rotate it to the desired angle and direction. 4. Return the locking lever to its original position. This prevents further rotation.
GOPXR646143 | ADD_PATIENT | Add | Interface | OPTIMA XR646 | How to Add Patient? | 1. On the Worklist screen click [ADD PATIENT]. 2. Enter the patient information. 3. Click [SAVE] or [START EXAM].
GOPXR646144 | EDIT_PATIENT_INFORMATION | Edit | Interface | OPTIMA XR646 | How to Edit Patient Information? | 1. Select the procedure from the Worklist. 2. Click [PATIENT INFORMATION]. 3. Edit the information as necessary. 4. Click [SAVE] to record the changes and return to the Worklist.
GOPXR646145 | GRID_LOADING_AND_REMOVAL | Removal | Assemble | OPTIMA XR646 | How to Grid Loading and Removal? | 1. Pull out the grid tray. 2. Install the grid with the “tube side” label facing up. . 3. Push in the grid tray.
GOPXR646146 | SEARCH | Feature | Interface | OPTIMA XR646 | How to search the features? | 1. Click the button on the Search By drop-down list to select the column you want to search. If the column you want is already selected, begin at step 3. 2. Select the column. For example, Patient Name. The list is sorted by the selected column in alphabetical or numerical order. 3. Type the search criteria into the text box.
GOPXR646147 | SORT_BY_COLUMN | Sort,column | Interface | OPTIMA XR646 | How to Sort by column? | 1. Click the column heading you want to sort, or choose the column in the Search By drop-down list. For example, if you want to see all the procedures that have a status of “Suspended”, click the “Status” column heading.  An arrow appears in the column heading to indicate which column is currently being sorted. 2. Click the column heading again to switch between ascending and descending order.  An up-pointing arrow indicates that the column is sorted in ascending order. That is, sorted in alphabetical order or numerical order from smallest to largest.  A down-pointing arrow indicates that the column is sorted in descending order. That is, sorted in reverse alphabetical order or numerical order from largest to smallest.
GOPXR646148 | FILTER_LIST | Filter | Interface | OPTIMA XR646 | How to filter the Worklist? | 1. Click [FILTER LIST] on the Worklist screen.  The Filter Screen appears 2. Select the Show list for option. 3. Select the Include options. 4. Select the Date option. 5. If you selected Range for the Date option, enter or select the From and To dates.
GOPXR646149 | SELECT_A_SINGLE_PROCEDURE | Select | Interface | OPTIMA XR646 | How to Select a Single Procedure? | 1. Close or suspend any open exams The Worklist screen appears. 2. Select the procedure from the Worklist.
GOPXR646150 | SELECT_MULTIPLE_PROCEDURES | Select | Interface | OPTIMA XR646 | How to Select Multiple Procedures? | 1. Sort the Worklist by Patient Name, if necessary. 2. Click on all the procedures that are to be selected. 3. Click on [START EXAM] 4. Select matching protocol from list in drop down menu. 5. Continue to select protocols to match the protocol in the drop down list. 6. Click on [ACCEPT] when all exams have had a protocol assigned.
GOPXR646151 | DELETE_A_SINGLE_PROCEDURES | Delete | Interface | OPTIMA XR646 | How to Delete a Single Procedure? | 1. Select the procedure to delete. 2. Click [DELETE]. If necessary, switch the button to [DELETE].  A message appears: “Are you sure that you would like to delete the selected items?” 3. Click [OK] to delete the procedures.  [CANCEL] retains the procedures and the Worklist appears.
GOPXR646152 | DELETE_SUSPENDED_PROCEDURES | Delete | Interface | OPTIMA XR646 | How to Delete Suspended Procedures? | 1. Select the suspended procedure to delete.  A message appears: “The patient entry you are trying to delete is still in progress. Would you like to mark the patient as ‘Completed’ / ‘Discontinued’ and proceed with deletion?” 2. Click [COMPLETE] or [DISCONTINUE].  [COMPLETE] changes the Status to of the procedures to “Completed.” If enabled, any acquired images are auto pushed, auto printed, and sent to PACS.  [DISCONTINUE] changes the Status to “Discontinued.” Any acquired images are marked as Discontinued and the information is sent to PACS.  [CANCEL] closes the message and returns you to the Worklist without deleting procedures.  The procedures are removed from the Worklist.
GOPXR646153 | CONDUCT_A_WALL_STAND_EXAM | Exam | Exam | OPTIMA XR646 | How to Conduct a Wallstand Exam? | 1. Select the Exam and View to perform from the protocol list. 2. Select the Patient Size. The system default is Medium Adult. 3. Select the wallstand Receptor, if necessary. 4. Select [AEC] or [FIXED] mode (if applicable for the protocol). 5. Confirm or adjust the Grid and SID status.  An alert icon appears if the current Grid or SID is not in the recommended status for the technique.  A checkmark appears if the current Grid or SID is in the recommended status for the technique.  Figure 10-13 shows the Grid not in the recommended status and the SID in the recommended status. 6. Make technique adjustments as necessary: kV, mA, Focal Spot, Cu Filter, and Ion Chambers (AEC mode only). 7. Position the patient in front of the wallstand. 8. Confirm or adjust the Patient Side field, if applicable. 9. Confirm or adjust the Patient Position field. 10. Confirm or adjust the Asymmetric Collimation, if applicable. 11. Collimate and shield as appropriate for the exam.
GOPXR646154 | CONDUCT_A_TABLE_TOP_EXAM | Exam | Exam | OPTIMA XR646 | How to Conduct a Table Top Exam? | 1. If multiple procedures were selected from the Worklist, select the procedure from the exam drop￾down list. 2. Select the Exam and View to perform from the protocol list. 3. Select the Patient Size. The system default is Medium Adult. 4. Select the Detector Icon [Table Top Mode]. 5. Make other technique adjustments as necessary. 6. Position the detector under the patient. 7. Confirm or adjust the Patient Side, if applicable. 8. Collimate and shield as appropriate for the exam. 9. Have the patient suspend respiration, if required. 10. Make exposure using the hand switch.
GOPXR646155 | CONDUCT_A_CASSETTE_EXAM | Exam | Exam | OPTIMA XR646 | How to Conduct a Cassette Exam? | 1. Click [CASSETTE EXAM] from the bottom of the Worklist screen. 2. Select the protocols to perform. 3. Click [ACCEPT]. 4. Select the Exam and View to perform from the protocol list. 5. Make technique adjustments as necessary for the appropriate body part being imaged: kV and mAs 6. Change the Cu Filter selection, if necessary. 7. Position the patient with the cassette as appropriate for the exam. 8. Collimate and shield as appropriate for the exam. 9. Have the patient suspend respiration, if required. 10. Make exposure using the hand-switch. 11. Click [CLOSE] to end the exam. 12. Process the cassette as necessary, depending on the media.
GOPXR646156 | CONDUCT_A_EMERGENCY_EXAM | Emergency,Exam | Exam | OPTIMA XR646 | How to Conduct an Emergency Exam? | 1. Click [EMERGENCY EXAM] from the bottom left of the Worklist. 2. Select the protocols for the exam. 3. Click [ACCEPT].  The Acquisition screen appears. 4. Select the protocol to perform from the Protocol List. 5. Acquire images.
GOPXR646157 | ACQUIRE_AEC_IMAGES | Images | Exam | OPTIMA XR646 | How to Acquire AEC Images? | 1. Open an exam from the Worklist and select the protocols.  The Acquisition screen appears. 2. Select the protocol to perform. 3. Change the Patient Size, if necessary. 4. Change the Receptor, if necessary. 5. Click [AEC], if necessary.  Depending on the current protocol, AEC may already be selected by default.  The AEC mode displays the Ion Chamber selections.  The Fixed mode removes the Ion Chamber selections.  You are able to alter this selection at any time to accommodate the requirements of a specific examination. 6. Change the selected Ion Chambers, if necessary.  Selected (active) AEC cells will appear highlighted on the acquisition screen. 7. Make other technique adjustments as necessary.  You are able to adjust the kV, mA, and Focal Spot. 8. Confirm or select the Cu Filter, if necessary. 9. Position the patient so that the anatomy of interest is centered over the selected AEC cells. Note: The body part must cover the selected ion chambers in order to achieve the appropriate expo￾sure. 10. Collimate the light field to encompass the detection area being used. 11. Collimate to the desired area. This ensures the receptor sensing area is aligned with the patient’s anatomy. 12. Confirm the Patient Position. Change if necessary. 13. Make the exposure. 14. If the AEC Back-up Reached message appears: Click [OK].  The message closes. 15. Observe technique settings and continue exam.
GOPXR646158 | CUSTOMIZE_SYSTEM_ANNOTATIONS | Customize | Customize | OPTIMA XR646 | How to Customize System Annotations? | 1. Select the Image Tools panel – Annotation tab if necessary. 2. Click [CUSTOM]. 3. Select (check) the annotations you want to appear. 4. Adjust the Font Size, if necessary. 5. Click [OK].  [CANCEL] closes the screen and leaves the selections unchanged.
GOPXR646159 | ADD_IMAGE_ANNOTATIONS | Image | Customize | OPTIMA XR646 | How to Add Image Annotations? | 1. Select the image to annotate, if necessary. 2. Click the button of the annotation to insert. Note: Selected image annotations are yellow with red handles. Unselected image annotations are aqua without handles. 3. Move, re-size, or change the angle of the annotation as described
GOPXR646160 | DELETE_IMAGE_ANNOTATIONS | Image | Customize | OPTIMA XR646 | How to Delete Image Annotations? | 1. Select the annotation. 2. Click [ERASE] or [ERASE ALL].  If [ERASE ALL] was clicked, a message appears: “Would you like to remove all annotations from the selected image?” 3. Click [YES].  All annotations are removed.
GACTS161 | RESET_THE_EMERGENCY_OFF_BUTTON | Emergency,Care,Disconnect | Safety | REVOLUTION ACTS | How to Reset Emergency OFF button? | 1. Press the Start button on the Main Disconnect Control.  Powerto the Power Distribution Unit (PDU), operator console, and system electronics will be restored. 2. Stab the E-Reset hole on the Control Panel or SCBSV.  Power to the gantry drives, X-ray system, and table drive will be restored. |  |  | ES7
GACTS162 | EMERGENCY_PATIENT_CARE_DURING_X-RAY_ON | Emergency,Care, | Safety | REVOLUTION ACTS | How to Care patient when x-ray is ON? |  Press STOP SCAN to abort X-ray and stop gantry/table movement.  Duringanexam,thesystempausesbetweenscans if youPressanybuttononthecontrol panelother than the alignment lights. It stops X-ray if you press the same buttons during a scan.  Select Resume on the screen to continue the exam. |  |  | ES7
GACTS163 | EMERGENCY_EGRESS | Emergency,Care | Safety | REVOLUTION ACTS | What will be do when the System operation maybe stopped due to powerfailure? | Use Cradle Unlatch button. |  |  | ES7
GACTS164 | TO_SAFELY_REMOVE_THE_PATIENT | Emergency,Care | Safety | REVOLUTION ACTS | How do we safely remove the patient? | 1. Press the Cradle Release gantry key or the Emergency Stop button (1) to disengage the clutch. 2. Pull the cradle to its out position, using the Cradle Lip or Cradle Handle (2). 3. Assist the patient off the table. |  |  | ES7
GACTS165 | SETTING_DOSE_ADMINISTRATOR_ROLE | Dose,Alert values | Dose Check | REVOLUTION ACTS | How to set Dose Administrator Role? | 1. ClickDoseCheck Managementfrom Protocol Management. 2. Click User Admin Tool. 3. From the EA3 Administration screen, type your Username and Password.  Consult your service engineer or Application Representative for user name and password. 4. From the Local Users tab, click Add Local User. 5. From the pop-up panel, type information for each of the following:  A unique userID  FullName  Password  ConfirmPassword 6. Click Add User. 7. From the Groups tab, click Add Local Group. 8. From the Add Local Group window, type and enter a unique group name 9. Click AddGroup  The group is highlighted in the Local Groups list box.  All information and buttons in the center panel referto the highlighted group. 10. To change a group's roles, select the Dose Admin Role option box and click Apply Roles. a. A green label confirms the applied roles. b. An error message box displays if it is unsuccessful. | https://drive.google.com/file/d/1zdcWJjNz4L471S1NTej2I_xcxSNqFKLl/view?usp=sharing |  | RS8
GACTS166 | SETTING_AV_EXCEEDING_USER_ROLE | Dose,Alert values,AV | Dose Check | REVOLUTION ACTS | How to set AV Exceeding User Role? | 1. ClickDoseCheck Managementfrom Protocol Management. 2. Click User Admin Tool. 3. From the EA3 Administration screen, type your Username and Password. 4. From the Local Users tab, click Add Local User. 5. From the pop-up panel, type information for each of the following:  A unique userID  FullName  Password  ConfirmPassword 6. Click Add User. 7. From the Groups tab, click Add Local Group. 8. From the Add Local Group window, type and enter a unique group name. 9. Click Add Group.  The group is highlighted in the Local Groups list box.  All information and buttons in the center panel referto the highlighted group. 10. To change agroup's roles, selecttheDoseCheck AV ExceedingUser option box and click ApplyRoles.  A green label confirms the applied roles.  An error message box displays if it is unsuccessful. | https://drive.google.com/file/d/1fbtu6QSgw_uH_xFViMX_F2F2Gu-atPKr/view?usp=sharing |  | RS8
GACTS167 | CONFIGURE_THE_SYSTEM_FOR_DOSE_CHECKING | Dose,Alert values,AV,NV | Dose Check | REVOLUTION ACTS | How do we check the system for Dose checking? | 1. ClickDoseCheck Managementfrom Protocol Management. 2. Set Checking for Notification Value (NV). a. In theNV Checking section Click CTDIvol and/orDLP. b. Type a value for CTDIvol and/or DLP.  The range for NV is 0 to ≤ to the value set for AV.  Or, the letter N can be entered to indicate that no checking will be done. The system default is NV checking CTDIvol "On" andDLP "Off". GE reference protocols will not contain anyNV values. 3. Set checking for Alert Value (AV). a. In the AV Checking section Click CTDIvol and/orDLP. b. Type a value for CTDIvol and/or DLP.  The maximum AV value for CTDIvol is 2000 mGy.  The maximum AV value for DLP is 400,000 mGy-cm.  ThesystemdefaultisAVcheckingbyCTDIvol"On"withavalueof1000mGy.DLPchecking"Off" with a value of 0 mGy-cm. Alert Value should be determined by a collaboration between the Radiologist and Physicist at the site. 4. Click Save. | https://drive.google.com/file/d/16mCgT_zI3VwJBElE7N1pp_8G6Ld6nm35/view?usp=sharing |  | RS8
GACTS168 | CONFIGURE_THE_SYSTEM_FOR_ALERT_VALUE_BY_AGE_THERSHOLD | Dose,Alert values,AV,NV,Age | Dose Check | REVOLUTION ACTS | What are the procedure to configure the system for Dose Checking by Alert Value(AV) Checking byAgeThreshold? | 1. Click Dose Check Management from Protocol Management. 2. In the AV (Alert Value) Checking section, click Add Age Threshold. 3. Type the desired Patient Age. 4. Type a value for CTDIvol and/or DLP.  The maximum AV value for CTDIvol is 2000 mGy.  The maximum AV value for DLP is 400,000 mGy-cm.  ThesystemdefaultisAV checkingbyCTDIvolOnwithavalueof1000 mGy.DLPcheckingOff,0 mGy  The value must be lower than the Adult AV value. 5. Click save. | https://drive.google.com/file/d/1l0zh2AKya0_xOdJKpoEMrFsGDrM21Xml/view?usp=sharing |  | RS8
GACTS169 | CONFIGURE_THE_SYSTEM_FOR_PROTOCOL_CHANGE_CONTROL | Dose,Alert values,AV,NV,Protocol | Dose Check | REVOLUTION ACTS | What are the procedure to configure the system for Protocol Change Control (PCC)? | 1. Click Dose Check Management from Protocol Management. 2. In the Protocol Change Control section Click Off to enable PCC. 3. Click save. | https://drive.google.com/file/d/1e3kOSqwCjOELrYR60ea7RHCdnAjzDyHM/view?usp=sharing |  | RS8
GACTS170 | BUILDING_PROTOCOLS_USING_PROTOCOL_CHANGE_CONTROL | Dose,Alert values,AV,NV,Protocol | Dose Check | REVOLUTION ACTS | What are the procedure to build protocols with Protocol Change Control (PCC)? | 1. ClickDoseCheck Managementfrom Protocol Management. 2. Click Logon Name. 3. Type the user name. 4. Click Password. 5. Type the User password. 6. Click Logon. | https://drive.google.com/file/d/1QV278wXB8wsSEA4OWzQs8M0_yW4XjWOw/view?usp=sharing |  | RS8
GACTS171 | BUILDING_PROTOCOLS_WITH_NOTIFICATION_VALUES | Dose,,AV,NV,Protocol | Dose Check | REVOLUTION ACTS | What are the procedure to set Notification Values (NV) in User protocols? | 1. From the scan monitor, click the Protocol Managementicon. 2. On the Auto Voice and Protocol Management screen, click Protocol Management. 3. On the Protocol Selection screen, click User. 4. On the User tab click an Anatomical area. 5. In the Anatomical Area click on a protocol. 6. Click Next Series to go to the first non-scout series in the protocol. 7. From the Dose Information area click Set Up. 8. Type the desired CTDIvol or DLP NV value by group.  Only one CTDIvol or DLP value can be set for SmartPrep scan groups. 9. Click Info to return to the Dose Information display. 10. Click Accept to save the protocol parameters.  The system compares all the CTDIvol and DLP values in the scan groups to the entered NV value.  The protocol parameters will not be accepted until the NV value is no longer exceeded or the user has entered a Diagnostic Reason for continuing to accept the scan parameters exceeding the NV Value.  Saving a protocol with a dose level exceeding an NV value demands a high degree of consideration for appropriateness. Therefore, in most cases, you will click Cancel to go back to adjust parameters. However, if there is a true clinical reason to save a protocol exceeding an NV value, then enter a diagnostic reason. 11. Click in theDiagnostic Reason text field. 12. Type the reason for scanning exceeding the NV value.  The maximum number of characters is 64.  Once the maximum number of characters is reached the system will not accept any additional character input. 13. Click Confirm to accept the protocol scan parameters. | https://drive.google.com/file/d/1fEVH502cV6rDMOIKx_Idrn6JkVhvw-gA/view?usp=sharing |  | RS8
GACTS172 | BUILDING_PROTOCOLS_WITH_NOTIFICATION_VALUES_BY_FACTOR | Dose,,AV,NV,Protocol,Factor | Dose Check | REVOLUTION ACTS | What are the procedure to set Notification Values (NV) in User protocols for all scan groups by a factor? | 1. From the scan monitor, click the Protocol Managementicon. 2. On the Auto Voice and Protocol Management screen, click Protocol Management. 3. On the Protocol Selection screen, click User. 4. On the User tab click an Anatomical area. 5. In the Anatomical Area click on a protocol. 6. Click Next Series to go to the first non-scout series in the protocol. 7. From the Dose Information area click Set Up. 8. Click NV column edit. 9. Type the desired factor. 10. ClickAccept. 11. Click Info to return to the Dose Information display. | https://drive.google.com/file/d/179FvWQIxSNaKdM3k3WiRgPuaspnBlfem/view?usp=sharing |  | RS8
GACTS173 | SCAN_USING_ALERT_VALUE_CHECKING | Dose,,AV,NV,Protocol,Checking | Dose Check | REVOLUTION ACTS | What are the procedure to scan using AV checking? | 1. Click in the Logon Name text field. 2. Type the UserName. 3. Click in the Password text field. 4. Type the user password. 5. Click in the Diagnostic Reason text field. 6. Type the reason for scanning exceeding the AV value.  The maximum number of characters is 64.  Once the maximum number of characters is reached the system will not accept any additional characterinput. 7. Click Confirm to proceed to scanning. | https://drive.google.com/file/d/1ehBPKyw-VRMdS3ptI7wmJ9ksGaYSTXHh/view?usp=sharing |  | RS8
GACTS174 | SCAN_USING_NOTIFICATION_VALUE_CHECKING | Dose,,AV,NV,Protocol,Checking | Dose Check | REVOLUTION ACTS | What are the procedure to scan using NV checking? | 1. Click in the Diagnostic Reason text field. 2. Type the reason for scanning exceeding the NV value.  The maximum number of characters is 64.  Once the maximum number of characters is reached the system will not accept any additional characterinput. 3. Click Confirm to proceed to scanning. | https://drive.google.com/file/d/1UGVpXfazVFsUeNgEigxyMQu72rHoBfJk/view?usp=sharing |  | RS8
GACTS175 | DISPLAY_A_DOSE_CHECK_LOG | Dose,,AV,NV,Protocol,Check log | Dose Check | REVOLUTION ACTS | What are the procedure to display a Dose Check Log? | 1. On the Auto Voice, Protocol Management and Dose Check Management screen, click Dose Check Management. 2. On the Dose Check Management screen, click Log List. 3. From the Dose Audit Tool, click Audit Tool. 4. On the Log List click on the desired log.  If the log data approaches the maximum size, the system will post a message directing the user to export the log data. |  |  | RS8
GACTS176 | DISPLAY_THE_PROTOCOL_SUMMARY | Dose,,AV,NV,Protocol,Summary | Dose Check | REVOLUTION ACTS | What are the procedure to display the Protocol Summary? | 1. ClickDoseCheck Managementfrom Protocol Management. 2. On the Dose Check Management screen, click Audit Tool. 3. In the Protocol Summary area, click All Protocols orInvalid Protocols. | https://drive.google.com/file/d/1-nardEeL4ma2sDrDgQbW-RwmiQ81_8bU/view?usp=sharing |  | RS8
GACTS177 | EXPORT_DOSE_CHECK_LOG_DATA | Dose,,AV,NV,Protocol,Log data | Dose Check | REVOLUTION ACTS | What are the procedure to export Dose Check log data? | 1. ClickDoseCheck Management from Protocol Management. 2. Insert CD –R media in the DVD RW drive or insert USB media in the USB port of the media tower. 3. On the Dose Check Management screen, click Audit Tool. 4. On the Dose Audit Tool, click Export. 5. On the Export Media screen, click media type CD/DVD or USB. 6. Onthe Export Contents, click 1 or moreof the following:DoseCheck Log, Protocol Summary or User Protocols  If Protocol Summary is selected click All Protocols or Invalid Protocols. 7. On the Export screen, click OK. |  |  | RS8
GACTS178 | COLOUR_CODING_FOR_KIDS_PROTOCOL_SELECTION | Colour,Kids,Protocol | Pediatrices | REVOLUTION ACTS |  |  | https://drive.google.com/file/d/1qVYLzrPjoyOodpymh4fdTc88jPl2I5bF/view?usp=sharing |  | PS9
GACTS179 | SETUP_A_PEDIATRIC_PATIENT_EXAM | Setup,Exam,pediatric,Small patient | Pediatrices | REVOLUTION ACTS | What are the procedures in time you start a new patient exam for a pediatric patient? | 1. On the scan monitor, click the Exam Rx icon.  The PATIENT INFORMATION screen displays the newExam Number.  The maximum Exam Number is 49, 999, which is reset by your Field Engineer. 2. From the PATIENT INFORMATION area, type data into the appropriate fields.  Press Enterto advance to the next field. Alternatively, use the mouse to navigate to each field.  Patient ID is a required field. If the patient does not have an identification number, type ? orthe wordtrauma. 3. Click PEDIATRIC. 4. From the Pediatric Protocol Selection screen, click an anatomical area.  With the patient's weight entered,the system automatically selects the appropriate pediatric color code area forthe anatomical area selected.  Use the arrows to scroll through the list.  Click a protocol from the list to download the scan parameter values. 5. Fromthe Pediatric Protocol Category screen, selectthe weight/length category based on the size of your patient or verify the correct Color category has been selected if you entered a weight.  The protocols in the selected color code are then displayed accordingly.  Thedefaultweight/colorselectordisplaysthepatientweightenteredonthePATIENTINFORMATION screen,orthe last weight/color selection(if no patientinformation was entered).  There is no reference protocolfor Pediatric areasNeck, Upper Extremity,Chest, Abdomen, Spine, Pelvis, and Lower Extremity.  If you enter a patient weight on the Patient Information screen and select a color/weight that is not consistent with the entered information, an error message displays. You must acknowledge that you have chosen a protocol that does NOT match the patient size. 6. Proceed with the Acquire a Scout scan procedure. |  |  | PS9
GACTS180 | ADJUST_PEDIATRIC_PATIENT_SCAN_PARAMETERS | Setup,Exam,pediatric,Small patient | Pediatrices | REVOLUTION ACTS | What are the parameters to be Adjusted in Pediatric patient Exam? | size, age, weight, height,also Recommended pediatric color coded protocols. |  |  | PS9
GACTS181 | AUTOmA | Setup,Exam,pediatric,Small patient | Pediatrices | REVOLUTION ACTS | What are the Procedures When Center the patient properly by using AutomA? |  Double check and verify the table height is centered to the patient.  Raise or lower the table as needed before taking the scouts.  Afterthescoutsaretakenandprescriptionisdone,verifythemAtablecalculationsbeforeconfirming the scan.  Make sure minimum mA and maximum mA values are set appropriately.  See Set the mA procedure in the Scan chapter. |  |  | PS9
GACTS182 | SFOV_FILTERS | Filter,Exam,pediatric,Small patient | Pediatrices | REVOLUTION ACTS | For small Patients what are the conditions to be maintained for SFOV Filters? |  Matching the appropriate SFOV bowtie filter to the size of the patient ensures the dose is delivered where it is needed, and filtered where it is not needed.  Small SFOV supports DFOVs up to 25 cm in diameter.  Ped Head SFOV supports DFOVs up to 25 cm in diameter and uses IBO processing to correct for beam hardening effects. It is particularly useful for infants 18 months or less in age. |  |  | PS9
GACTS183 | SCBSV_BUTTONS | Buttons,Equipment | Equipment | REVOLUTION ACTS |  |  | https://drive.google.com/file/d/1FoV0WFIiVl9SPJwOblCCxwul6hnvYYGU/view?usp=sharing
GACTS184 | OPERATE_THE_MOUSE_CONTROLS | Mouse,Controls | Equipment | REVOLUTION ACTS |  |  | https://drive.google.com/file/d/1ID4y1Mj1hlky89LHAstVN6DUxoAFjLK7/view?usp=sharing
GACTS185 | OPERATE_THE_GANTRY_CONTROLS | Gantry,controls | Equipment | REVOLUTION ACTS |  |  | https://drive.google.com/file/d/1HyRABfbcAifC2smA6vLjicpNqJM_US4Z/view?usp=sharing
GACTS186 | SYSTEM_LOGIN | System,login | Power | REVOLUTION ACTS | How to login to the system? | 1. From the logon screen, type your assigned logon name. 2. Typeyourpassword.  Click Emergency Logon only if you do not have a valid account set up. The Emergency Logon screen is launched, which does not require a ID or password. 3. Click Logon. |  |  | ES1
GACTS187 | SYSTEM_LOGOUT | System,logout | Power | REVOLUTION ACTS | How to log off the system? | 1. From the monitor, click the Shutdown icon. 2. From the Attention Shutdown pop-up screen, click Logout User andOK. |  |  | ES1
GACTS188 | ROUTINE_DAILY_START_UP | System,Start up | Power | REVOLUTION ACTS | Daily start up? | 1. From the monitor, click Shutdown. 2. ClickRestart. 3. ClickOK. |  |  | ES1
GACTS189 | STARTUP_FROM_POWER_OFF | System,Start up | Power | REVOLUTION ACTS | How to start up from power off? |  If powerwas turned off via the computer switch andtheMainDisconnectControlA1 PowerPanel (Breaker),then a. Turn on power to the Main Disconnect Control. b. Turn on power to the computer. c. Press Reset on the gantry controls.  SeeTubewarm-upprocedureforpotentialAttentionmessages.ClickOKandfollowtheAttention instructions.
GACTS190 | SHUTDOWN_AND_RESTART_THE_SYSTEM | System,Shutdown,Restart | Power | REVOLUTION ACTS | What are the procedures to Shutdown and restart the system? | 1. Disconnect any USB devices from the USB drive before performing a reboot or shutdown.  If a save/restore of scan data is in progress, wait for the process to complete, exit Recon Management, and disconnect the USB device before you reboot or shutdown. If for some reason the system should shutdown unexpectedly while a USB drive is connected to the system, disconnectthe USB drive before restarting the system. 2. From the monitor, click Service 3. FromService,click Shutdown. 4. From the Attention Shutdown pop-up screen, click Shutdown or Restart and OK. 5. For systemshutdown,whenthe"SystemHalted" message displaysonthe monitorthesystemcan be powered off. Press the power switch on the front of the operator’s console to the off position. |  |  | ES1
GACTS191 | PATIENT_SCHEDULE | Patient,Schedule | Patient Schedule | REVOLUTION ACTS | How to view the Patient Schedule Screen? | Click the Patient Schedule icon to display the Schedule screen. | https://drive.google.com/file/d/18hjlplmjn7ubIjM-keWa7uX7CkAoHHUI/view?usp=sharing |  | DS4
GACTS192 | UPDATE_PARAMETERS_SCREEN | Parameter,Patient | Patient Schedule | REVOLUTION ACTS | How to update your Patient Schedule list with information from HIS/RIS? | Click the Patient Schedule icon and selectUpdate Parameters to display the Update Parameters screen, where you can update your Patient Schedule list with information from HIS/RIS. Get Patient List For  This system populates the Patient Schedule for the current scanner you are on.  All CT Systems poppulates the Patient Schedule for all of the CT systems on the HIS/RIS connection.  All Systems populates the Patient Schedule for all the systems on the HIS/RIS connection. DataRange Populates teh Patient Schedule from the defined date range. Use the month/day/year format. Requested Proc ID, AccessionNumber, PatientName, PatientID Additional ways to populate the Patient Schedule. ContinueUpdate Applies the criteria chosen and updates the Patient Schedule. CancelUpdate Closes the Update Parameters screen without saving any new selections. | https://drive.google.com/file/d/1x4OWh7dQKT16fmjzXNw7eFPCoJYcKCVB/view?usp=sharing |  | DS4
GACTS193 | PATIENT_SCHEDULE_PREFERENCES_SCREEN | Preferences,Patient | Patient Schedule | REVOLUTION ACTS | How to select preferences in patient Schedule Screen? | Click the Patient Schedule icon andselectPreferences. | https://drive.google.com/file/d/1p1JFiG6DbtMkL114eq-7EELlszwYveva/view?usp=sharing |  | DS4
GACTS194 | ADD_A_PATIENT_TO_THE_SCHEDULE_MANUALLY | Manually,Patient,Schedule | Patient Schedule | REVOLUTION ACTS | How to Add a patient to the schedule manually? | 1. Click the Patient Schedule icon. 2. From the Schedule screen, click Add Patient. 3. From the Schedule Patient screen, complete all relevantfields.  Entering even weights may be rounded up to the next odd pound because weights are stored as kilogram units then converted back to pounds.  If adding a patient where Exam Split is going to be used, an exam description must be entered for those entries to be used with Hard Exam Split.  If desired, add the number of the specific protocol under Protocol Number.  If desired,type in the date and time in appropriate fields to be able to sort the Patient Schedule by scheduledtimes. 4. Click Accept to add the patient to the schedule. |  |  | DS4
GACTS195 | EDIT_A_PATIENT_IN_THE_SCHEDULE | Edit,Patient,Schedule | Patient Schedule | REVOLUTION ACTS | How to Edit a patient in the schedule? | 1. Click the Patient Schedule icon. 2. From the Schedule screen, select a patient. 3. Click Edit Patient. 4. Change the patientinformation as needed. 5. ClickAccept. |  |  | DS4
GACTS196 | DELETE_A_PATIENT_FROM_THE_SCHEDULE | Delete,Patient,Schedule | Patient Schedule | REVOLUTION ACTS | How to Delete a patient from the schedule? | 1. Click the Patient Schedule icon. 2. From the Schedule screen, select patients to be removed. 3. ClickDelete Selected.  Alternatively, choose one of the following options.  Click Delete All to remove all the patients from the list.  Click Delete All Completed to remove all the completed patients marked with "C" from the list. 4. Click OK to the confirmation prompt. |  |  | DS4
GACTS197 | UPDATE_THE_PATIENT_SCHEDULE | Update,Patient,Schedule | Patient Schedule | REVOLUTION ACTS | How to Update the patient schedule? | 1. Click the Patient Schedule icon. 2. From the Schedule screen, click Update. 3. From the Update Parameters screen, select one of the following options from the Get Patient List to populate the schedule from yourHIS/RIS.  Select This system to pull the patient schedule forthe current scanner.  Select All CT Systems to pull the patient schedule for all of the CT systems on the HIS/ RIS connection.  Select All Systems to pull the patient schedule for all the systems on the HIS/RIS connection. 4. Enter values for the Date Range. 5. Enter a Requested Procedure ID, accession number, patient name orID to narrow the search. These are optional fields. 6. Click Continue Update to continue or click Cancel Update to exit and make no changes. |  |  | DS4
GACTS198 | VIEW_MORE_PATIENT_INFORMATION_ON_THE_SCHEDULE | Information,Patient,Schedule | Patient Schedule | REVOLUTION ACTS | How to View more patient information on the schedule? | 1. Click the Patient Schedule icon. 2. From the Schedule screen, select a patient. 3. Click View More Info. The screen is populated from your HIS/RIS. 4. Click Cancel to close the More Information screen. |  |  | DS4
GACTS199 | SET_PATIENT_SCHEDULE_PREFERENCES | Information,Patient,Schedule | Patient Schedule | REVOLUTION ACTS | How to Set patient schedule preferences? | 1. Click the Patient Schedule icon. 2. From the Schedule screen, click Preferences. 3. From the Preferences screen, complete fields as needed. |  |  | DS4
GACTS200 | SELECT_A_PATIENT_FROM_THE_SCHEDULE | Information,Patient,Schedule | Patient Schedule | REVOLUTION ACTS | How to Select a patient from the schedule? | 1. Click the Exam Rx icon. 2. Click the Patient Schedule icon. 3. From the Patient schedule list, select the desired patient and click Select Patient.  The Patient Information screen fields are populated.  Multiple records forthe same patient can be selected.  Multiple records must be selected in order to use Exam Split.  If more than one record with the same Patient ID is found in the Patient schedule list a dialog message displays. To avoid having multiple Patient Records with the same Patient ID, set the Delete Completed Exams preference to zero. This will ensure that only new records are in the Patient Schedule list. 4. If a protocol is not tied to the patient, select the protocol. The system displays the first series. |  |  | DS4
GACTS201 | CHECK_A_PATIENTS_STATUS | Information,Patient,Schedule,Status | Patient Schedule | REVOLUTION ACTS | How to Check a patient's status? | 1. Click the Patient Schedule icon. 2. View the Status column.  N = New Record or Not Completed  C = Completed 3. Click or arrows to page through the list. |  |  | DS4
GACTS202 | IMAGES_FAIL_TO_RECONSTRUCT | Images,Scan,Reconstruct | Scan | REVOLUTION ACTS | What are the steps if images fail to reconstruct? | 1.Click Recon Status to display Recon Management screen. 2.From the Recon Management screen, click Unsuspend Entries. 3.If images still fails to recon click Restart Queue. 4.If images still fails to recon, shutdown and reboot the system.
GACTS203 | PATIENT_INFORMATION_SCREEN | Scan,Patient,Scanning | Scan | REVOLUTION ACTS | How to check the Patient lnformation scanning screen? | Click Exam Rx , to check the scanning screenn
GACTS204 | PATIENT_INFORMATION_PRESET_SELECTION_SCREEN | Scan,Patient,Scanning,Preset | Scan | REVOLUTION ACTS | How to set Patient Information preset selection screens? | Exam Description Referring Physician Operator Radiologist Information in the Referring Physician List is listed in the order it is entered with last entry displayed at the top of the list. Change Allows you to change an item on the list. Add Allows you to add a name to the list. Delete Allows you to remove an item from the list. Sort Allows you to sort the Exam Description list. Text field Displays the item that is to be changed, added, or deleted.
GACTS205 | PATIENT_POSITION_OR_SERIES_LEVEL_FUNCTIONAL_AREAS | Scan,Patient,ScanninG,Areas | Scan | REVOLUTION ACTS | How to setup Patient Position/Series Level Functions area? | The Patient Position/Series Level Functions area is on the Scan Setup screen. Auto Store/Auto Transfer If these features were set for the scout image, they are automatically set for this series. If they were not selected for the scout series, they may be activated in this series and the scout images are included in the storage/transfer. Click Auto Store to automatically send the image data to the system storage device. Click Auto Transfer and a window displays where you can select the IP address to send the images. Dose Report Auto Transfer Auto Transfer for Dose Report can be set individually. Click Dose Report Auto Transfer to select a remote host to automatically send the reports. Auto Scan Automatically moves the table and scans the confirmed series. If Auto Scan is disabled, you must press Move to scan for every scan before Start Scan becomesavailable. If you select Auto Scan for a group prescription, it remains on for every group in that series. Auto Scan is automatically enabled for a SmartPrep series.
GACTS206 | PATIENT_POSITION_IN_CORRECT_MANNER | Scan,Patient,Scanning,Position | Scan | REVOLUTION ACTS | How to set the correct patient position? | click on the arrow direction on top of the patient's body to change the head forward / foot forward direction, click on the arrow direction on top of the patient's body to rotate the model, for a rotation of 90°.
GACTS207 | COPY_PATIENT_POSITON_AND_ORIENTATION | Scan,Patient,Scanning,Orientation | Scan | REVOLUTION ACTS | How to copy the Patient Position and Patient Orientation? | click Copy Pt.Orient Pt.Position Anat.Ref. When the protocol is used in New Patient, the fields where Copy Forward have been defined will be outlined in green. If copy forward for Patient Position/Patient Orientation is selected, Copy Pt. Orient. Pt. Position Anat. Ref. will be highlighted.
GACTS208 | CREATE_NEW_SERIES | Scan,Patient,Scanning,Series | Scan | REVOLUTION ACTS | How to Create New Series? | Opens the Select New Series Type screen, where you can create a new series in the exam.
GACTS209 | SCAN_TASK_LIST | Scan,Patient,Scanning,Series | Scan | REVOLUTION ACTS | If we scan task list what will happened? | Scan tasks List shows all scan sequences included in the current examination, of which the highligted item is the current scan sequence, with symbol to indicate that it has completed the scan sequence.
GACTS210 | VIEW_EDIT_SCREEN | Scan,Patient,Scanning,Edit | Scan | REVOLUTION ACTS | Need more about view edit screen? | The View Edit screen is an area on the Scan Setup screen. The contents of the ViewEdit screen changes dependent on the active application. End Exam Completes the exam and sends any prescribed PMR1 Recon 2 and 3 for the last series to be reconstructed.It also starts AutoStore and AutoTransfer by Exam. Next Series Transitions to the next series if additional series are programmed in the protocol. Confirm Takes you out of the prescription preview mode so you can begin a scan. Confirm is not available if the patient has not been landmarked or if invalid parameters exist on the ViewEdit screen.
GACTS211 | SCAN_WORK_FLOW_OR_EXAM_WORK_FLOW | Scanning,patient setup,Image management | Scan | REVOLUTION ACTS | What is the overview procedures for Scan or Exam work flow? | Set up and position the patient. 1.Set up the patient's information. 2.Acquire a scout. 3.Define the Series setup parameters. Set the Auto Store option. a.Set the Auto Transfer option. 4.Acquire the scan data. Set the scan parameters. a.Adjust the Graphic Rx. b.Set the timing parameters. c.Set up Auto Voice. d.Set the Recon parameters. e.Set the film parameters. f.Apply any additional scan features. g.Enter contrast descriptions. h.Start the scan. Check the following before you Confirm: the start/end locations for each recon direction of scan mA table for Auto or SmartmA Repeat a series, if necessary. 5.End the exam. 6.View the images. 7.Manually film the images. 8.Manually archive the images. 9.Manually network the images. 5641787-
GACTS212 | SET_UP_AND_POSITION_THE_PATIENT_IN_SCAN | Scan,Position,setup | Scan | REVOLUTION ACTS | Before going into the Scanning room,What are the procedures to be done to the patient positioning? | 1.Connect the head holder or foot extension on the cradle facing the gantry. If the head holder cannot be securely locked onto the cradle, please contact your servicerepresentative to install the parts inside the tool bag 2327335 (P9230JV). 2.When preparing to move the patient onto the table, step on the 'home' pedal under the bed to adjustthe height of the table. The cradle moves to home and to a minimum height of 44 cm. 3.Transfer the patient to the table. Make sure to follow the manufacture directions for use of a protective table slicker whenscanning, otherwise artifacts may be seen in images. Blocked reference channels may cause image artifacts. Make sure that there are no blankets,clothing, tubing, or straps hanging down below the table when scanning that could cause ablockage. Please use the patient positioning strap to secure miscellaneous items beside the table. 4.Make the patient comfortable and immobilize as needed. Connect body straps to the cradle andwrap them around the patient. For head scans, confirm that the patient's head is completely in the head holder. If the patient hasnot entirely entered into the head holder, then light and dark images can be seen at both ends ofthe collimated rays during the two rotations. When the scanning is at the head of the bed surface, make sure that the head is not at theconnection of the cradle's extension. The head should be fully placed on the cradle. Always use the head holder straps to immobilize the patient's head during head scans. Make sure items such as blankets and straps are not left hanging off the table or head holder,allowed to touch, or be dragged through the gantry while scanning otherwise artifacts may beseen in images. 5. From the gantry controls, press the following buttons as needed to refine the patient's position.  Cradle moves out  Cradle moves in 6. On the gantry control panel, click the laser alignment light button to switch on the laser: align the laser with the center of the anatomic site.  A poorly positioned and centered patient can impact the mA values calculated for AutomA/ SmartmA.  Do not position the patient with the laser lights in their eyes. 7. On the gantry control panel, click the external icon or the internal icon. Confirm remains gray (unselectable) until the landmark button has been pressed.  The landmark sets a known anatomical reference for the radiologist to correlate anatomy.  The landmark sets the zero location. When scanning towards the patient’s head, you are scanning superior to the zero location. When scanning towards the patient’s feet, you are scanning inferior to the zero location. You should set the zero location to known anatomy. For example, when scanning a head, the landmark or zero location is typically the orbital meatal line.  A landmark or zero reference point is required before you scan a patient. It can be set before or after you click Exam RX.  When using the laser alignment light for patient positioning purposes, be aware that the patient’s elevation may be slightly lower with the cradle extended than with the cradle fully retracted. The cradle might bend slightly because of the patient's weight. This difference should be taken into consideration for applications where patient position information is critical, such as Treatment Planning. 8. From SCBSV, set patient-to-operator and operator-to-patient volume. Press Talk control and test the volume levels. AutoVoice communication control volume is adjusted on the software interface. 9. Proceed to Set up the patient's information.
GACTS213 | SETUP_PATIENT_INFORMATION | Patient,Setup,Information | Scan | REVOLUTION ACTS | Before entering the patient to the scan,What are the procedures to be done? | 1. Click Exam Rx icon or Emergency Patient . * The Patient Information screen displays the new Exam Number. * The maximum Exam Number is 49, 999, which is reset by your Field Engineer. Emergency Patient screen only: * Emergency Patient screen automatically assigns the Patient ID and Patient Name. Both fields can be modified. 2. From the Patient Information area, type data into the appropriate fields. *Press Enter to advance to the next text field. Alternatively, use the mouse to navigate to each field. *The mouse cursor must be within the Patient Information area for the input to be accepted. *Patient ID is a required field. If the patient does not have an identification number, type ? or theword trauma. *Once an identification number has been assigned to the patient, the exam information may beedited using Edit Patient after End Exam. *Set Delete Completed Exams to zero when Patient Information is updated from the HIS/RIS. Thisassures that Completed Accession numbers are not inadvertently selected for scanning a secondtime causing Patient reconciliation issues on the PACS1 . *DICOM2 users must input the patient name according to the following method: surname, givenname, middle name, use "^" to separate each field. For example: Doe^John^M. 3 When all the desired patient information is entered, select a protocol from the Anatomical Selectorarea, using one of the following methods. *Type a protocol number in the Protocol Number field. *From the Protocol Selection screen, click an anatomical area. Click a protocol from the list todownload the scan parameter values or click a reference protocol. 4.Proceed to Acquire a Scout.
GACTS214 | CHANGE_AN_ITEAM_LIST_ON_PRESET_FIELD | Patient information | Scan | REVOLUTION ACTS | How to Change an item on the list in preset field? | 1.From the Preset fields screen, select the item you want to change. 2.Type the information in the text field. 3.Click Change.
GACTS215 | ADD_AN_ITEAM_TO_THE_LIST | Patient information | Scan | REVOLUTION ACTS | How to Add an item to the list in preset field screen? | 1.From the Preset fields screen, enter new information in the text field. 2.Click Add. * 100 is the maximum number of entries. * Items are listed with the newest entry at the top of the list.
GACTS216 | DELETE_AN_ITEAM_FROM_LIST | Patient information | Scan | REVOLUTION ACTS | How to Delete an item from the list in Preset field screen? | 1.From the Preset fields screen, select the item you want to delete. * The item is displayed in the text field. 2.Click Delete.
GACTS217 |  | Patient information | Scan
GACTS218 |  | Patient information | Scan
GACTS219 |  |  | Scan
GACTS220
GACTS221 | failure in gandry
GACTS | setting,scan,run
n no of system
unlacth the table

## CH9.2
TID (unique) | KEYWORD(unique) | TAG(Separate with comma) | CATEGORY | PID(Separate with comma if more than one) | QUESTION | ANSWER | Documentation link | YouTube solution link | RSID

## Sheet2
TID (unique) | KEYWORD(unique) | TAG(Separate with comma) | CATEGORY | PID(Separate with comma if more than one) | QUESTION | ANSWER | Documentation link | YouTube solution link
GRGXR61 | OPERATOR_CONSOLE | operator console,indicators,indications,radiography,ge | indicator informations | XR6000,pioneer | What does the signals on indicators tell us?. | 1.X-ray ON Indicator Yellow. When system console receives the x-ray on signal, system console will light up the x-ray on indi- cator on the system console, and the x-ray on lamp on the room. In the meantime, system console will turn on the alarm sound. 2.Tube Overheat Indicator Red. When system console receives the tube overheat signal, system console will light this indicator up, and system exposure will be inhibited. When the light is off after appropriate cooling, exposures can be taken. 3.Inhibit Indicator Red. The system inhibit indicator lights up means there is an error. It includes: - Examination room door is not closed - Various inhibition errors on the system - Technique parameters are over the limit 4.Generator Ready Indicator Green. This indicator will light up when system is ready for exposure after pressing prep. If the prep switch is released, the Generator Ready indicator will be off." | https://drive.google.com/file/d/1hROr2w1CToY5uViR9M_WCYJda0AwQJwa/view?usp=sharing
GRGXR62 | WALL_STAND_MODE | Operating mode 1,radiography,ge | modes of operation | XR6000,pioneer | How to apply Wall Stand Mode? | 1. System power ON. 2. Tube warm-up. 3. Set technique parameters. a.) Select wall stand mode . b.) Select focal spot . c.) Adjust exposure parameters. 4. Place Cassette. 5. Adjust wall stand height by patient size. 6. Patient positioning. 7. Adjust tube angulation and SID. 8. Take exposure. 9. Take out cassette, and develop films. | https://drive.google.com/file/d/1FJnqNbkvuWtrrIgBsv1hYsuY6cI6It8R/view?usp=sharing
GRGXR63 | TABLE _MODE | operatig mode 2,radiography,ge | modes of operation | XR6000,pioneer | How to apply Table Mode? | 1. System power ON. 2. Tube warm-up. 3. Set technique parameters. a.) Select table mode . b.) Select focal spot . c.) Adjust exposure parameters.4. Place cassette onto the cassette tray under the table. 5. Patient positioning on the radiographic table.6. Adjust tube angulation and SID.7. Take exposure. 8. Take out cassette, and develop films. | https://drive.google.com/file/d/1JrpHEBzTn4eMrv86Bj7Ww2Kt9cfgznZc/view?usp=sharing
GRGXR64 | TABLE _TOP_MODE | operating mode 3,radiography,ge | modes of operation | XR6000,pioneer | How to apply Table Top Mode? | 1. System power ON. 2. Tube warm-up. 3. Set technique parameters. a.) Select table top mode . b.) Select focal spot . c.) Adjust exposure parameters.4. Patient positioning on radiographic table, and place cassette to the appropriate location. 5. Adjust tube angulation and SID.6. Take exposure. 7. Take out cassette, and develop films. | https://drive.google.com/file/d/1cNFN3-UgLi31TrGLH-qmNStqHJ4-493G/view?usp=sharing
GRGXR65 | STRETCHER_TABLE_MODE | operating mode 4,radiography, ge | modes of operation | XR6000,pioneer | How to apply Stretcher table mode? | 1. System power ON. 2. Tube warm-up. 3. Set technique parameters. a.) Select table top mode . b.) Select focal spot . c.) Adjust exposure parameters. 4. Rotate Tube Stand 180 Degree. 5. Patient positioning on stretcher table, and place cassette to the appropriate location.6. Adjust tube angulation and SI7. Take exposure. 8. Take out cassette, and develop films. | https://drive.google.com/file/d/1xoOaukIM0A1iN67BZuacmmKuG7CuYNjb/view?usp=sharing
GRGXR66 | SYSTEM_POWER_ON | system power options,power on, radiography,ge | sytem power options | XR6000,pioneer | How to power on the system? | Connect the Wall Brake, the POWER ON green indication lamp located on the right side of the POWER ON button on system console lights up. To turn ON the system, press the “POWER ON“ button on the top-right of system console. | https://drive.google.com/file/d/19E7zF9pjH0Onq_gxbSPtEgT0MwNXvAnj/view?usp=sharing
GRGXR67 | SYSTEM_POWER_OFF | system power options,power off, radiographyge | sytem power options | XR6000,pioneer | How to power of the system? | To turn OFF the system, press the “POWER OFF“ button on the top-right of system console. All of the subsystems power will be cut off at the same time. System console POWER ON green indication lamp will not be off until the Wall Brake is cut off. | https://drive.google.com/file/d/19E7zF9pjH0Onq_gxbSPtEgT0MwNXvAnj/view?usp=sharing
GRGXR68 | SAVING_CHANGES | saving changes,system status,radiography,ge | saving changes | XR6000,pioneer | How to save changes in system parameters? | After change the parameters, press “Save APR“ + “Retrieve APR“ (press “Save APR“ button first) for 3 seconds, system console will save the current parameters as the new system power-on initial mode.
GRGXR69 | TUBE_WARMUP_REQUIREMENTS | tube warm up,requirements,radiography,ge | tube warm up requirements | XR6000,pioneer | What are the tube warm up requirements? | 1. Use the larger focal spot of the tube, and the seasoning will be done with the circuit specified in technical data. 2. The operating tube voltage and current should be monitored all exposures during this seasoning operation. 3. The anode rotation should be utilized at normal speed (3000 or 3600 min-1). 4. Do not exceed either the generator or the tube ratings.
GRGXR610 | TUBE_WARMUP-PROCEDURE | tube warm up, procedures, radiography,ge | tube warm up procedures | XR6000,pioneer | What are the tube warm up procedures? | 1. Provide continuous input of small tube current (Approx.1 mA) and increase the tube voltage stepwise according to the time of the Seasoning Schedule at installation (See Figure 4-1). For any X-ray equipment without the fluoroscopy circuit, perform seasoning of short-time input from 50 kV. 2. Use the larger focus, rotate the anode, set the tube current to 50% of the maximum permissible tube current at maximum tube voltage, apply short-time input once a minute with exposure time of 0.1s, and increase the tube voltage stepwise. When unstable, states of the tube current are observed in the course of increasing tube voltage, reduce the tube voltage to the extent that no variation occurs, and then increase it more slowly than previously. 3. Stain like impact traces might appear on the X-ray tube target surface by the slight discharge in the course of the seasoning. The phenomena are one process to recover the withstanding voltage performance at that time, Even if unstable states appear in the course of the seasoning, the stable operation is obtained at the maxi- Direction 5275602-100 Rev 16 94 mum tube voltage of succeeding seasoning, the tube unit can be used without any interference to its electrical performance. (The maximum tube testing voltage must be strictly limited to the maximum rated tube voltage of the tube unit.) 4. When the X-ray tube unit is used after a dormant period of a week or more, use the tube unit after performing the Seasoning Schedule at Starting . 5. Every day morning, when you start to use X-ray tube, you are requested to warm-up the tube to your necessary high voltage condition for stabilized equipment operation. a.) At least 80kV (Step 1) is requested for all operation. b.) Need voltage up to your necessary daily voltage (Step 2 or higher step). c.) If discharge happened, re-warm up again under the previous condition.
GRGXR611 | CHANGE_APR_PROTOCOL | apr protocol,apr parameters,apr,radiography,ge | change apr protocol | XR6000,pioneer | How to overwrite an existing protocol? | 1. Select the body anatomy button under the APR area on the XR 6000 console. 2. Select any button that needs to be changed (receptor, focus spot size, kVp, mA, mAs, mSec, AEC, exposure view, or patient size). Once a button is selected that is not in the protocol, the body anatomy icon on LCD will flash notifying a change. 3. Press APR Save button for 3 seconds, a beep from system means the new protocol is saved. And the new protocol parameters will be saved as default in system console.
GRGXR612 | TO_MAKE_AN_HANDSWITCH_EXPOSURE | handswitch,handswitch exposure,radiography,ge | to make an handswitch exposure | XR6000,pioneer | How to make an handswtich exposure? | The handswitch connected to the system console is a three-position push button switch. Its three positions are OFF, Prep and Expose. The handswitch is normally in the OFF position. Press the handswitch halfway to the Prep position for 1~1.5 seconds. This prepares the X-ray tube for exposure. Then press the handswitch all the way down to the Exposure position and hold until the exposure is complete. If release the handswitch during this period, exposure will stop immediately. X-ray on indicator off and beep sound terminated notify that the exposure is complete. Prep on handswitch can be used again. If the operator does not release the Exposure switch after exposure operation, the next exposure will be inhibited. Only after release the Prep switch, a new exposure is permited.
GRGXR613 | TO_MAKE_AN_SYSTEM_CONSOLE_EXPOSURE | system console,system console exposure,radiography,ge | to make an system console exposure | XR6000,pioneer | How to make an system console exposure? | On the lower right hand corner of the system console under Radiography is where the Prep and Exposure buttons are located. Prep: Press this button for 1~1.5 seconds, system will prepare the tube for exposure. Exposure: If the ready indicator lights up, after press this button, the system will generate the X-ray. And system console will light up the x-ray on indicator on the system console and the x-ray on lamp on the room until the exposure is finished. In the meantime, the buzzer will sound until the exposure is finished. | https://drive.google.com/file/d/1_tOgFWmtaOeg_7UVdm00I57Z0QSMQlW0/view?usp=sharing
GRGXR614 | TABLE_TOP_MOVEMENT_OPERATING_INSTRUCTIONS | table top,table top movement,operating instructions table top | table top movement operating instructions | XR6000,pioneer | How to move the table top? | The tabletop is locked in place by electromagnetic locks when power is applied. To position the tabletop, the operator must first unlock the electromagnetic locks by pressing the foot pedal located along the base of the patient table two consecutive times . Hold the pedal down releases the longitudinal and lateral locks. To re-lock the tabletop into position, remove your foot from the pedal.
GRGXR615 | CASSETTE_TRAY_MOVEMENT_INSTRUCTIONS | cassette tray,movement instructions | cassette tray movemnt instuctions | XR6000,pioneer | How to move the cassette tray? | The cassette tray can be moved by pressing in the gray lock switch located under the hand grip. Then with the hand grip, move the cassette tray to the desired position. Release the lock switch or the hand grip after the desired position is reached | https://drive.google.com/file/d/1ybh-wOowVQdtXlyXDKbZZCblY2HVD96j/view?usp=sharing
GRGXR616 | CASSETTE_LOADING | cassette,cassette loading,cassette tray | cassette loading | XR6000,pioneer | How to load a cassette? | To load a cassette, hold the cassette handle and pull out the cassette tray. Place the cassette into the two clamps. Make sure the cassette’s center line and center of cassette tray align.When suitably positioned, lock the cassette by turning the knob clockwise until it is locked. To release and adjust, turn the knob counter clockwise. Finally, push the tray in. | https://drive.google.com/file/d/1amC3RrCZpsjwoUh3vVXUVkSCzeUlVubT/view?usp=sharing
GRGXR617 | CASSETTE_REMOVAL | cassette,cassette tray,cassette removal | cassette removal | XR6000,pioneer | How to remove a cassette? | To remove a cassette from the manual cassette tray, pull the tray out fully by its handle. Turn the knob counter clockwise to release the clamps. Push the cassette in and enlarge the space between the two clamps. The cassette is now free to be removed.
GRGXR618 | X-RAY_TUBE_ASSEMBLY_VERTICAL_MOVEMENT | xray tube,assembly,vertical movement, | xray tube assembly vertical movement | XR6000,pioneer | How vertical movement of xray tube assembly is achieved? | Press in the VERT switch (or ALL) on the Control panel, then hold the handles on the panel and move the tube assembly to the desired position. Notice the mark placed on the tube stand and chose desired SID. After the desired position is reached, release the switch on the tube control panel to re-engage the tube stand vertical locks
GRGXR619 | X-RAY_TUBE_ASSEMBLY_STAND_TRANSVERSE_MOVEMENT | xray tube,assembly,stand transverse movement, | xray tube assembly stand transverse movement | XR6000,pioneer | How transverse movement of xray tube stand is achieved? | Press in the TRANS switch (or ALL) on the Control panel , then hold the handles on the panel and move the tube stand to the desired position. After the desired position is reached, release the switch on the tube control panel, to re-engage the tube stand transverse locks. Remember when the tube stand is moved out of the range of the cassette tray, the cassette tray will decouple from the tube stand
GRGXR620 | X-RAY_TUBE_ASSEMBLY_STAND_ROTATION | xray tube,stand rotation | xray tube asssembly stand rotation | XR6000,pioneer | How to rotate xray tube stand? | The tube stand can rotate to any desired angle +/-180 degrees for off table radiographic images. To rotate the tube stand, grasp the handles on the x-ray tube stand control panel and turn the tube stand to the desired position. The rotation is automatically notched every 90 degrees. | https://drive.google.com/file/d/1PXmBTigeAy8pCB1KUzQtsCy2VrDpC6JA/view?usp=sharing
GRGXR621 | X-RAY_TUBE_STAND_ANGULATION | xray tube,angulation | xray tube stand angulation | XR6000,pioneer | How xray tube stand angulation is achieved? | The X-ray tube can angulate +/-120 degrees for angulation views, and there are detents at 0 degree, +/- 90 degrees, and +/-120 degrees. The distance from tube focal spot to the ground is from 555 mm to 1760 mm. A small window on the control panel shows the angulation of the tube. To change the X-ray tube angulation press in and hold the control panel ANG (or ALL) switch and position the tube assembly to the desired angulation. Then release the ANG switch when the desired angulation is reached. | https://drive.google.com/file/d/1DrPjvTtDxX-7BzGji1g00JBVkE6aHVFt/view?usp=sharing
GRGXR622 | X-RAY_TUBE_TILT | xray tube,tilt,movement | xray tube tilt | XR6000,pioneer | How is xray tube tilted? | The X-ray tube can tilt forward more than 30 degrees, and backward more than 20 degrees. Behind the tube are two mechanical locks. Release the locks, the tube is now free to be moved. Position the tube to the desired location and lock the tube in place. To replace the tube to the normal position, release the locks, and position the tube. On the side of the tube, there is an angulation indicator and an arrow to align the tube into correct position. | https://drive.google.com/file/d/1d9WVWobnSdA4_d09RBj6yRaEDhbsdwsU/view?usp=sharing
GRGXR623 | ALLIGNMENT_FIELD_LIGHT | field light,allignment,collimator | allignment field light | XR6000,pioneer | How is field light alligned? | Press the field light button located on the collimator face to activate the high-intensity field lamp.The field light is timed and automatically turns off 30 seconds after your release the switch. To activate the field light for another 30 seconds, press the button again when the light goes out.
GRGXR624 | POSITIONING_OF_COLLIMATOR | postion,collimatore,allignment | positioning of collimatore | XR6000,pioneer | How is a collimator positioned? | Position the collimator and tube using the field light and crosshair shadow as guides. The crosshair shows the field center. Collimate by adjusting the size of the field with the two control knobs located on the front of the collimator.. The left knob controls the transverse dimension of the blades, and the right blade knob controls the longitudinal dimension of the blades. A diagram on the face of the colllimator shows which knob controls which dimension. Field sizes from 0x0 up to 17” x 17” at a 102 cm SID can be obtained. The knob selectors indicate the field size for a selected SID. The maximum field size is 17” x 17”. Dial numbers on the calibrated scale are for 100 cm and 180 cm scale SID. A ruler is located on the left side of the collimator with the measurements in centimeters.
GRGXR625 | COLLIMATOR_ROTATION | rotation,collimator,allignment | collimator rotation | XR6000,pioneer | How is a collimator rotated? | 1. Unlock the Fix Knob located on the top of the collimator and turn the collimator to the desired position. 2. After the collimator is positioned, turn the knob back to the locked position. 3. Unlock the Fix Knob and turn the collimator back to the normal position. Tighten the Fix Knob. Make sure the collimator is in alignment with the cassette tray by turning on the collimator light and seeing that the light is aligned. | https://drive.google.com/file/d/19bM59xQrTR6OBIwyyNlhrMvbfkQWNldu/view?usp=sharing
GRGXR626 | DAP_METER_OPERATION | dap meter,operation | dap meter operation | XR6000,pioneer | How is a dap meter operated? | By Pressing [Reset] button, all the internal measured value memory is deleted and the display is reset to zero (0.0 or 0.00). The meter is ready for measurement again after [Reset] button is pressed. | https://drive.google.com/file/d/1Ren4kIuQJbwhZosT5KBmOj3HEyR8ZO-B/view?usp=sharing
GRGXR627 | ADVANCED_WALL_STAND_VERTICAL_POSITIONING | wall stand,allignment,positioning | advanced wall stand vertical positioning | XR6000,pioneer | How are vertical postioning of advanced wall stand achieved? | For advanced wall stand, follow the procedures: 1. Press and hold the switch button to release the lock. 2. Move the wall stand to the desired position. 3. Release the switch button to lock the wall stand.
GRGXR628 | WALL_STAND_CASSETTE_LOADING | wall stand,allignment,positioning,cassette loading | wall stand cassette loading | XR6000,pioneer | How to load a cassette in a wall stand? | 1. Pull the cassette tray out by holding the handle of the cassette tray. 2. Insert the cassette into the tray, and the spring on the cassette tray will automatically center the cassette. 3. Push the cassette tray back into the Tray holder.
GRGXR629 | ADVANCED_WALL_STAND_CASSETTE_LOADING | wall stand casstte tray, cassette loading | advance wall stand casstte loading | XR6000,pioneer | How to load a cassette in an advanced wall stand? | 1. To insert a cassette, pull the tray out of the Bucky to the tray stop. 2. Insert the shelf into the centering scale slots that correspond to the cassette size selected. 3. Lift the clamping lock handle to unlock it. 4. Slide the clamping apart to insert a cassette on the shelf. Clamp and center the cassette transversely on the tray. 5. Push the clamping lock against the cassette and lock it by pressing down the lock handle. 6. Push the cassette tray all the way into the Bucky.
GRGXR630 | STANDARD/ADVANCED_CASSETTE_REMOVAL | wall stand,cassette tray,cassette removal | standard advanced casette removal | XR6000,pioneer | How to remove a cassette ina wall stan? | To remove a cassette from the cassette tray, pull the tray out fully by its handle.
GNMDIVH31 | EMERGENCY_POWER_OFF_BUTTONS | safety,emergency power off,nuclear medicine | emergency power of buttons | NM discover VH | How does emergency power buttons work? | The camera is provided with two spring-loaded emergency buttons. 1.A wall mounted emergency button, located near the operator station according to the site layout 2. A red emergency button located on the ON/OFF pnale on the rear of the Gantry When either button is pressed, it stops any motion of the Detector Heads and powers down the system. To resume system motions, the depressed emergency button must be released by turning it a quarter turn clockwise, and the patient, if any, must be safely removed
GNMDIVH32 | PRESSURE_SENISITIVE_DEVICES_[PSD] | safety,emergency, pressure sensitive devices,nuclear medicine | pressure senisitive devices | NM discover VH | How does pressure senisitive devices work? | A Pressure Sensitive Device is provided on the surface of each collimator. A force greater than 2 kg (4.4 lbs) pressing on the PSD activates it. Activation of the PSD stops and disables all Gantry and Patient Table motions. To re-enable system’s motions: 1. In H-Mode: press the RADIAL <OUT> key of the Remote Control Unit together with the ENABLE lever .2. In L-Mode: press the LATERAL <OUT> key of the Remote Control Unit together with the ENABLE Lever
GNMDIVH33 | STRETCHER_RELEASE_HANDLE | safety,emergency,stretcher release handle,nuclear medicine | stretcher release handle | NM discover VH | How does stretcher release handles work? | The Stretcher Release Handle is located at the head of the Patient Table, that is the end farthest from the Gantry. When pressed, the Stretcher Release Handle disengages the Stretcher’s clutch to enable moving the Stretcher forwards and backwards manually. When the Handle is pressed, it stops the motion of the Detector Heads.
GNMDIVH34 | ENABLE_LEVER | safety,emergency,enable lever,nuclear medicine | enable lever | NM discover VH | How does enable lever work? | An ENABLE Lever is provided on the right hand side of the Remote Control Unit (RCU). To prevent movement of the Gantry when one of the RCU keys is accidently pressed, the system responds to an RCU key only when the ENABLE Lever is depressed. 1. Static, Dynamic and Gated acquisitions can be started by simply pressing <START/STOP> without depressing the ENABLE Lever. 2. During Whole Body SCAN, SPECT or GSPECT acquisitions, pressing the ENABLE Lever or any RCU key stops Gantry motions.
GNMDIVH35 | EMERGENCY_REMOVAL_OF_PATIENTS_WHEN_SAFETY_EVENT_OCCURED | safety,emergency,emergency removal of patients,safety event | emergency removal of patients when safety event is occured | NM discover VH | How to remove the patient when a safety event occured? | 1. In H-Mode: Press the RADIAL <OUT> key for each Head on the Remote Control Unit (RCU) to move the Detector Heads outwards away from the patient. In L-Mode: Press the LATERAL <OUT> key on the Remote Control Unit (RCU) to move the Detector Heads outwards away from the patient. 2. Press the Stretcher Release Handle to disengage the Stretcher’s clutch, and pull the Stretcher away from the Gantry 3. Assist the patient to get down from the Patient Table
GNMDIVH36 | EMERGENCY_REMOVAL_OF_PATIENTS_WHEN_SYSTEM_POWER_IS_DOWN | safety,emergency,emergencyremoval of patients when system power is down | emergency removal of patient when system power is down | NM discover VH | How to remove patients when sysytem power is down? | The system power can be down either due to a power failure or due to activation of one of the Emergency buttons. When system power is down, press the Stretched Release Handle and pull the Stretcher out and assist the patient to get down from the Table
GNMDIVH37 | EMERGENCY_REMOVAL_OF_PATIENT_DURING_BRAIN_SCAN | safety,emergency removal,brain scan, | emergency removal of patients during brain scan | NM discover VH | How to remove the patient during brain scan? | During a Brain scan, the Stretcher is already in its OUT position. So no special procedure is required to move the patient. Assist the patient to get down from the Patient Table.
GNMDIVH38 | EMERGENCY_REMOVAL_OF_PATIENT_DURING_ANY_OTHER_AQUISITION | safety,emergency,patient removal,data aquisition | emergency removal of patients during any other aquisition | NM discover VH | How to remove patients when during any other aquisition | 1. Press the Stretcher Release Handle to disengage the clutch and pull the Stretcher to its OUT position. 2. Assist the patient to get down from the Patient Table
GNMDIVH39 | ON_AND_OFF_PANEL | on and off,panel,system operation | on and off panel | NM discover VH | How to turn on and off the system? | The ON/OFF Panel, located on the upper left side at the rear end of the Gantry features three buttons: 1. Green ON button (marked I). When pressed, the system is set to the ON Mode. 2. Orange OFF button (marked O). When pressed, an Orange LED indicator on the this button is turned ON, and the system is set to STANDBY Mode. In this mode, the system can not be used but High Voltage is supplied to the Detectors, keeping them in operational state. 3.Red Emergency Stop button. When pressed, disconnects the power to all the system units (including the High Voltage to the Detectors), stops all motions, and turn off the STANDBY LED indicator. | https://drive.google.com/file/d/1pCS99pag4X7NQtp6QV4Jig1JePG2trPb/view?usp=sharing
GNMDIVH40 | SYSTEM_POWER_UP | power,system working | power up the system | NM discover VH | How to power up the system? | 1. Ensure that the system is connected to the power supply. 2. Ensure that both the Wall-Mounted emergency button and the Red emergency button on the ON/OFF Panel at the rear of the Gantry are released. 3. Lift up the Circuit Breaker located at the rear (lower right) of the Gantry to the ON position. 4. Press the ON button on the ON/OFF Panel located on the rear side of the Gantry. 5. Press the POWER button on the Computer. 6. Press the POWER button on the Monitor. When the system is restarted, it automatically verifies the integrity of all files.
GNMDIVH41 | SYSTEM_SHUTDOWN_AT_THE_END_OF_THE_WORKDAY | power,system working,system shutdown | system shutdown | NM discover VH | How to shutdown the system at the end of the workday? | 1. Verify that no removable disk is present in any of the disk drives. 2. Run the Shutdown procedure, that is, first close the System window and then perform system shut down.
GNMDIVH42 | CLOSE_AQUISITION_STATION_WINDOW | system shutdown,aquisition window | close aquisition window | NM discover VH | How to close the aquisition window? | 1. Click on the acquisition station window to make it active. 2. From the File Menu, choose the Exit option; this action closes the System window and all other opened System windows.
GNMDIVH43 | PERFORM_SYSTEM_SHUTDOWN | system shutdown | perform system shutdown | NM discover VH | How to perform system shutdown? | 1.Move the pointer to an empty part of the screen and click with right mouse button; the OS/2 Desktop Pop up menu appears. 2.Select the Shut down option. 3.The system prompts the user to confirm closing each window.4. A shutdown completed message appears on the screen. 5.Turn Monitor off.
GNMDIVH44 | RESTART_THE_SYSTEM | restart,system operation | restart the system | NM discover VH | How to restart the system? | 1. If the Gantry is in STANDBY mode, press the ON button on the ON/OFF Panel, located on the rear panel of the gantry. 2. If the Monitor was powered off, turn it on. 3. If the computer was turned off, switch the power on, otherwise, press <Ctrl>+<Alt>+<Del>.
GNMDIVH45 | SYSTEM_POWER_DOWN | power,system working,power down | system power down | NM discover VH | How to power down the system? | 1. Run the Shutdown procedure as described in step 2 of System Shutdown on 2. Press the OFF button on the ON/OFF Panel (located on the rear panel of the Gantry). The system is set to the STANDBY mode. 3. Lower the Circuit Breaker to the OFF position
GNMDIVH46 | LEAF_THROUGH_THE_PAGES | pages,parameter values | leaf through the pages | NM discover VH | How to leaf through the pages? | Click the appropriate page label at the base of the notebook OR Click the arrows (left or right) on the bottom right corner of the Notebook
GNMDIVH47 | MOVE_FROM_ONE_FIELD_TO_THE_OTHER | field ,parameter values | move from one field to other | NM discover VH | How to move from one field to the other? | Use the tab keys OR Roll the trackball to position the pointer at the required field and click the Left button
GNMDIVH48 | MOVE_WITHIN_A_FIELD | field,parameter values | move within a field | NM discover VH | How to move within a field? | Use the arrow keys.
GNMDIVH49 | DROP_DOWN_LISTS | drop down lists,parameters | drop down lists | NM discover VH | How does drop down lists work? | 1. Click on the arrow to the right of the field to display the drop down list. 2. Select the required element from the list. Use the arrows to scroll up and down the list.
GNMDIVH50 | SELECT_A_CHECK_BOX | lists,check box | select a check box | NM discover VH | How to select a check box? | Click the check box; a tick mark appears to indicate that it is selected.
GNMDIVH51 | CLEAR_CHECK_BOX | lists,check box | clear a check box | NM discover VH | How clear a checkbox? | Click the check box; the tick mark disappears.
GNMDIVH52 | SELECT_A_RADIO_BUTTON | radio buttons,lists | select a radio button | NM discover VH | How to select a radio button? | Click on the button; the button is filled in indicating that it is active.
GNMDIVH53 | DESELECT_A_RADIO_BUTTON | radio buttons,lists | deselect a radio button | NM discover VH | How to deselect a radio button? | Click on a radio button that is not currently selected; the button is no longer filled in.
GNMDIVH54 | ACTIVATE_A_BUTTON | buttons,display panel,aquisition panel | activate a button | NM discover VH | How to activate a button? | Position the mouse pointer on the button and click with left mouse button.
GNMDIVH55 | OPEN_DISPLAY_CONTROL_PANEL | buttons,display panel | open display panel | NM discover VH | How to open display panel? | Position the mouse pointer on the [DISPLAY] button and click the left mouse button | https://drive.google.com/file/d/1SvVdIO2tkLg9AxtvCiElC7pmZlEuN1r-/view?usp=sharing
GNMDIVH56 | SELECT_DESELECT_WINDOWS_IN_IMAGE_AREA_FOR_DISPLAY_MANIPULATION | display panel, display manipulation,display controls | select/deselect windows in image area for display manipulation | NM discover VH | How to select/deselect windows in image area for display manipulation? | Press the key and double-click the right mouse-button.
GNMDIVH57 | DISPLAY_CONVERION_OPTIONS | display panel, display conversion,display controls | display conversion options | NM discover VH | How to display conversion options? | 1. Click with the left mouse button on the arrow; the drop down list-box displays the available Conversion Tables. 2. Select the required Table
GNMDIVH58 | DISPLAY_AVAILABLE_COLOR_PALETTES | colour palattes,display | display available colour palettes | NM discover VH | How to display available colour palattes? | 1. Click with the left mouse button on the arrow; the drop down list-box displays the available Color Palettes. 2. Select the required Palette
GNMDIVH59 | ACTIVATE_FOV_ROI | fov,roi | activate fov roi | NM discover VH | How to activate FOV ROI? | • Click with the left mouse button on the checkbox to select this option | https://drive.google.com/file/d/1PEUd54pCbiguRByHH7ml78pE3LkkbfWu/view?usp=sharing
GNMDIVH60 | SELECT_A_DESIRED_AQUISITION_AREA | fov,roi | select a desired aquisiton area | NM discover VH | How to select a desired aquisition area? | You can select the desired acquisition area by changing the Size, Rotation and Position of the ROI. The mouse pointer changes shape when it is over the ROI. The shape determines the operation. • Press the right mouse button of the trackmouse and roll the trackball to manipulate the ROI.
GNMDIVH61 | CHANGE_THE_SIZE_OF_ROI | fov,roi | change the size of roi | NM discover VH | How to change the size of roi? | • Position the mouse pointer over the ROI; the mouse pointer changes into a double arrow. • Holding down the right mouse button, roll the trackmouse.
GNMDIVH62 | MOVE_THE_ROI | fov,roi | move the roi | NM discover VH | How to move the roi? | • Position the mouse pointer within the ROI; the mouse pointer changes into a four-sided arrow. • Hold down the right mouse button, roll the trackball/mouse.
GNMDIVH63 | ROTATE_THE_ROI | fov,roi | rotate roi | NM discover VH | How to rotate roi? | • Position the mouse pointer over the ROI compass; the mouse pointer changes into a circular arrow. • Hold down the right mouse button, roll the trackball/mouse.
GNMDIVH64 | INCOPORATE_CHANGES | fov,roi | incoporate changes | NM discover VH | How to incoporate the changes? | • Click the [PERSIST] button to incorporate the changes, OR • Deselect the FOV ROI option.
GNMDIVH65 | SELECT_PHA_OPTIONS | pha,display | select pha options | NM discover VH | How to selsct pha options? | Click with the left mouse button within the checkbox. The PHA histogram appears on the lower half of the Image Area The Dialog Window automatically displays Page 2 (EXTENDED)
GNMDIVH66 | REPOSITION_THE_MARKERS | pha,reposition markers,display | reposition the markers | NM discover VH | How to reposition the markers? | 1. Roll the trackball to position the pointer on the marker to be repositioned; the pointer changes shape. 2. Press and hold the right mouse button and roll the trackball to move the marker in the desired direction. 3. Release when at the desired location. Any change made manually is reflected on-line in the Energy Parameters on Page 2 (EXTENDED) of the Dialog Window.
GNMDIVH67 | CANCELLING_THE_PROTOCOL | cancellling the protocol | cancelling the protocol | NM discover VH | How to cancel the protocol? | Click the [CANCEL] button. The system prompts from the Cancel Confirmation window.  Click [YES] to exit from the protocol; the system closes the Acquisition Dialog Window.  To continue the acquisition procedure, click [NO]. | https://drive.google.com/file/d/1srthJumttCB2gNDy9-48BGQxCbdQLDMR/view?usp=sharing
GNMDIVH68 | STARTING_AQUUISITION_PROCEDURES | data aquisition | starting aquisition procedures | NM discover VH | How to start aquisition procedures? | • Click on [START] with the left mouse button. The [STOP] button replaces the [START] button. Dialog Window disappears from screen. Message Area displays:acquiring Image Area displays images acquired from each Head. | https://drive.google.com/file/d/1OOudEN6K_cuNLxkXqvPagwt09Gs-FuZU/view?usp=sharing
GNMDIVH69 | STOPPING_AQUISITION_PROCEDURES | data aquisition | stopping aquisition procedures | NM discover VH | How to stop aquisition procedures? | • Click the [STOP] button. The system sounds a beep. Message Area displays:aquisition terminated The ]NEXT VIEW] and [ICONIZE] and [TRANSFER] buttons are enabled. If automatic CLOSE DATA function is not activated, the [CLOSE DATA] button also appears. | https://drive.google.com/file/d/1Txjf491QrlLTbbEx3RWIrPyXcphqR5qE/view?usp=sharing
GNMDIVH70 | ACTIVATING _PAUSE | data aquisition,aquisition window | activating pause | NM discover VH | How to activate pause in aquisition window? | • Click the [PAUSE] button. In PERSISTence mode, temporarily freezes the persistence image. When a STATIC acquisition is in progress:  Temporarily freezes the acquisition procedure.  Replaced by the [CONTINUE] button.  Message Area displays:paused | https://drive.google.com/file/d/1fNpA811Tq2o0RTwQqD_Mj-2veJyUDrvm/view?usp=sharing
GNMDIVH71 | ACTVATING_CONTINUE | data aquisition,aquisition window | activating continue | NM discover VH | How to activate continue in aquisition window? | • Click [CONTINUE]. The button is replaced by the [PAUSE] button. In PERSISTence mode, clears and restarts the persistence image. When a STATIC acquisition is in progress:  Continues the acquisition from the point at which it was interrupted.  Message displayed before [PAUSE] was activated re-appears in the Message Area.
GNMDIVH72 | PERFORMING_ABORT | data aquisition,aquisition window | performing abort | NM discover VH | How to perform abort? | • Click the [ABORT] button. The current acquisition type is shown (e.g. Static/Dynamic/ECT/Scan/SYMA/ GSPECT)  Click [YES] to abort protocol; the system returns to the Acquisition Protocols window.  Otherwise, click [NO] or [Cancel] to continue acquisition. | https://drive.google.com/file/d/1JprdBRtv_gWzFAK5Cs2CAz_sDfN1pkaj/view?usp=sharing
GNMDIVH73 | ACTIVATING_NEXT_VIEW | data aquisition,aquisition window | actiavating next view | NM discover VH | How to activate next view? | • Click the [NEXT VIEW] button. Restores the Dialog Window if it has been iconized. Displays the previous Dialog; the cursor by default blinks at the VIEW field. Image Area displays the persistence image using the dialog values. [START] button is enabled. Message Area display:ready for start | https://drive.google.com/file/d/1k5BYz3EGtkJqSXxqMEwQejLZaUh0Lenr/view?usp=sharing
GNMDIVH74 | ACTIVATING_ICONIZE | data aquisition,aquisition window | activating iconize | NM discover VH | How to activate iconize? | • Click the [ICONIZE] button. Minimizes Acquisition Window and displays only the Header
GNMDIVH75 | RESTORING_AN_ICONISED_WINDOW | data aquisition,aquisition window | restoring iconised window | NM discover VH | How to restore iconised window? | • Click on the maximize button at the upper right corner of the Title Bar, OR Double-click the title bar.
GNMDIVH76 | SWIVELING_THE_PATIENT_TABLE | patient table,manual operations | swiveling the patient table | NM discover VH | How to swivel the patient table? | The Table has to be moved sideways during collimator exchange in order to position the Collimator Cart between the Detector Heads. 1. With the Stretcher positioned at its OUT limit, depress the Foot Clutch and pull the Patient Table along the Table Track using the Hand Grips. 2. Pull the Table until the Table mounts the Floor Stopper at the end of the Table Track
GNMDIVH77 | RELEASING_THE_TABLE | patient table,manual operations | releasing the table | NM discover VH | How to release the patient table? | The Table has to be moved sideways during collimator exchange in order to position the Collimator Cart between the Detector Heads. 1. With the Stretcher positioned at its OUT limit, depress the Foot Clutch and pull the Patient Table along the Table Track using the Hand Grips. 2.Pull the Table until the Table mounts the Floor Stopper at the end of the Table Track
GNMDIVH78 | MOVING_THE_STRETCHER_MANUALLY | patient table,manual operations | moving the stretcher manually | NM discover VH | How to move the strectcher manually? | Though the Table can be moved inwards/outwards using the dedicated keys on the RCU, it can also be physically moved as follows: Press the Stretcher Release Handle inwards while pushing or pulling the Stretcher.
GNMDIVH79 | ADJUSTING_THE_HEIGHT_OF_THE_TABLE | patient table,automatic operations | adjusting the height of the table | NM discover VH | How to adjust the height of the table? | • Press the <UP/DOWN> keys on the RCU; the current height of the Table is displayed on the gantry display unit (GDU).
GNMDIVH80 | MOVING_THE_STRETCHER_TOWARDS_OR_AWAY_MANUALLY | patient table,automatic operations | moving the patient table towards or away manually | NM discover VH | How to move the patient table towards or away manually? | Press the <IN/OUT> keys on the RCU; the current position of the Table is displayed on the gantry display unit (GDU).
GNMDIVH81 | SETTING_HOME_POSITION_FOR_COLLIMATOR_EXCHANGE | collimator exchange,position | setting home position for collimator exchange | NM discover VH | How to set home position for collimator exchange? | 1. Pull the Stretcher to its maximum OUT position. 2. Position the Detector Heads at 0o and 180o in order to avoid collision between the Patient Table covers and the Head covers when the Table is set to the Home Position for collimator exchange. 3. Press the Foot Clutch of the Patient Table; use the Hand Grip and pull the table sideways along the table track on the floor, until the table is mounted on the floor stopper. 4. When the table is fully mounted on the stopper, the Gantry Display Unit displays the following message.bed is mounted in slanting position If not, the GDU displays:bed is misplaced Check to see if there is any obstruction on the table track. Continue pulling until you hear a click (which indicates that the table is mounted) and the message “Bed is mounted in slanted position” re-appears. 5. Verify that the Home Position displayed on the GDU is:coll exch Press the <HOME> key on the RCU to position the Detector Heads at the Home Position required for the exchange procedure.
GNMDIVH82 | REMOVING_COLLIMATORS | collimator exchange,position | removing collimators | NM discover VH | How to remove collimators? | 1. When the Home Position for the collimator exchange procedure is reached, the GDU displays the message:place coll cart.prsess<stop> to quit To quit the procedure, press the <STOP> key on the RCU; the Gantry prompts the operator to perform the P.S.D. tests, after which it returns to IDLE. Otherwise, to continue, roll an empty Collimator Cart towards the Gantry and position it between the Detector Heads. Step on the Cart Locking Pin to lock the cart in place (see Figure 4-7 on page 4-46). 2. The Gantry detects the presence of the cart and the GDU displays the message:cart placed.press <cont>or<stop> To quit the procedure, press the <STOP> key on the RCU; the GDU displays the message:remove collimator cart Step on the black Footplate to release the Cart Locking Pin. Roll the empty cart completely away from the Gantry. The Gantry returns to IDLE. Otherwise, to continue, press <CONT> on the RCU; the Detector Heads start moving inward, toward the cart and stop when they reach the correct position. During the motion, the GDU displays the message:caution automatic radial motion3. When the Heads reach the correct position the GDU prompts: open head handles or press <continue> or <stop> To open each Head Handle, pull the handle out and move it to the “Open” position (as indicated on Label). The pin must be inserted into the hole in order for the collimator to be fully in the “Open” position. To quit, press the <STOP> key on the RCU; the Detector Heads then move outward, away from the Gantry with the collimators still loaded on the Heads.During the motion, the GDU displays the message:caution automatic radial motion and then prompts remove collimator cart Step on the black Footplate to release the Cart Locking Pin. Wheel the empty cart away from the Gantry. The Gantry returns to IDLE.If one of the Head handles is in the open position and the <STOP> key is pressed, the GDU displays: Coll #1 or 2 Unlocked! If at least one of the handles of the heads is closed and the <CONT> key is <cont>or <stop> To continue the procedure, open the Handles of both Heads and press <CONT>. The Radials move out. During motion, the GDU displays:caution:automatic radial motion When the Heads return to their position, the GDU displays: replace next cart .press<cont> to continue 4. Step on the black Footplate to release the Cart Locking Pin and roll the loaded cart to the storage area. The Detector Heads are temporarily without collimators. | https://drive.google.com/file/d/1wWg7kP2_zkQpXtGVEtRwXVomGpFxa2t8/view?usp=sharing
GNMDIVH83 | RELOADING_COLLIMATORS | reloading collimators,collimator exchange | reloading collimators | NM discover VH | How to reload collimators? | 1. Wheel the Collimator Cart with the required set of collimators towards the Gantry and position it between the Detector Heads. 2. Step on the Cart Locking Pin to lock the Collimator Cart in place and press the <CONT> key on the RCU to confirm. In case the Gantry does not ‘detect’ the Cart, the GDU displays: Position the Cart correctly and press <CONT>. 3. The Detector Heads move inward toward the Cart and stop at the correct position (approximately 16.5 cm) when the pins of the Detector Heads attach themselves to the pins of the loaded collimator. The GDU displays: 4. Close the handles and verify that the handles are fully locked. 5. Press <CONT> to continue. Note Do not press the <HOME> key on the RCU since it will initiate Gantry motion. Cart not in place. Press <CONT> to cont. Close Head Handles. Press <CONT> or <STOP> If at least one of the head handles is open, the GDU displays After 5 seconds the GDU displays a blinking message: The Gantry waits for full unlocking (both head handles must be opened). After opening the head handles the GDU displays Press <Cont> on the RCU. The detector heads move outward without the collimators. Try to load new collimators with a new cart. Coll #1 or 2 Unlocked! Open Handle #1 or 2 Press <Cont> to Repeat Operation 6. If both head handles are closed, the detector heads move outward with the loaded collimators during the motion, the GDU displays: 7. After Radial motion has stopped, the GDU displays: Step on the black Footplate to release the Cart Locking Pin and wheel the empty cart completely away from the Gantry to the storage area. 8. The GDU prompts to test the PSDs. Press the PSD on each Head. 9. The GDU prompts: 10. Once the table has been moved from the slanted position to the center, the GDU prompts: After CONT is pressed, a rotation of 270o will be performed in order to verify that the collimators are firmly anchored in place. Wait until the rotation is complete. Caution: Automatic Radial Motion Remove the collimator cart Move the table to center position Press <CONT> to start rotation
GNMDIVH84 | SELECTING_A_HOME_POSITION | collimator exchange,position | selecting a home position | NM discover VH | How to select a home position? | Home Positions can be selected in either of two ways: z Through the specific acquisition protocol, OR z Using the <MODE> key on the RCU to scroll through the available Positions. Only one User Home Position is available through the RCU. This is the User Home Position loaded to the Gantry via the Gantry Handler. In both cases the user has to press the <HOME> key on the RCU to set the Position and move the Detector Heads to the mode required for the selected Home Position
GNMDIVH85 | DISPLAY_A_DIAGNOSTIC_GROUP_CONTAINER | display group,aquisition protocol | display a diagnostic group container | NM discover VH | How to display a diagnostic group container? | From the Acquisition Station window, double-click the icon of a diagnostic group; a window that includes the list of protocols belonging to this group appears.
GNMDIVH86 | CHANGING_SETTINGS_OF_PROTOCOL_WINDOW | changing view settings,aqusition protocols | changing settings of protocol window | NM discover VH | How to change settings of protocol window? | The protocols windows have a Menu Bar with a number of options that can be used to change various settings of these windows. It should be noted that the settings changed through these menu options, do not for the most part affect or influence the actual acquisition procedures. To effectively change a procedure, the user has to modify parameter values from within the relevant protocol (that is, the protocol has to be activated) and then save the changes. Most of the menu options in the protocol windows are identical to all other windows.
GNMDIVH87 | SAVING_CHANGES_TO_VIEW_SETTINGS | changing view settings,aqusition protocols,saving changes to view settings | saving changes to view settings | NM discover VH | how to save changes to view settings? | Changes made to the appearance of a window through the options of the View menu are described below. These changes may be saved; the next time you open the window, it opens with the new settings. To save changes, from the File menu, select Save. To exit without saving, select Quit.
GNMDIVH88 | CHANGING_VIEW_SETTINGS | changing view settings, aquisition protocols,saving changes | changing view settings | NM discover VH | how to change viewing settings? | The options provided in the View menu pertain to changing the appearances of the contents in various ways. Some of the options are identical to those of the Acquisition Station window, In addition, the View menu offers the Detail, Sort, Filter, Refresh now, Hide and Show All options, described in the following sections
GNMDIVH89 | SORTING_CONTENTS_OF_PROTOCOL_CONTAINERS | aquisition protocol,sorting contents | sorting contents of protocol containers | NM discover VH | How to sort contents of protocol containers? | The Sort option allows you to organize the contents of the containers according to your convenience. The Dialog Box enables you to define the field and the direction of the sort. 1. From the View menu, select Sort; the Sort Parameters Dialog Box appears: 2. Select the sort Field by clicking the up and down arrow of the spin button. You can sort according to:• Date – denotes the date on which the protocol was created. • Clinical – application (diagnostic group) to which it belongs. • Type – acquisition type - Static, Dynamic, ECT, SYMA, GSPECT, SCAN. • Name – protocol name. • Icon – icon that appears in the diagnostic group container. • Label – protocol label - this string appears alongside theicon in the diagnostic group containers. • Frame Size – acquisition Frame Size. • Isotope – isotope used in protocol 3. Under Direction, click either the Ascending or the Descending radio-button to select the order of the sort: 4. Click [OK] to perform the sort operation. Otherwise, click [Cancel].
GNMDIVH90 | FILTERING_CONTENTS_OF_PROTOCOL_CONTAINERS | aquisition protocol,filtering contents | filtering contents of protocol containers | NM discover VH | How to filter contents of protocol containers? | The user can select the type of entries to be displayed in a window through the filter option in the View menu. This option also enables the user to search for protocols according to specific criteria. Similar to the Sort command, the user has to define the key field, that is, the Filter to be used. 1. From the View menu, select the Filter option; the Filter Parameters Dialog Box appears. 2. Select the filter Field by clicking the up and down arrow of the spin button. The Filter fields (identical to the Sort fields) are: • Date – denotes the date on which the protocol was created. • Clinical – application (diagnostic group) to which it belongs. • Type – acquisition type – Static, Dynamic, ECT, SYMA, GSPECT, SCAN, PET, TE • Name – protocol name. • Icon – icon that appears in the diagnostic group container. • Label – protocol label – this string appears alongside the icon in the diagnostic group containers. • Frame Size – acquisition Frame Size. • Isotope – isotope used in protocol. 3. Type a prefix to designate the prefix to be matched against the values of the selected field by the filter operation; for example, with Label as the filter field, in order to display only studies whose labels begin with the letters ‘st’, type ‘st’ as the Prefix. 4. Click [OK] to confirm; otherwise click [Cancel].
GNMDIVH91 | RESTORING_FILTERED_CONTAINER | aquisition protocols,restoring filtered container | restoring filftered container | NM discover VH | How to restore filtered container? | 1. Open the Filter/Find Parameters Dialog Box once again. 2. Select the Filter field last used; leave the Prefix field empty. 3. Click [OK]. The original contents are restored.
GNMDIVH92 | HIDING_AND_REVEALING_CONTAINER_PROTOCOLS | aquisition protocols,hiding and revealing container protocols | hiding and revealfing container protocols | NM discover VH | How to hide or reveal container protocols? | The user can hide one or more protocols included in a container, by selecting them and then selecting the Hide option from the View drop-down menu To show all the hidden protocols, select the Show All option from the View dropdown menu.
GNMDIVH93 | MODIFYING_PROTOCOL_SETTINGS | aquisition protocols,modifying protocols | modifying protocol settings | NM discover VH | How to modify protocol settings? | Changing the protocol identification through the Menu options Changing parameter values from within the protocol itself | https://drive.google.com/file/d/1xTyzsfVfov7lCvJb16exMWRcANwItwJr/view?usp=sharing
GNMDIVH94 | SELECT_SINGLE_PROTOCOL | aquisition protocols,single protocol | select single protocol | NM discover VH | How to select single protocol? | Click on the protocol icon; the selected protocol is highlighted
GNMDIVH95 | SELECT_ALL_PROTOCOLS | aqusition protocols,all protocols | select all protocol | NM discover VH | how to select all protocols? | From the Edit menu, click Select All; all protocols in the window are highlighted.
GNMDIVH96 | DESELECT_SELECTED_PROTOCOL/S | aqusition protocols,all protocols | deselect selected protocol/s | NM discover VH | How to deselect selected protocol/s? | From the Edit menu, click Deselect All; no protocol are highlighted.
GNMDIVH97 | MODIFY_PROTOCOL_IDENTIFICATION_SETTINGS | aquisition protocol,modify settings | modify protocol identification settings | NM discover VH | How to modify protocol identification settings? | 1. From the Acquisition Station window, double-click a diagnostic group to display the list of acquisition protocols. 2. Select the protocol to be modified. 3. From the Edit menu, select the Modify Protocol option; the Modify Protocol dialog box appears: 4. Type in new values for the following fields:• Protocol Name User assigned, identifying name of up to eight characters. • Clinical Application Diagnostic group to which protocol belong. • Protocol Label Descriptive label that appears alongside the icon in the diagnostic group container window • Study Label Default label used for a new Study created by the protocol. This is used to determine the organ so that the Study will appear in the appropriate catalog.. 5. Click [OK] to confirm; the relevant values are saved and will be used until further changes are made. If relevant, the icon is moved to the new diagnostic group or the label alongside the icon is updated. Otherwise click [Cancel].
GNMDIVH98 | ACTIVATING_PROTOCOLS | aquisition protocol,modify settings | activate protocols | NM discover VH | How to activate protocols? | From the Acquisition Station window, double-click the diagnostic group. Double-click on the protocol icon; the selected protocol is activated. OR Click on a required protocol to select it, and from the Menu Bar select Activate.
GNMDIVH99 | MODIFYING_PARAMETER_VALUE | aquisition protocol,modify settings | modifying parameter value | NM discover VH | How to modify parameter values? | 1. With the required protocol activated, modify values as required. The PERSIST button is illuminated indicating that the Persistence image no longer reflects the Dialog values. 2. When all values have been modified, click the PERSIST button once again to incorporate changes.
GNMDIVH100 | MASKING_PROTOCOL _PARAMETERS | aquistion protocol,masking protocol | masking protocol parameters | NM discover VH | How to mask protocol parameters? | 1. From the System Menu at the top-left corner of the Acquisition Dialog Notebook, select Mask mode (if not selected). 4. When Mask Mode is ‘ON’, the menu item is checked (9 Mask Mode). Parameters on the first MAIN page can now be masked. 2. Move pointer to required parameter field and click with right mouse button; the parameter appears dimmed. For example, to mask the parameter Isotope, position pointer over the field and click with right mouse button. If multiple parameters are to be masked, click parameters as required using right mouse button. 3. From the System Menu, select Mask Mode again; Mask Mode is now ‘OFF’ and the menu item is not checked (Mask Mode). The masked parameter(s) disappear from the Dialog Notebook.
GNMDIVH101 | SAVING_CHANGES | aquisition protocols,all protocols,saving changes | saving changes | NM discover VH | How to save changes in a protocol? | After having modified parameter values or masked one or more parameters, the user can save the changes in one of two ways: Save changes to the current protocol, provided that this is a user defined protocol. Save changes to a new user defined protocol.
GNMDIVH102 | SAVING_CHANGES_TO_THE_CURRENT_PROTOCOL | aquisition protocols,all protocols,saving changes | saving changes to the current protocol | NM discover VH | How to save changes to the current protocol? | It is not possible to modify factory provided protocols. The Save As... option should be used in this case to create new user protocols.1. From the System Menu at the top-left corner of the Acquisition Dialog Notebook, select Save Protocol; the ‘Save acquisition protocol’ Dialog Box appears: 2. If necessary in the Label field, type in a different protocol label. This label will appear alongside the icon in the diagnostic group container. 3. In the Preset field add the name of the preset to be activated at the end of the acquisition. 4. Click [OK] to confirm; otherwise click [Cancel]. | https://drive.google.com/file/d/1aDgZDv37_52-aYmPBhkuwq2xal7XX3ld/view?usp=sharing
GNMDIVH103 | SAVING_CHANGES_TO_THE_NEW_PROTOCOL | aquisition protocols,all protocols,saving changes,new protocol | saving changes to the new protocol | NM discover VH | How to save changes to the new protocol? | 1. From the System Menu at the top-left corner of the Acquisition Dialog Notebook, select the Save protocol As... menu entry; the ‘Save acquisition protocol’ Dialog Box appears: In the File Name entry field, type the new name for the protocol. The name must be unique for the acquisition type and can be up to eight characters long. The existing label appears in the Label field; type a new label. This label will appear alongside the icon in the diagnostic group container. If necessary, select a diagnostic group from the Clinical Application drop down list box. 2. In Preset field add the name of the preset to be activated at the end of the acquisition. 3. Click [OK] to confirm all changes; a new icon for the new protocol is added within the diagnostic group specified. Otherwise click [Cancel]. | https://drive.google.com/file/d/1Tipb9xx5Pd62IdYtbHPb4TlG00RGbNiT/view?usp=sharing
GNMDIVH104 | BACKUP_AND_RESTORE | aquisition protocols,all protocols,backup,restore | backup and restore | NM discover VH | How to backup and restore protocol? | To backup user defined protocols: 1. On the Desktop, double-click the Service icon to open the Service folder. 2. Double-click the BACKUP icon, then double-click the USER icon and follow the displayed instructions. To restore user defined protocols: 1. On the Desktop, double-click the Service icon. 2. In the opened Service folder, double-click the RESTORE icon, then double-click the USER icon and follow the displayed instructions.
GNMDIVH105 | DELETE_ PROTOCOLS | aquisition protocols,delete protocols | delete protocols | NM discover VH | How to delete protocol? | One or more user defined protocols can be deleted at a time. 1. Click on the protocol name or icon to be deleted. To select multiple protocols, hold down the <CTRL> key and click each protocol. To select all protocols in the list, from the Edit menu, click Select All. 2. From the Edit menu, select Delete protocol; the Delete Protocol(s) window appears. Each protocol is identified according to the protocol type, name and diagnostic group, separated by periods. If necessary, use the up/down arrow-heads to scroll through the list. 3. Click [OK] to confirm delete; otherwise click [CANCEL] | https://drive.google.com/file/d/1SPZmOp4V-QIxC60MXTFMcySisVrg6MXZ/view?usp=sharing
GNMDIVH106 | SELECTING_A_PROTOCOL | selecting a protocol,general aquisition operation | selecting a protocol | NM discover VH | How to select a protocol? | From the Acquisition Station window, double-click the desired diagnostic group icon. The diagnostic group container is opened. Double-click on the protocol icon. The selected protocol is activated; the Acquisition Notebook Dialog with the values last used for this protocol appears at the left side of the screen, while the Acquisition Window appears at the right side of the screen. For system response to the activation of a specific acquisition type, refer to the relevant chapter
GNMDIVH107 | ENTERING_PATIENT_INFORMATION | entering patient information,general aquisition operation | entering patient information | NM discover VH | How to enter patient information? | There are four options to enter patient information: z Manually, for a new study. In this case, first the protocol is activated, and then patient information is typed into the relevant input fields on the MAIN Page and on the Personal Page. (mandatory)The three patient information fields on the MAIN page (Patient ID, Last and First) are mandatory From the Work List, if the study is defined there, as follows:  Activate the protocol.  Click the [WorkList...] button located on the MAIN page of the Acquisition Notebook Dialog to open the catalog. The local catalog is opened, showing only the TODO studies for the current protocol.  Select the desired study. The Local catalog is closed and the patient information is loaded into the Acquisition Notebook dialog.  Click the [PERSIST] button to incorporate the updated information. From an existing study in the Local Catalog, for acquiring a new group into an existing study. In this case, after activating the protocol:  Type in the number of the relevant study into the Study ID field, OR  Click the [Catalog...] button on the MAIN page of the Acquisition Notebook Dialog to open the catalog, and then select the desired study. The patient information from the selected study is loaded and a new group number is assigned (the next free number with regard to the group(s) in the selected study.  Click the [PERSIST] button to incorporate the updated information.
GNMDIVH108 | PATIENT_POSITIONING | patient positioning,general aquisition operation | patient positioning | NM discover VH | How to position the patient? | There are two ways of positioning the patient on the Patient Table: LEGS IN. LEGS OUT. to acquire STATIC and DYNAMIC studies on: Patients sitting in a chair or wheelchair Patients in a standing position Special care must be taken when positioning the patients for Brain Scans
GNMDIVH109 | SITTING_ON_A_CHAIR_OR_WHEELCHAIR_OR_STANDING | patient positioning,general aquisition operation | siting on a chair or wheelchair or standing | NM discover VH | How is a patient seated or put in a standing position? | 1. Move the table to SLANTED position. 2. Activate the required STATIC or DYNAMIC protocol. When in Persistence phase, select the HOME position VERTICAL using the MODE button on the RCU and press HOME. This Home position allows Lateral and Radial movements of the heads. It does not allow rotation. 3. Locate the patient between the Laterals. See Figure 6-2 on page 6-21. 4. Adjust the HEADS to the desired position using the Lateral buttons on the RCU. Adjust Radial position if necessary. The Lateral movement limits are from 65 cm to 150 cm. 5. Activate the protocol by pressing START. 6. When the acquisition has terminated, the patient should be removed. The software will ask for confirmation before moving to next HOME position. The operator is asked to press CONT to confirm that patient was removed. | https://drive.google.com/file/d/16vCpQzBEBQXnBhgElmH78pvpvgn2p37g/view?usp=sharing
GNMDIVH110 | PATIENT_POSITIONING_FOR_BRAIN_SCANS | patient positioning,general aquisition operation,brain scans | patient positioning for brain scans | NM discover VH | How is a pateint positioned for brain scan? | To position the patient for brain scans: Retract the Patient Table completely. Remove the mattress from the stretcher. Place the Head Holder on the stretcher so that only the part supporting the patient’s head protrudes from the stretcher, and attach the hook bands on the bottom of the Head Holder to the loop bands on the stretcher. Place the mattress on the stretcher.Select the appropriate home position via the <MODE> key on the RCU. Press the <HOME> key on the RCU to bring the Detector Heads to the required position Carefully position the patient on the Head Holder, so that his/her head is placed on the head support. If needed to prevent patient motion, cover the head bands with a disposable tissue and use the head bands to secure the patient’s head.z Move the Stretcher IN (via the RCU), so that the patient’s head is centered within the Field of View (FOV). Rotate the Gantry carefully, while checking Detector Heads clearance. Pay special attention to the patients shoulders | https://drive.google.com/file/d/1XzLrpYzK_b6Rnu9LWd7Y6-pKc2CpE_75/view?usp=sharing
GNMDIV111 | USING_THE_REMOTE_CONTROL_UNIT_DURING_AQUISITION | patient postioning,general aquisition operation | using the remote control during aquisition | NM discover VH | How to use remote control during aquisition? | The keys required to perform acquisitions are the: <START/STOP> keys <PAUSE/CONT> key 1. During SPECT, Whole Body Scan, SYMA and GSPECT acquisitions, pressing any key on the RCU pauses (or stops) the movement of the Gantry and the acquisition procedure. 2. The gantry display unit (GDU) prompts to continue or terminate the acquisition. 3. Any movement that may cause collision between the Pallet and the roller is automatically inhibited. 4. The user has to depress the ENABLE lever located on the right hand side of the RCU to activate any key on the RCU 5. In this chapter, reference is made to the RCU function keys without mentioning the ENABLE lever each time. 6. Pressing the <PAUSE/CONT> button twice during Persistence, clears the image area.
GNMDIVH112 | PRESSING_ THE_PSD _DURING_AQUSITION | patient positioning,general aquisition operation,PSD | pressing the PSD during aquisition | NM discover VH | How to press PSD during aquisition ? | During ECT or Whole Body Scan acquisitions, pressing the PSD (Pressure Sensitive Device) will temporarily interrupt the acquisition. Proceed as follows: 1. The Gantry Display Unit prompts:PSD activated To release the PSD, in H-mode, press the RADIAL <OUT> key of the affected Head on the RCU. To release the PSD in L-mode, press the LATERAL <OUT> key on the RCU. 2. The gantry display unit (GDU) prompts:Press <STOP> to quit <CONTINUE> to resume To continue the acquisition, press the <PAUSE/CONT> key on the RCU. To terminate the acquisition, press the <START/STOP> key on the RCU or the [STOP] button on the Acquisition Control Panel. The data is stored on the disk and the gantry display unit (GDU) displays the message:Acquisition terminated
GNMDIVH113 | PATIENT_REMOVAL | patient positioning,general aguisition operation,patient removal | patient removal | NM discover VH | How to remove patient? | To remove the patient at the end of acquisition, proceed as follows: 1. Using the RCU <MODE> key scroll to the Release position. 2. Press the RCU <HOME> key. In response, the Detector Heads move OUT, the Stretcher moves OUT and the Table moves DOWN. 3. Assist the patient off the Table | https://drive.google.com/file/d/1hnzl5pM1YHbgJ0ClV9oCQBG7RASb3glQ/view?usp=sharing
GNMDIVH114 | SPECT_PROCEDURE | spect aquisition | spect procedure | NM discover VH | How is spect performed? | Accessing Protocol Modifying Protocol Values Moving the System to Home Position. Positioning Patient
GNMDIVH115 | WHOLE_BODY_SCAN_PROCEDURE | scan parameters | whole body scan procedure | NM discover VH | How is whole body scan performed ? | Accessing Protocol z Modifying Protocol Values z Positioning Patient z Moving the System to Home Position z Executing Protocol in Step & Shoot Mode z Executing Protocol in Continuous Mode
GNMDIVH116 | GATED_SYMA_PROCEDURE | acessing protocols,gated procedure | gated syma procedure | NM discover VH | How is gated syma procedure performed? | Accessing Protocol Modifying Protocol Values Executing Protocol
GNMDIVH117 | GSPECT_PROCEDURE | acessing protocols,gspect procedure | gspect procedure | NM discover VH | How is gspect perfrormed? | Accessing Protocol Modifying Protocol Values Moving the System to Home position Positioning Patient Executing Protocol
GNMDIVH118 | VOLUMETRIC_COINCIDENCE_AQUISITION_PROCEDURE | aquisition procedure,volumetric coincidence | volumetric coincidence aquisition procedure | NM discover VH | How is volumetric coincidence aquistion procedure performed? | 1. Patient Positioning and Clearance Test 2. Checking the emission count rate, using the Loss utility. 3. Acquiring the Study , including on-line generation of NM slices. 4. Rebinning of the acquired data into projections, which is done automatically, once the acquisition has been completed.
GNMDIVH19 | PATIENT_POSITIONING_FOR_TORSO_COINCIDENCE_IMAGING | volumetric coincidence imaging,patient positioning | patient positioning for torso coincidence imaging | NM discover VH | How is the patient positioned for torso coincidence imaging? | Retract the Patient Table completely. Select the COITORSO Home Position via the RCU <MODE> key.. Help the patient lie down. Press <HOME> key on the RCU to bring the Detector Heads to the required position.
GNMDIVH120 | AQUIRING_VOLUMETRIC_COINCIDENCE_STUDY | volumetric coincidence imaging,patient positioning | aquiring volumetric coincidence imaging study | NM discover VH | How is patient volumetric coincidence imaging aquired? | 1. On the Acquisition Station, double-click on the icon of the relevant diagnostic group (Neurology or Tumor and Inflammation); the list of subprotocols appears: For Neurology, the relevant protocols are: CoDe Brain – Narrow Field CoDe Brain – Narrow Field, Not Centered For Tumor, the only relevant protocol is: CoDe Torso – Narrow Field. 2. Double-click on the icon of the required acquisition protocol. The system enters PERSISTENCE mode as per standard SPECT acquisition. The RATE COUNTERS on the top of the images show the rate of true coincidence events and random events per Head. The RATE COUNTER within the Header displays the sum of the rates from both the Heads 3. Enter study and patient parameters in the opened acquisition notebook.4. Bring the Gantry and the Patient Table to the required home position by pressing the RCU <HOME> key.5. Start the acquisition from the RCU (Remote Control Unit) by pressing the ENABLE lever and the <START/STOP> key concurrently. As the acquisition progresses, the NM slices created by the on-line COSEM reconstruction appear on the screen and are continuously updated on the screen.6. Wait for the acquisition to complete. When the acquisition has been completed, pre-processing is invoked automatically. Re-binning takes several seconds, as indicated by the progress indicator displayed in a Message Window. Once the procedure has been completed, a Message Window appears displaying the following message: Data rebinned into group S <No.> G <No> DONE 7. Once the data pre-processing has been completed, a new acquisition using the same protocol and the same parameters can be initiated from the [NEXT Window] button on the control pane.
GNMDIVH121 | TRANSMISSION_EMISSION_AQUISITION_PROCEDURE | transmission emission imaging,aquisition procedure | transmission emission aquisition procedure | NM discover VH | How is transmission emmision imaging performed? | z If needed, mount the Septa shields onto the NM Detectors. z Help the patient onto the Table. . z Double click the desired TE icon from within the relevant clinical container, to display the TE Acquisition Notebook z Fill in Patient Information z Select the setup mode (Survey or Emission Persistence) z Select the Emission Protocol z Select the Transmission Protocol z Select the Emission Transmission order (Transmission First or Emission First) z Define patient positioning and orientation z When done, click the [START] button on the Acquisition Window Control Panel z The procedure starts by instructing the user to move the Gantry to Home position by pressing <HOME> on the RCU, and to press <START> to begin the first phase of the procedure (setup phase). From this point, the procedure progresses automatically according to the defined sequence and consists of the stages listed below. Operator intervention is required only in the setup phase and when confirming X-ray activation. All motion (rotation and table) is performed automatically. 1. Setup, either Emission or Survey 2. Performing the acquisitions: z Transmission Scan followed by Emission Scan OR z Emission Scan followed by Transmission Scan 3. Transmission data processing. The procedure varies slightly according to the defined sequence
GNMDIV122 | STARTING_XRAY_TRANSMISSION_SCAN | transmission emission imaging,aquisition procedure ,xray scan | starting xray transmission scan | NM discover VH | How is xray transmision scan performed (transmission followed by emission? | Starting the X-ray Transmission Scan In case of Emission Setup only: The message Press Start To Confirm Auto Motion, appears on the GDU and the message Radial Movement Press Start on the Gantry appears at the bottom of the Acquisition Control Panel. z Press <START/STOP> + <Enable> lever on the RCU. Then the message Radial Movement Please Wait appears in the Message window of the Acquisition Control Panel and the radials move to their outer limits. Then, or for Survey Setup: The following message appears in the Message window of the Acquisition Control Panel for both Emission Setup and Survey Setup. Table Movement. Please Wait... The Table moves to Transmission position (according to the ROI defined during Setup). When the table reaches Transmission position, the Transmission scan starts automatically. Wait for the Transmission acquisition parameters to appear on the Acquisition Window Control Panel The following acquisition parameters appear on the Acquisition Control Panel: Voltage Current CTDIw is the estimated X-ray dos. For more information. See CTDIw on page 18-3. Slice No. 0/X, where X is the total number of slices, calculated according to the ROI defined in the setup procedure. Time 0/X, where X is the total acquisition time in seconds, calculated by the system (X is usually slightly higher than actual acquisition time). z When the acquisition is ready to begin, the following confirmation window will appear: z Click [OK]. Starting the Emission Scan The following message appears on the Message window of the Acquisition Control Panel: Table Movement. Please wait... The Table moves to position (according to the ROI defined in Setup). The Emission Acquisition Window Control Panel is displayed. While the emission images are acquired an Acquiring sign is displayed in the Acquisition Window Control Panel. When acquisition stops, the Acquisition Window Control Panel shows Combined Acquisition completed. The post-acquisition processing begins automatically and takes up to 30 seconds. When this ends, a pop up window appears saying that the sequence has completed. Click [OK] to close the window.
GNMDIVH123 | STARTING_THE_EMISSION_SCAN | transmision emissin,emission scan,transmission scan | starting emission scan | NM discover VH | How is xray emission scan(emission followed by transmission) performed? | Starting the Emission Scan In case of Survey Setup only: The following message appears in the Message window: Table Movement. Please Wait... The Table moves to Emission position, according to the ROI defined during Setup. Then, or for Emission Setup: When the table reaches or is at the Emission position, the Emission scan is ready to begin automatically. While the emission images are acquiring an Acquiring... sign is displayed on the Acquisition Window Control Panel. When acquisition stops, the Acquisition Window Control Panel shows Continuing... Starting the X-ray Transmission Scan When the Emission scan is done, the Gantry will move the radials to their outer limit (if needed) and the bed to transmission position and the following message appears on the Message window of the Acquisition Control Panel: Table Movement. Please wait... z When the acquisition is ready to begin, the following confirmation window will appear:z Click [OK]. The Acquisition Control Panel appears on the screen
GNMDIVH124 | HAWKEYE_COINCIDENCE_AQUISITION_WORKFLOW | hawkeye coincidence,transmission emission image | hawkeye coincidence aquisition workflow | NM discover VH | What is hawkeye aquisition workflow? | 1. Verify that the Septa shields are mounted. 2. With the Stretcher is fully retracted, position the patient on the table with the area of interest within the Hawkeye scan limits (between the two labels), preferably with Legs OUT. If a Survey Setup is to be performed with the table moving IN and Survey length = 80 cm, verify that the area of interest is below 105 cm on the table. 3. Check the Loss rate by activating the LOSS utility from the CoDe Utilities container. . Verify that the Emission Single Rate is within 400 – 900 Kc/s. 4. Activate the desired TE protocol from the relevant container. 5. Enter Patient information. Refer to Entering Patient Information on page 6-17. 6. Select the desired Setup (default = Survey). 7. Select the desired X-ray protocol (recommended: XRAY Table OUT with matrix = 256). 8. Select the desired CoDe protocol (with FAM extension)be transferred to the predefined processing station.. 9. Select the scan order (default = Transmission first) 10. Select patient orientation (recommended Legs OUT) 11. Click [PERSIST] to incorporate protocol updates. 12. Press <HOME> on the RCU to move the system to TE Home Position. 13. Follow the system prompts. At the end of the acquisition sequence six groups are automatically created on the Acquisition Station: z XRAY SURVEY – Survey (planar X-ray) image; cannot be transferred to eNTEGRA. z XRAY SLICES – CT slices (optionally with raw data as extension); cannot be transferred to eNTEGRA z X-RAY for FUSION – This is one of the two input groups to the Fusion package. This includes a full cube of X-ray slices registered with a full cube of NM slices. z ATT MAP – This group contains the attenuation map to be used for the attenuation correction included in the Automatic COSEM/OSEM Processing package. z NM OTF IRAC or NM OTF IRNC – Emission Slices, created by COSEM iterative reconstruction (128 frames) with Attenuation Correction (AC) or without attenuation correction (NC), depending on the acquisition sequence, as follows:  Transmission first sequences results with attenuation corrected slices.  Emission first sequence results with non-attenuation corrected slices. z Fourier proj – Emission projections (64 frames). The last four groups should be transferred to the eNTEGRA workstation for further processing, by clicking the [TRANSFER] button. To transfer a TE study from the Local Catalog: 1. Open the Local Catalog and select the TE study to be transferred. 2. Drop-down the Selected menu and select the entry: Transfer TE Study to remote station. All the study groups except XRAY SURVEY (if any) and XRAY SLICES will be transferred to the predefined processing station.
GNMDIVH125 | SLICE_XRAY_AQUISITION_FLOW | xray aquisition,work flow | slice xray aquisition workflow | NM discover VH | How does slice xray aquisition work? | 1. Starting to rotate the Gantry. 2. Turning ON the X-rays according to the operator defined parameters. 3. Starting X-ray acquisition. 4. Continuing the acquisition until the Gantry has rotated through the specified rotation angle (216 or 360 degrees) 5. Stopping the slice acquisition 6. Continuing the Gantry rotation, while moving the table to next position (10 mm slice increment). 7. Starting to scan the next slice, repeating Steps 2 through 6, above. This acquisition will acquire transmission data that will be corrected and reconstructed to form the X-ray slices. Two new groups are generated from the X-ray slices: z X-ray Slices for Fusion z Attenuation Map use for attenuation correction
GNMDIVH126 | PERFORMING_DAILY_QC_IN_H_MODE | quality control,hmode, 57Co Flood Source | performing daily qc in h mode | NM discover VH | what are the steps to implement quality control in h mode using 57Co Flood Source? | 1. Prepare the source. The Count Rate should not exceed 20 Kc/s for either Detector Head. 2. Position the Detector Heads in H-Mode. 3. Use the RADIAL <IN> key to bring the Heads to the maximum IN position. 4. Place the Co57 flood source between the Heads with collimators mounted. 5. From the Maintenance container, double-click the Daily QC icon to activate the procedure; the system displays a Dialog Notebook. The first page titled DLQC PARAMETERS contains the settings of the parameters used in the procedure. 6. From the drop down list, for Isotope, select the appropriate isotope and press <ENTER> or click [OK].
GNMDIVH127 | PERFORMING_DAILY_QC_IN_L_MODE | quality control,l mode, 99mTc Point Source | performing daily qc in l mode | NM discover VH | What are the steps to implement quality control in l mode using 99mTc Point Source ? | 1. Prepare the source. The Count Rate should not exceed 20 Kc/s for either Detector Head. 2. Position the Detector Heads in L-Mode. 3. Place the Tc99m point source at a distance of 2.5 m from the mounted decoy collimator. 4. From the Maintenance container, double-click the Daily QC icon to activate the procedure; the system displays a Dialog Notebook. The first page titled DLQC PARAMETERS contains the settings of the parameters used in the procedure. 5. From the drop down list, for Isotope, select the appropriate isotope and Head number, and press <ENTER> or click [OK]. Do not change the settings of the following check box parameters: • STATIC Auto Start • ITER Auto Start • Show DATA • Show REPORT During the test all functions are implemented automatically in the following sequence: z PHA acquisition 300 Kcounts and GAUSSFIT of PHA histogram for the defined Head (1 or 2) z Flood STATIC (4 million counts) acquisition, simultaneously for the defined Head (1 or 2). z Flood image integral uniformity calculation for the defined Head (1 or 2). z A Daily QC Report is created and stored in the Archive in a Study labeled Daily QC. 6. Repeat the procedure for the second Head.
GNMDIVH128 | XRAY_DAILY_QC | quality controls,xray | xray daily qc | NM discover VH | What are the steps to implement xray daily qc? | z XRAY Warm-up – warming up of the X-Ray Tube by acquiring a few slices in low Voltage (120 KV). z Air Calibration – acquisition of a single full rotation slice, using the nominal X-Ray values. This slice is used to create the Air Table. z Survey Acquisition – used to define the Phantom slice for QC Analysis ition – used to check the quality of the X-ray image. This step requires the uniform Water Phantom. z Water Uniformity and Noise Analysis – used to check the quality of the X-ray image. This step requires the uniform water slice image.
GNMDIVH129 | XRAY_WARM_UP | quality controls,xray | xray warm up | NM discover VH | What are the steps to implement xray warm up? | Fully retract the Stretcher so that there is nothing between the X-ray Source and the Detector. z Double click the XRAY DLQC Acquisition icon in the XRay Utilities container (accessed from the Acquisition Station container). The warm-up X-ray acquisition parameters will be displayed on the screen: Voltage = 120 kV Current = 2.5 mA No. of Slices = 3 (full rotation of 1 rpm) z Press <HOME> + Enable lever on the RCU. z Press <START> + Enable lever on the RCU. z Immediately return to the operator console. The X-Ray confirmation pop-up is displayed. (the pop-up will disappear after two minutes) z Click [OK] to start the warm-up. The system will rotate for three revolutions with X-rays ON. This takes about three minutes.
GNMDIVH130 | AIR_CALLIBRATION | quality contols,xray | air callibration | NM discover VH | What are the steps to implement air callibration? | 1. Air Calibration Acquisition must be done when there is nothing between the X-ray Source and the Detector. Any object that disturbs the creation of the Air Table will result in artifacts in all X-ray images. 2. The Air Calibration must be acquired after warm-up of the X-ray tube. Air Calibration acquired with a cold tube will result in artifacts. z Take the table out and click [OK]. z Wait for Air calibration acquisition parameters to be displayed on the operator console: Voltage = 140 kV Current = 2.5 mA z When the message Ready for Start appears at the bottom of the acquisition dialog, press <START> + Enable lever on the RCU. z Immediately return to the operator console The X-Ray confirmation pop-up is displayed. z Click [OK] to start the air scan. When the new Air table processing is complete, the following message appears in the message window
GNMDIVH131 | ADDING_ISOTOPE | maintanece and setup,isotopes | adding isotopes | NM discover VH | How to add isotopes? | 1. Display the List of Isotopes by double-clicking the LISO icon residing within the Maintenance container. 2. From the Edit menu, select the Add Isotope entry. 3. In the opened Dialog, type the name of the new isotope and click [OK]. 4. In the opened dialog, click [YES] to confirm the name of the new isotope. 5. In the new opened dialog, type the name of a known isotope which is to be used as a basis for the definition of the new isotope. The Create (modify) Isotope Definition (CISO) Dialog Notebook opens on the ‘ISOTOPE’ page, with the name of the new isotope displayed in the current ISOTOPE NAME entry field . The other parameter fields such as number of peaks etc. have to be assigned appropriate values. 6. On the ISOTOPE page : z Click the appropriate radio-button to specify the number of peaks. z For each peak, enter the Energy Peak value, its location and its Half Life (HL) Time. z Click [OK] to cancel the incorrect settings. 7. On the ENERGY page : z Define Low %, High % by typing values in the entry fields. The Map No. drop down list box displays the maps available in the system. ‘255’ denotes that no map was defined. z Assign a map from the list by clicking on the required number; it appears in the entry field. 8. On the LINEARITY page , define Slow and Fast Linearity correction maps, as follows: z Click the arrow next to the relevant entry fields; the Map No. drop down list box displays the maps available in the system. z Click the desired number; the Map no. appears in the appropriate entry fields. ‘255’ indicates that no Map is defined 9. On the COLLIMATORS : z To add a collimator to the Defined list:  Click on the desired collimator from the list of Available collimators; the name appears in the entry field.  Click the [Add] button to include it in the Defined list of collimators. z To remove a collimator from the Defined list:  Click on the desired collimator in the Defined list; the name appears in the entry field.  Click the REMOVE button z To define Sensitivity Maps for a given collimator:  Click on the collimator name for which sensitivity maps are to be viewed or modified; the drop down list displays the Slow and Fast Rate sensitivity maps for this collimator (for each Head and Peak). The values may be modified.  Click on the desired Sensitivity Map numbers; the selected numbers appear in the appropriate entry fields. 10. On the LOCATIONS z Add the scatter labels (SC) to identify the scatter windows
GNMDIVH132 | DELETING_ISOTOPE | maintanece and setup,isotopes | deleting isotopes | NM discover VH | How to delete isotopes? | 1. Display the List of Isotopes by double-clicking the LISO icon residing within the Maintenance container. 2. Select the isotope(s) to be deleted. To select multiple isotopes, hold down the <Ctrl> key and click on the isotopes. 3. From the Edit menu, select the Delete Isotope entry; the name(s) of the selected isotope(s) appears and the system prompts: Press OK to delete the following isotopes OK Cancel
GCTREV133 | SET_UP_A_PEDIATRIC_OR_SMALL_PATIENT_EXAM | pediatric,small patient scans | set up a pediatric or small patient scan | NM discover VH | How to set up a pediatric or small patient scan? | 1. From the scan monitor, click the Plus icon , in the upper left corner. ○ The Patient Information screen displays the new Exam Number. ○ The maximum Exam Number is 49,999, which is reset by your Field Engineer. 2. From the Patient Information area, type data into the appropriate fields. ○ Press <Tab> to advance to the next field. Alternatively, use the mouse to navigate to each field. ○ Patient ID is a required field. If the patient does not have an identification number, type ? or the word trauma. 3. Click [Pediatric]. 4. From the Pediatric Protocol Selection screen, click an anatomical area. ○ With the patient's weight entered, the system automatically selects the appropriate pediatric color code area for the anatomical area selected. ○ Use the arrows to scroll through the list. ○ Click a protocol from the list to download the scan parameter values. 5. From the Pediatric Protocol Category screen, select the weight/length category based on the size of your patient or verify the correct Color category has been selected if you entered a weight. ○ The protocols in the selected color code are then displayed accordingly. ○ The default weight/color selector displays the patient weight entered on the Patient Information screen, or the last weight/color selection (if no patient information was entered). ○ If you enter a patient weight on the Patient Information screen and select a color/weight that is not consistent with the entered information, a red alert displays next to the protocol name. You must acknowledge that you have chosen a protocol that does NOT match the patient size. Alert message: Selected color is not suitable for the entered patient weight of 29 Kg (65 lbs). 6. Proceed with the Acquire a Scout scan procedure
GCTREV134 | SET_A_DOSE_CHECK_ADMINISTRATOR_ROLE | patient care | set a dose check administrator role | revolution ct | How to set the dose check administrator role ? | 1. From the image monitor, click . 2. From the Mode menu, click Access Controls, then User Admin Tool. 3. From the EA3 Administration window, enter your username and password. NOTE: The user should be someone with administrative privileges. 4. Click the Local Users tab. 5. From the Local Users tab, click [Add Local User]. 6. From the Add User window, enter the following: ○ User ID ○ Full Name ○ Password ○ Confirm Password NOTE: The User ID must be unique from others used on the system. 7. Click [Add User]. 8. Click the Groups tab. 9. From the Groups tab, click [Add Local Group]. NOTE: Alternately, the system can be configured for enterprise groups if enterprise authentication has been enabled. Refer to the Access Controls chapter for information on configuring enterprise authentication. 10. From the Add Local Group window, type a unique group name. 11. Click [Add Group]. ○ The group is highlighted in the Local Groups list. ○ All information and buttons in the center panel refer to the highlighted group. 12. From the Roles area, select Dose Check Administrator and click [Apply Roles]. ○ A green label confirms the applied roles. ○ An error message displays if it is unsuccessful.
GCTREV135 | SET_THE_DOSE_CHECK_AV_EXCEEDING_USER_ROLE | patient care | Set the Dose Check AV Exceeding User Role | revolution ct | How to set the dose check av exceeding user role? | 1. From the image monitor, click . 2. From the Mode menu, click Access Controls, then User Admin Tool. 3. From the EA3 Administration window, enter your username and password NOTE: The user should be someone with administrative privileges. 4. Click the Local Users tab. 5. From the Local Users tab, click [Add Local User]. 6. From the Add User window, enter the following: ○ User ID ○ Full Name ○ Password ○ Confirm Password NOTE: The User ID must be unique from others used on the system. 7. Click [Add User]. 8. Click the Groups tab. 9. From the Groups tab, click [Add Local Group]. NOTE: Alternately, the system can be configured for enterprise groups if enterprise authentication has been enabled. Refer to the Access Controls chapter for information on configuring enterprise authentication. 10. From the Add Local Group window, type a unique group name. 11. Click [Add Group]. ○ The group is highlighted in the Local Groups list. ○ All information and buttons in the center panel refer to the highlighted group. 12. From the Roles area, select Dose Check AV Exceeding User and click [Apply Roles]. ○ A green label confirms the applied roles. ○ An error message displays if it is unsuccessful
GCTREV136 | CONFIGURE_THE _SYSTEM_FOR_DOSE_CHECK | patient care | configure the system for dose check | revolution ct | How to configure the system for dose check? | 1. From the image monitor, click . 2. From the Mode menu, click Access Controls, then Dose Check Management. 3. From the Dose Check Management window, in the NV (Notification Value) Checking area, select CTDIvol and/or DLP.4. From the Dose Check Management window, in the AV (Alert Value) Checking area, select CTDIvol and/or DLP. NOTE: The Alert Value (AV) should be determined by the site’s radiologist and physicist. 5. Enter a value for CTDIvol and/or DLP. ○ The maximum AV value for CTDIvol is 2,000 mGy.○ The maximum AV value for DLP is 400,000 mGy-cm. ○ The system default is AV checking by CTDIvolOn with a value of 1,000 mGy and/or DLPOff with a value of 0 mGy-cm. 6. Click [Save] | https://drive.google.com/file/d/13kpJCQvDIXXEb9WKB8zUEPrLBzEOPvL_/view?usp=sharing
GCTREV137 | CONFIGURE_THE_AV_BY_AGE_THRESHOLD | patient care | configure the av by age threshold | revolution ct | How to configure the av by age threshold? | 1. From the image monitor, click . 2. From the Mode menu, click Access Controls, then Dose Check Management. 3. From the Dose Check Management window, in the AV (Alert Value) Checking area, select Age Threshold and enter the patient’s age. Illustration 2: AV (Alert Value) 4. Enter a value for CTDIvol Alert (mGy) and/or DLP (mGy-cm)○ The maximum AV (Alert Value) for CTDIvol is 2,000 mGy. ○ The maximum AV (Alert Value) for DLP is 400,000 mGy-cm. ○ The system default is AV checking by CTDIvolOn with a value of 1,000 mGy and/or DLPOff with a value of 0 mGy-cm. ○ The value must be lower than the adult AV (Alert Value). 5. Click [Save]. | https://drive.google.com/file/d/1mVtjNBkMx0P6oITqKkuHRKf3AQ2G0Unz/view?usp=sharing
GCTREV138 | SCAN_USING_ALERT_VALUE_CHECKING | patient care | scan using alert value checking | revolution ct | How to scan using alert value checking? | Adjust the scan parameters for the series until the exceeded AV value is below the prescribed AV value. If the AV value is still exceeded after you click [Confirm], a confirmation message will display. The scan cannot be confirmed until the AV value is no longer exceeded, or a user with a Dose Check AV Exceeding user role has authorized the scan. Performing a scan exceeding an AV value demands a high degree of consideration for appropriateness. In most cases, you will click [Cancel] to go back to adjust parameters. If there is a true diagnostic need to perform a scan exceeding an AV value, enter the diagnostic reason. 1. Type the Logon Name. 2. Type the Password. 3. In Diagnostic Reason field, enter the reason for exceeding the AV value. 4. Click [Confirm] to proceed to scanning | https://drive.google.com/file/d/13fncIQsRD9iDuf-pnyu5wBIRcKmyZHhc/view?usp=sharing
GCTREV139 | SCAN_USING_NOTIFICATION_VALUE_CHECKING | patient care | scan using notification value checking | revolution ct | How to scan using notification value checking? | Adjust the scan parameters for the scan group until the exceeded NV value is below the prescribed NV value. If the NV value is still exceeded after you click [Confirm], a confirmation message will display. The scan cannot be confirmed until the NV value is no longer exceeded, or you have entered a diagnostic reason for continuing the scan with the exceeded NV value. 1. In Diagnostic Reason field, enter the reason for exceeding the NV value. 2. Click [Confirm] to proceed to scanning | https://drive.google.com/file/d/1u6KHeftUd6AS2lE21joN8R8ctItDdhJb/view?usp=sharing
GCTREV140 | PREPARE_THE_SYSTEM | patient care,hardware | prepare the system | revolution ct | How to prepare the system for scan? | • Clean the accessories and check for damage. • Check and remove dried contrast agent: ○ Scan window (around the gantry opening) ○ Table extension and cradle surfaces — especially the patient restraint plastic channels on the table ○ Accessories (head holders, pads and cushions, et cetera) • Check supplies.
GCTREV141 | LOGGING_ON_AND_OFF | patient care,hardware | logging on and off the system | revolution ct | How to log on and off the system? | Log On If Access Controls is enabled, use the following procedure to log on to the system. 1. From the Logon window, enter your logon name. 2. Enter your password.3. Click [Logon]. Log Off If Access Controls is enabled, use the following procedure to log off from the system. 1. From the image monitor, click . 2. From the Mode menu, click Shut Down, then Switch User
GCTREV142 | TUBE_WARMUP | patient care,hardware | tube warmup | revolution ct | How to warm up the tube? | 1. Clear the gantry area of all objects. Any obstruction in the gantry may lead to artifacts in scanned images. 2. From the image monitor, click the Utilities icon . 3. From the Utilities menu, click Daily Prep, then Tube Warmup. The Service-Calibration Tools window opens. 4. Select Tube Warmup from the drop-down and click [Run]. 5. Alternatively, click , then [Tube Warmup]. NOTE: Refer to the Safety chapter regarding using non-GE tubes. 6. Press <Start Scan>. 7. The system returns to the Service-Calibration Tools window when the tube warmup is complete. Click [Dismiss] to exit the window.
GCTREV143 | ADD_A_PATIENT_TO_SCHEDULE | patient care,patient schedule | add a patient schedule | revolution ct | How to add a patient schedule? | 1. Click the Plus icon in the upper left of the left monitor. , or From the scan monitor, click the Patient Schedule drawer. Then select Add Patient from the Schedule dropdown menu in the top right corner of the window. 2. Complete all relevant fields on the Patient Information window If you are using Exam Split with Hard Exam Split mode to add a patient, you must enter an exam description. 3. From the pull down, click Save to Schedule to add the patient to the schedule or Cancel to close the window without adding the patient.
GCTREV144 | VIEW _MORE_PATIENT_INFORMATION_ON_THE_SCHEDULE | patient care,patient schedule | view more patient information on the schedule | revolution ct | How to view more patient information on the schedule? | 1. From the scan monitor, click the Patient Schedule drawer. 2. From the Schedule window, click the Information icon for the patient. The screen is populated from your HIS/RIS. It also displays information entered locally. If there is a caution for the patient, you will see the [Caution Info] icon . 3. Click [Close] to close the Patient Information window.
GCTREV145 | SELECT_A_PATIENT_FROM_THE_SCHEDULE | patient care, patient schedule | select a patient from the schedule | revolution ct | How to select a patient from the schedule? | 1. Click anywhere on the Patient Schedule drawer. 2. From the Patient schedule list, select the desired patient and click [Select Patient]. The Patient Information screen fields are populated. ○ Multiple records for the same patient can be selected. ○ Multiple records must be selected in order to use Exam Split. 3. Select desired protocol and then click [Accept]. The system displays the scan series tasks in the protocol.
GCTREV146 | PATIENT_STATUS | patient care,patient schedule | patient status | revolution ct | How to check patient status? | 1. From the scan monitor, click the Patient Schedule drawer. 2. View the Status column. ○ N – New Record or Not Completed ○ C – Completed 3. Click and drag the scroll bar to browse through the list.
GCTREV147 | SET_UP_AND_PATIENT_POSITIONING | patient care,patient positioning | setup and patient positioning | revolution ct | how to set up and position the patient? | 1. Attach the head holder or foot extender to the end of the cradle toward the gantry, if needed. 2. Press and hold the Unload table foot pedal or use the gantry controls to adjust the table height for transferring the patient onto the table. The cradle moves to home and to a minimum height of 49 cm. 3. Transfer the patient onto the table. ○ Make sure to follow the manufacturer’s directions for use of a protective table slicker when scanning, otherwise artifacts may appear in images. ○ Blocked reference channels can cause image artifacts. Make sure that there are no blankets, clothing, tubing, or straps hanging down below the table when scanning. Keep items wrapped close to the table by using the patient restraint straps. 4. Press and hold the Load table foot pedal. ○ The table moves up to a height of 210 mm and advances the table and cradle into the gantry. If a protocol has been selected, the table will move to a preset height and location based on the anatomical reference in the protocol○ Release the foot pedal when the patient is properly positioned. Use the gantry controls to further adjust the position. 5. Make the patient comfortable and immobilize as needed. Connect body straps to the cradle and wrap them around the patient. ○ Use the head holder straps to immobilize the patient's head. ○ Make sure items such as blankets and straps are not hanging off the table or allowed to touch or be dragged through the gantry during the scan, otherwise artifacts may appear in the images. 6. If needed, attach the IV pole. 7. Use the following gantry controls to refine the patient's position: Table Up Table Down Cradle In Cradle Out 8. From the gantry controls, press the <Laser Alignment Light> button to turn on the two laser lights: internal and external (axial, sagittal, and coronal). Align the laser lights with the target anatomical reference. Do not position the patient with the laser lights shining into his eyes. 9. From the gantry controls, press the <Internal Landmark> button or the <External Landmark> button. You can not click [Confirm Settings] until after you have pressed a landmark button. ○ The landmark sets a known anatomical reference for the radiologist to correlate anatomy. ○ The landmark sets the zero location. When scanning towards the patient’s head, you are scanning superior to the zero location. When scanning towards the patient’s feet, you are scanning inferior to the zero location. You should set the zero location to known anatomy. For example, when scanning a head, the landmark or zero location is typically the orbital meatal (OM) line. ○ A landmark or zero reference point is required before you scan a patient. It can be set before or after you start a new patient exam. ○ When using the external laser alignment light to position the patient, be aware that the patient’s elevation may be slightly lower with the cradle extended than with the cradle fully retracted. The cradle may bend slightly under the patient’s weight. Take this difference into consideration for applications where patient position is critical, such as treatment planning. To minimize these effects, after positioning the patient using the external laser alignment, advance the patient to the CT scan plane. Turn on the CT alignment lights to determine if they line up with the markers on the patient. Compensate for the bend in the cradle by elevating the table. When the CT alignment lights line up with the markers, reset the landmark for the scan using the Internal laser alignment light. 10. From the Scan Control Interface, set volumes with the <Patient-to-Operator> and <Operator-to-Patient >, or <AutoVoice> communication controls. Press the <Talk> control to test the <Operator-to-Patient > volume level.
GCTREV148 | START_THE_SCAN | pateint care,patient status | start the scan | revolution ct | How to start the scan? | Use this procedure to start a scan. 1. Review the following parameters for each acquisition before confirming the scan: a. Start/End locations b. Direction of scan c. Projected Noise Index and mA table for SmartmA and SmartmA/ODM 2. Click [Confirm Settings]. 3. Press <Move to Scan> 4. Deliver breathing and table move instructions, as needed. 5. Press <Start Scan>. Considerations • Press <Stop Scan> to stop scans between multiple helical groups and do not press <Stop Move>. • Scan aborts may occur during Scout, Axial or Helical scanning. Always be aware of the scan progress during an exam and select [Resume] to continue. • The scan may fail to confirm, posting a message that not enough image space exists, even though the disk image space icon indicates there is enough space or File Manager may not have exams listed. This occurs because images are stored on the system disk in more than one partition. Remove consecutive exams to free up image space to confirm and proceed with the scan.
GCTREV149 | STOP_A_SCAN | patient care,patient status | stop the scan | revolution ct | How to stop the scan? | Abort a scan in progress Use this procedure to abort the X-ray and stop gantry and table movement. 1. On the scan control interface, press <Stop Scan>. 2. Click <Resume> from the Scan Progress window to resume the scan. Pause a scan in progress Use this procedure to pause the current scan. This finishes the current scan, then pauses the next scan. 1. On the scan control interface, press <Pause Scan>. 2. Click [Resume] from the Scan Progress window to resume the scan. Emergency stop scan, electronics, and x-radiation Use this procedure in the event of a patient related emergency or if the cradle, table, or gantry starts to move unexpectedly. 1. On the scan control interface or gantry, press <Emergency Stop>. All table and gantry motions and X-ray exposures are stopped. 2. On the gantry, press <Reset> to clear the Emergency Stop.
GCTREV150 | SCAN_PARAMETER_WORKFLOW | patient care,patient status | scan parameter workflow | revolution ct | What is the scan parameter workflow? | 1. From the Scan Settings window, select a scan task from the protocol. NOTE: You must execute each scan task in order. Click [Confirm Settings] for each scan task, and then scan that scan series. You may then click [Continue] to move to the next scan task. Scan tasks can be deleted if you want to skip them. Scan tasks can also be reviewed when completed. 2. Make sure the patient orientation is the same as it was for the scout. If necessary, click the patient image in the Patient Position area to change the patient orientation and position. NOTE: There may be times when you change the patient’s position from the way it was built in the protocol. In this case, you will have to change the patient’s orientation each time you advance to the next series unless you use the Copy Forward feature. 3. Choose the Scan Type. ○ For an Axial or Helical scan, see the Technical Reference Manual; System Operational Modes section for detailed information about the scan types. ○ For a Hi Res scan, see Acquire a Hi Res scan. ○ For a Cardiac scan, see Select a Cardiac scan. 4. Set the Coverage Speed options. 5. Set the Primary Recon options. See Recon Parameters for details. 6. Set kV and mA Control. ○ Set the mA. ○ Set the kV. 7. Set the Anatomy Selection. a. Set the Start and End locations. ■ To graphically set the locations, see Adjust the Graphic Rx. ■ To explicitly set the locations, see Set the Start and End locations. b. Set the SFOV. c. Set the DFOV. d. Set the A/P center and R/L center. 8. Set the Contrast options. Scan Parameters a. Select a contrast mode. b. Select contrast presets. 9. Set the Timing parameters. 10. Set the Gating and Trace options. 11. Set the Secondary Recon options.
GCTREV151 | SCAN_TYPE | pateint care,patient status | scan type | revolution ct | How to choose scan type? | 1. From the Scan Settings window, click Scan Type. 2. From the Scan Type collection, choose the desired scan type. ○ Click Axial for a step and shoot method. ■ Routine head studies typically use Axial scans acquired with a 1 second rotation time. ■ Hi Res chest studies may use Axial scans acquired with a sub-millimeter (0.625 mm) slice. ○ Click Helical for a continuous table movement method. ■ A wide range of exam types use Helical scans. ■ To stop a scan between multiple Helical groups, press <Stop Scan> not <Stop Move>. ○ Click Cine when one location needs to be scanned over a period of time, such as for hemangioma studies. Full is continuous exposure that supports table movement equal to the beam collimation or no table movement, where the scan is taken at one table position. ○ Click Cardiac to help freeze motion of the heart and vascular structures with a Helicalgated or Cine-gated scan. If the Scan Type is changed from another scan type, the mA value automatically updates. Review it to be sure it is acceptable. 3. For Hi Res Mode, click On in combination with the Axial, Helical, or Cardiac scan type to produce images with improved resolution in the x and y directions. ○ Hi Res Mode is supported to enable collection of more views in the x and y directions to support increased resolution using HD algorithms or improvement in IQ as a result of having more views, such as for off-center DFOV imaging. ○ The standard set of reconstruction algorithms are compatible with Hi Res acquisitions, however, the inherent resolution will be reduced to match non Hi Res acquisitions. ○ The Hi Res Modes supported include the following: ■ Axial – 40, 80, 120, 140, 160 mm detector coverage ■ Helical 0.516:1 pitch 40 mm detector coverage 0.984:1 pitch 40 mm detector coverage 0.992:1 pitch 80 mm detector coverage ■ Cardiac Axial – 40, 80, 120, 140, 160 mm detector coverage 4. Click a Rotation Time. ○ Axial – 0.28 s, 0.35 s, 0.5 s, 1.0 s ○ Cine – 0.28 s, 0.35 s, 0.5 s, 1.0 s ○ Helical – 0.28 s, 0.35 s, 0.5 s, 1.0 s ○ Cardiac Axial – 0.28 s, 0.35 s 5. Click [Save].
GCTREV152 | COVERAGE_SPEED_OPTIONS | patient care,patient status | coverage options | revolution ct | What are the coverage options and how to set them? | Use this procedure to set the coverage speed options for detector coverage, image slice thickness, pitch, table speed, and rotation speed. 1. From the Scan Settings window, click Coverage Speed. 2. Choose the Detector Coverage. 3. For a Helical scan type, choose the Pitch. NOTE: Total Exposure Time is updated automatically when a pitch is selected. ○ The pitch is the table travel in millimeters per rotation divided by the detector coverage. ■ 4 Row Interleaved provides a 40% mAs reduction, is 1.5 to 3 times faster than single slice helical and has minimal helical artifacts, but interleaved pitch 0.5:1 provides only one third of the coverage compared to interspaced mode. ■ 4 Row Interspaced provides data acquisition 2 to 6 times faster than single slice helical, but requires more interpolation, has more helical artifacts, and only a 20% mAs reduction. ○ Helical pitches include: 0.516:1, 0.984:1, 0.992:1 and 1.375:1. 4. For Axial, Cine and Cardiac scan types, the detector coverage and slice thickness determine the images per rotation. ○ The choice for one determines the choices for the other. Some combinations are not allowed. ○ The choice of detector coverage and slice thickness determines the images per rotation. Some combinations are not allowed. 5. For an Axial or Cine scan type, the interval is set equal to the detector coverage. The interval value can be increased or decreased. If the interval is greater than the detector coverage, the gap is between each rotation of data, not between the slices. For an Axial scan type, the interval or spacing defaults to the equal number of images per rotations multiplied by the slice thickness. ○ The interval for Axial scanning can be zero, equal to, or greater than the width of the detector configuration. ○ Axial interval with skip refers to a gap between scan groups. This can be useful, for example, when performing a survey exam, such as a high resolution chest exam. 6. Note the Total Exposure Time. 7. Click [Apply].
GCTREV153 | SET_THE_MA | patient care,patient status | set the mA | revolution ct | How to set the mA? | Use this procedure to set the mA for the scan prescription. 1. From the Scan Settings window, click kV and mA Control. 2. Choose an mA Mode. ○ Set to Manual mA and enter an mA value. ○ Set to SmartmA or SmartmA and ODM to enable mA modulation in the x-y and zdirections. ■ Select a Noise Index value that provides acceptable diagnostic quality. As the Noise Index increases the required mA decreases and image noise increases. The optimum strategy is to find the highest Noise Index sufficient for the clinical task and let SmartmA select the mA without using significant constraints. Use a slightly lower Noise Index for small patients and a slightly higher Noise Index for large patients. ■ Enter Min mA and Max mA values, which constrain the mA range used during the scan. ○ For cardiac scans, see guidelines for AEC with Cardiac Modes. 3. Click [Apply]. 4. Review: For SmartmA and ODM scans, review if the scan is meeting the prescribed Noise Index from the Image Quality widget to view the calculated X and Y, or A, P, R and L values. If not, click [Adjust] to be presented with a list of solutions with dose changes to realize the noise versus dose in the prescribed Noise Index. ○ This can only be done during scan set up and not in Protocol Management. ○ The system will display an indication when one or more rotations are at the Min mA or Max mA value specified for SmartmA. If the system indicates this condition, check the scan parameters in the mA Table to ensure the acquisition will provide the expected results.
GCTREV154 | SET_THE_KV | patient care,patient status | set the kv | revolution ct | How to set the kv? | Manual kV Use this procedure to manually prescribe the kV. 1. From the Scan Settings window, click kV and mA Control. 2. Click kV Mode and select Manual kV. 3. Select a kV value. Typical applications are listed below. ○ Click 70 kV for pediatric imaging. ○ Click 80 kV for Bone Mineral Densitomery and Perfusion imaging in the brain and pediatric imaging. ○ Click 100 kV for pediatric and small patient imaging. ○ Click 120 kV for routine imaging of the chest, abdomen, and pelvis areas. ○ Click 140 kV for imaging the posterior fossa, thick areas, and heavy patients. 4. Click [Apply]. kV Assist Use this procedure to prescribe a scan with kV Assist. See the Protocols chapter for information about kV Assist. 1. Select a protocol and verify the Manual kV, mA Mode/SmartmA (Noise Index minimum/ maximum mA) or Manual mA, Slice Thickness, Rotation Time, Pitch, SFOV, and WW/WL values are acceptable.2. From the kV and mA Control collection, set the kV Mode to kV Assist. NOTE: Auto Adjust will be “On” in the Primary Recon collection by default for all recons to optimize the protocol’s reference WW and WL based on the expected change in image contrast and noise when kV Assist is enabled. Select “On” to use the WW and WL values set in the Primary Recon. To change the WW and WL values, kV mode must be set to Manual kV. 3. Specify the Clinical Task for the study. 4. Review the resulting kV, Noise Index or Manual mA and WW/WL values to ensure they are acceptable. NOTE: Manual mA or Smart mA parameters cannot be changed when kV Assist is on. Be sure to set these settings prior to turning on kV Assist. Secondary recons will not automatically adjust WW and WL.
GCTREV155 | SET_THE_START_AND_END_LOCATIONS | anatomy selection,patient care | set the start and end locations | revolution ct | How to start and end locations? | 1. From the Scan Settings window, click Anatomy Selection. 2. For a Helical scan type, click Start. For Axial, Cine or Cardiac scan types, click Full FOV Start 3. Type a beginning point. ○ The numbers must be preceded by the correct designation of S for superior (toward the head) from the centering point, or I for inferior (toward the feet). ○ Plus + can be used for superior and minus - for inferior. 4. For a Helical scan type, click End. For Axial, Cine or Cardiac scan types, click Full FOV End 5. Type an end point. ○ The numbers must be preceded by the correct designation of S for superior (toward the head) from the centering point, or I for inferior (toward the feet). ○ Plus + can be used for superior and minus - for inferior.
GCTREV156 | SET_THE_SCAN_FIELD_OF_VIEW | anatomy selection,patient care | set the scan field of view | revolution ct | How to set the scan field of view? | 1. From the Scan Settings window, click Anatomy Selection. 2. Choose an SFOV. ○ Click Ped Head to image infants 18 months or less in age. It allows you to enter up to a 32 cm DFOV and is limited to 42 kW. It has a special processing to correct for beam hardening effects. ○ Click Ped Body to image to image infants 18 months or less in age. It allows you to enter up to a 32 cm DFOV and is limited to 42 kW. ○ Click Small Head to image small heads up to 32 cm. It has a special processing to correct for beam hardening effects. ○ Click Head to image adult heads up to 32 cm. It has a special processing to correct for beam hardening effects. ○ Click Small Body to image small bodies, extremities that are centered in the middle of the gantry, and the QA phantom. It allows you to enter up to a 32 cm DFOV. ○ Click Medium Body to image most medium body work. If you measure the anatomy that you are scanning and it measures over 30 cm, you need to use Medium Body. It allows you to enter up to a 36 cm DFOV. ○ Click Large Body to image most adult body work. If you measure the anatomy that you are scanning and it measures over 30 cm, you need to use Large Body. It allows you to enter up to a 50 cm DFOV. ○ Click Cardiac Small, Cardiac Medium or Cardiac Large for cardiac studies.
GCTREV157 | SET_THE_DISPLAY_FIELD_OF_VIEW | anatomy selection,patient care | set the display field of view | revolution ct | How to set the display field of view? | 1. From the Scan Settings window, click the Anatomy Selection collection. 2. Click DFOV. 3. Enter the value of the patient's measurements. ○ Measure the patient at the widest point and add 2 cm. This will show all the anatomy and the soft tissue around it. ○ If an R/L value or A/P value is more than half of the DFOV, then the R/L, A/P image annotation does not display on the images. For example, if the DFOV is 10 cm and the R value is 56 mm, the image annotation displays R. Use an R/L value or A/P value that is less than half of the DFOV in mm to avoid this. ○ The minimum DFOV is 5.0 cm. ○ The maximum DFOV depends on the selected SFOV
GCTREV158 | SET_THE_R/L_CENTRE_COORDINATES_FOR_PRIMARY_RECON | anatomy selection,patient care | Set the R/L Center coordinates for the primary recon | revolution ct | How to set the r/l centre coordinates for primary recon? | 1. From the Scan Settings screen, click the Anatomy Selection collection. 2. Click R/L Center. 3. To find the coordinates, place the mouse over the Show Localizer Anterior/Posterior (AP) scout image and look at the Right (R) and Left (L) readout at the bottom of the image. 4. Type an R or L prefix and the R/L center coordinates, in millimeters. ○ The range of values can be from 0 to one-half the SFOV (example: Head SFOV is 32 cm so the maximum offset R/L is 16 cm or 160 mm). Typically, you would not want the offset to exceed one-half the DFOV or the resulting image does not show a right or left marker; it does show markers as R-R or L-L. ○ Entering a value other than zero off centers the image in the right and left axes of the patient. ○ Plus <+> can be used for R values and Minus <-> for L values for faster data entry using the ten key pad.
GCTREV159 | SET_THE_A/P_CENTRE_COORDINATES_FOR_PRIMARY_RECON | anatomy selection,patient care | Set the A/P Center coordinates for the primary recon | revolution ct | How to set the a/p centre coordinates for primary recon? | 1. From the Scan Settings screen, click the Anatomy Selection collection. 2. Click A/P Center. 3. To find the coordinates, place the mouse over the Show Localizer lateral scout image and look at the Anterior (A) or Posterior (P) readout at the bottom of the image. 4. Type an A or P prefix and the A/P Center coordinates, in millimeters. ○ The maximum offset for A/P Center is one half the SFOV from isocenter selected (example: Head SFOV is 32 cm so the maximum offset A/P is 16 cm or 160 mm). ○ Entering a value other than zero off centers the image in the anterior and posterior axes of the patient. ○ Plus <+> can be used for Anterior (A) values and Minus <-> for Posterior (P) values for faster data entry using the ten key pad
GCTREV160 | SET_UP_PREP_DELAY_TIME | anatomy selection,patient care,time parameters | set up prep delay time | revolution ct | How to set up prep delay time? | 1. From the Scan Settings window, click Timing. 2. From the Timing collection, click Prep/Group Delay. 3. Type a scan delay time (in seconds).
GCTREV161 | SET_INTER_SCAN_DELAY_TIME | scan parameters,anatomy selection,patient care,time parameters | set inter scan delay time | revolution ct | How to set inter scan delay time? | 1. From the Scan Settings window, click Timing. 2. From the Timing collection, click Inter Scan Delay. 3. Type a scan delay time (in seconds) between each Axial scan. ○ The value is dependant on the detector coverage selected when One or More is selected. ○ When the scan interval is zero, the value is set by the system based on the selected Rotation Time and Minimum Time Between Passes. ○ The valid range is 1 to 300 seconds. ○ Typically, the ISD is set at 1 to 1.7 seconds so the exam is done as quickly as possible.
GCTREV162 | SLICE_THICKNESS | scan parameters,anatomy selection,patient care,primary recon parameters | slice thickness | revolution ct | How to set the slice thickness? | 1. From the Scan Settings window, click Primary Recon. 2. Choose the Thickness. ○ The smaller the slice thickness, the more technique you need to use. The thinner the slice thickness, the higher the noise will be in the slice. Adjust protocol factors to provide desired image quality and noise level for the acquisition. ○ A thinner slice provides better detail. However, contrast resolution will be decreased.
GCTREV163 | IMAGE _INTERVAL | scan parameters,anatomy selection,patient care,primary recon parameters | image interval | revolution ct | How to set the image interval? | 1. Click Interval. 2. Type an interval value in mm. For a Helical scan type, the table moves in mm per rotation, while exposing the patient. ○ The image interval defaults to equal the slice thickness (for primary reconstructions only). ○ The maximum interval is twice the slice thickness. ○ The interval can be less than or greater than the slice thickness. ○ IQ Enhance is compatible with 0.625 mm and 1.25 mm slice thicknesses. Interval for acquisitions with IQ Enhance must be equal to the slice thickness (0.625 mm or 1.25 mm) or an overlap of 50% (0.312 mm or 0.625 mm).
GCTREV164 | RECON_TYPE | scan parameters,anatomy selection,patient care,primary recon parameters | recon type | revolution ct | How to set the recon type? | 1. From the Scan Settings window, click the Primary Recon collection. 2. Click Recon Type. 3. Select the appropriate algorithm for the primary or first reconstruction. The algorithms from left to right increase spatial resolution and decrease low contrast detectability. ○ Click Soft for tissues with similar densities, but not useful for non-enhanced scans. ○ Click Stnd for routine exams, e.g., chest, abdomens, and pelvis scans. ○ Click Detail for post myelograms, where hybrid tissue detail and bone edges are important. ○ Click Lung for interstitial lung pathology. ○ Click Bone for high resolution exams and sharp bone detail. ○ Click Bone Plus for sub mm detailed head work. It can be used for any study that normally used the bone algorithm, but is very useful in cases where the Edge algorithm was used. This is because the Bone Plus algorithm has no reconstruction penalty and is very close in standard deviation to Edge. ○ Click Edge for small bone work in the head, as well as high resolution scans. ○ Click Chest for mediastinum and lung detail studies. It provides soft tissue resolution and contrast when viewing the images in a soft tissue/mediastinal W/L and high resolution of the lung tissue when viewing the images in a lung W/L.4. If Hi Res mode is enabled, select a higher resolution algorithm that has similar characteristics to the non-Hi Res algorithms. ○ HD Stnd ○ HD Lung ○ HD Detail ○ HD Bone ○ HD Bone Plus ○ HD Edge ○ HD Ultra 5. If Cardiac Hi Res mode is enabled, select a higher resolution cardiac algorithm that has the same characteristics as the non-Hi Res cardiac algorithms. ○ HD Soft ○ HD Stnd ○ HD Stnd Plus ○ HD Detail ○ HD Detail Plus ○ HD Lung ○ HD Edge
GCTREV165 | AUTO_VOICE_LANGUAGE | auto voice,recon parameters | auto voice language | revolution ct | How to set auto voice language? | Use this procedure to set the language for Auto Voice. 1. From the Scan Settings window, in the Patient Position area, click Auto Voice. 2. Select the desired Auto Voice language from the drop-down menu. This selection will apply to all Auto Voice selections for the exam.
GCTREV166 | ADD_OR_DELETE_A_GROUP | scan parameters,add,delete | add or delete a group | revolution ct | How to add or delete a group? | Add a group Use this procedure to insert another set of images following the prior group with all of the same factors, except for the Start and End locations within the same series. 1. From the Scan Settings window, right-click on the last task in the series. 2. Click Add Group. • Each time you click Add Group a new group is added. • The start location of the new group is automatically set contiguous to the end of the prior group. • The end location is determined by the number of slices, slice thickness, and scan interval (Axial scan type) or image interval (Helical scan type). Delete a group To remove an entire group from a series, follow these steps. 1. Right-click on the group you want to delete. 2. Click Delete Group.
GCTREV167 | AQUIRE_A_SCAN | scan parameters,asir-v | aquire a scan | revolution ct | How to aquire a scan using asir-v? | 1. Prescribe the scan. 2. To enable ASiR-V and select an ASiR-V level in percentage open the Primary Recon settings collection by clicking the Primary Recon title or icon. Click [ASiR-V]. 3. Select the desired ASiR-V percentage level or use SmartmA. If you want to use ASiR-V to maintain image quality but reduce dose, from the kV and mA Control settings collection, reduce mA. 4. Start the scan
GCTREV168 | SMART_PREP_SERIES | scan parameters,smart prep | smart prep series | revolution ct | how to set up a smart prep series? | Use this procedure to set up a SmartPrep series. 1. Add a SmartPrep sub-step for a selected series by selecting the diagnostic group to which the SmartPrep task is to be added and using the right mouse menu to add a SmartPrep task to the end of the task list for the series. 2. If the diagnostic scan tasks are not confirmed, confirm the settings for these groups. After the last is confirmed, the SmartPrep Setup window opens. 3. From the scout localizer on the display monitor, click and drag the slice line to define the Monitor location. Alternatively, type in the S/I location in the Monitor Location text field in the Baseline and Monitor Collection, for example, S110. 4. Expand the Baseline and Monitor Collection and enter the following values:○ mA (range 10 to 140 in 10 mA increments), This is the mA used during the Monitor and Baseline Phases. Typically enter 50 mA for most studies. The system defaults to a low mA value, which keeps heat units to a minimum. You can adjust mA value depending on the patient size. The mA selected should be adequate to provide images that allow the detection of contrast but do not need to be of high diagnostic quality. ○ kV is a non-editable field and is the same as the value set for the first diagnostic scan. If this needs to be changed, go back to the Diagnostic scan task and change the settings there. ○ Monitoring Delay (0 to 60 seconds in 1-second increments) sets the time before monitoring scans begin. The time varies based on the anatomy being monitored. ○ Monitoring Inter Scan Delay (1 to 60 seconds in 1-second increments) sets the time between monitoring scans. ○ Enhancement Threshold (0 to 1,000 in steps of 1) sets the bar on the graph to monitor threshold enhancement. When Dynamic Transition is On, the Scan Phase starts automatically when the HU value of the Transition ROI reaches the Enhancement Threshold. This is the magnitude of the change in HU level with respect to the HU level in the baseline scan. ○ Auto Minimum sets the Diagnostic Delay to the minimum value possible, taking into consideration location of the monitor phase, start phase location, and AutoVoice premessage length. When Auto Minimum is set, the Diagnostic Delay value is read only. Use this to see the minimum delay value for your specific scan. ○ Diagnostic Delay, (up to 60 seconds in one-second increments) typically set to the minimum value (automatic) for arterial studies and 10 seconds or longer for venous studies. The minimum value is constrained to the time needed to move the table and play the auto-voice message. The minimum for scans without AutoVoice and no table move is about 1.1 seconds. ■ Allow enough time to deliver breathing instructions if you are not using AutoVoice. ■ The minimum time may vary if AutoVoice is on and depending on the monitor location relative to the start location. ○ AutoVoice is available if AutoVoice is prescribed in the diagnostic scan. If AutoVoice is enabled for the series, the pre-message will be played as part of the diagnostic delay. The diagnostic delay will be increased to accommodate the length of the AutoVoice message. Autovoice is not editable from the Baseline setup screen. Make any changes to it in the diagnostic settings. 5. Set the contrast settings. If the integrated injector option is enabled, see the Enhanced Xtream Injector section in this chapter. If the integrated injector option is not enabled, see the Enter Contrast Description section in the Scan chapter. 6. From the SmartPrep Setup window, click [Acquire Baseline]. 7. Review and respond to the potential alert message that indicates that the projected dose efficiency had fallen below 70%. 8. Proceed to Acquire a scan without Dynamic Transition or Acquire a scan with Dynamic Transition.
GCTREV169 | AQUIRE_A_SCAN_WITHOUT_DYNAMIC_TRANSISTION | scan parameters ,dynamic transition | aquire a scan without dynamic parameters | revolution ct | Baseline Phase without Dynamic Transition The baseline scan is taken as a standard axial scan with one rotation and one image and allows you to establish an area to monitor contrast enhancement. 1. After the SmartPrep settings are confirmed, the Baseline scan is taken. When the Baseline scan is complete, the Monitor Phase Setup window displays. 2. Review the values on the Scan Timeline for the monitor and scan phases. 3. If the baseline scan is not correct (for example if the patient moved or the scan location is not right), click [Retake Baseline] and return toSet up a SmartPrep series.4. Press [Move to Scan] to advance the table to the monitoring location. 5. Press [Start Scan]. 6. Once the scan is complete, from the reconstructed image on the display monitor, click [Ellipse ROI] or [Box ROI] to place a maximum of three ROIs over areas of interest. Monitor ROIs are colored blue. 7. Click [Begin Monitoring] to continue to the Monitor Phase Monitor Phase without Dynamic Transition The Monitor phase acquires images at the monitoring location during the delivery of intravenous iodinated contrast material and graphically displays the images, charts the enhancement thresholds, and clock with the time since monitoring began. Use the contrast graph, contrast table, and/or monitor images to determine when the diagnostic scan can begin. 1. After the Baseline scan settings are confirmed, the [Start Scan] button begins to flash. 2. Simultaneously press [Start Scan] and start the IV contrast injection. A maximum of 40 monitoring scans is available. Move To Scan is typically not required since the last scan was at the baseline location. NOTE: If using an integrated injector, the step to start the IV contrast injection is not needed as the command will be sent by the scanner to the injector to start the injection. 3. The system waits the time set in the Monitoring Delay area, then begins acquiring images at the time set for the interscan delay (ISD). The total SmartPrep monitor scan time shows next to the elapsed timer. The ISD countdown displays on the Real Time Information area at the bottom of the Monitor Phase window. 4. Using the tools on the Monitor Phase window to determine when the contrast has reached a sufficient level to take the diagnostic scan. When this occurs, use the [Begin Scan Now] button to transition to the Scan Phase
GCTREV170 | AQUIRE_A_SCAN_WITH_DYNAMIC_TRANSISTION | scan parameters,dynamic aquisition | aquire a scan with dynamic aquisition | How to aquire a scan with dynamic aquisition? | Baseline Phase with Dynamic Transition The Baseline phase acquires one non-contrast scan and allows you to establish an area to monitor contrast enhancement. 1. After the SmartPrep settings are confirmed, the Baseline scan is taken. After the Baseline scan is complete, the Monitor Phase Setup window displays. 2. Review the values on the Scan timeline for the monitor and scan phases. 3. Press [Move to Scan] to advance the table to the monitoring location. 4. Press [Start Scan]. 5. In the case that the baseline scan is not correct (for example if the patient moved or the scan location is not right), click [Retake Baseline] and return to Set up a SmartPrep series. 6. Once the scan is complete, from the reconstructed image on the display monitor, click [Ellipse ROI] or [Box ROI] to place a maximum of three ROIs over areas of interest. Monitor ROIs are colored blue. For example, if scanning a liver, place an ROI in liver parenchyma away from vessels, if evaluating a vessel place the ROI over the vessel. Use, as needed, the zoom, roam, display normal, hide/show graphics, erase, or explicit mag (factor range: 0.5 to 2.0) from the SmartPrep display control panel. 7. Select one of the ROIs to be used for the dynamic transition. The dynamic transition ROI is colored orange-yellow. 8. Click [Begin Monitoring] to continue to the Monitor Phase. Monitor Phase with Dynamic Transition The Monitor phase acquires images at the monitoring location during the delivery of intravenous iodinated contrast material and graphically displays the images, charts the enhancement thresholds, and clock with the time since monitoring began. The contrast graph, contrast table, and/or monitor images can be used by the technologist to determine that the diagnostic scan can begin. 1. After the Baseline scan settings are confirmed, the [Start Scan] button will begin to flash. 2. Simultaneously press [Start Scan] and start the IV contrast injection. A maximum of 40 monitoring scans are available. Move To Scan is typically not required since the last scan was at the baseline location. NOTE: If using an integrated injector, the step to start the IV contrast injection is not needed as the command will be sent by the scanner to the injector to start the injection. 3. The system waits the time set in the Monitoring Delay area and then begins acquiring images at the time set for the ISD. The total SmartPrep monitor scan time and Interscan Delay time between monitor images is displayed on the Real Time Information screen on the Monitor Phase screen. 4. Use the tools on the Monitor Phase window to determine when to switch to the diagnostic scan phase. In all the viewports with ROI information, the Dynamic Transition ROI will be visually distinct from the other ROIs. 5. The system will auto-transition when the HU Enhancement level has been reached on the ROI selected as the transition ROI. 6. Using the tools on the Monitor Phase window, monitor the progress of the contrast. a. If it looks like the dynamic transition will happen too early, uncheck the Dynamic Transition checkbox to disable this function. Note that once this is disabled, it cannot be re-enabled. Once this is done, you must click [Begin Scan Now] to transition to the diagnostic scan. b. If it looks like the dynamic transition will happen too late, click [Begin Scan Now] to transition to the diagnostic scan. If you click [Skip to Diagnostic] to transition to the diagnostic scan, the SmartPrep task is marked as skipped. Clicking [Begin Scan Now] sets the SmartPrep task to completed.
GCTREV171 | SCAN_PHASE | scan parameters | scan phase | revolution ct | what is scan phase? | The Scan Phase is the start of the actual scan prescription following the Monitor Phase. 1. After you clicked [Begin Scan Now] or Dynamic Transition has occurred, deliver the first breath hold instructions if Auto Voice Pre-Message is not turned on. 2. The system will: a. Move to the scan prescription start location (no need to press Move to Scan). b. Play the auto-voice message at the appropriate time if Auto Voice Pre-Message is turned on. c. Start the scan (no need to press Start Scan). d. Save an image of the last monitoring phase screen including the images, graphs, and table and save in series 99 for the study. e. Save the baseline image and monitoring images in series 200. If an extra monitoring phase was added, this will be in series 201. 3. When the scan is complete close the study by clicking [Done Scanning]. The Reports step displays.
GCTREV172 | ENHANCED_XTREAM_INJECTOR | scan parameters | enhanced xtream injector | revolution ct | How to set up enhanced extreme injector? | Before you start • A cable from the injector must be connected to the gantry. • Confirm that the injector is connected to the system and power is on. • Setup the syringes. • Prepare the Injector as described in the injector operator manual. To set up the Enhanced Xtream Injector after the patient and the injector are prepared, follow these steps. Review the injector status messages. 1. Prepare the injector. 2. From the Contrast collection, select Enhanced Xtream Injector from the IV Contrast field dropdown menu. ○ The button turns on and synchronizes the start of the scan and injector. ○ IV contrast is automatically turned on when Enhanced Xtream Injector is turned on. 3. Set injector settings on the Injector Settings window. To open this window, click the icon next to the Injector Settings field. The injector parameters can also be defined in Protocol Management. 4. Click [Apply] on the Injector Settings window when you have adjusted the settings. 5. From the Scan Settings screen, set the scan parameters. 6. Click [Confirm Settings]. ○ If there is an alert next to the Contrast Collection, check and resolve the injector parameters on the system or the Injector. Confirm Settings is only active when both the system and the injector are ready. Some errors may require a power cycle of the injector to resolve. ○ Complete any required fields and note the injector status information on the Scan Progress screen. 7. Press [Start Scan] to start the scan and injection.
GCTREV173 | CARDIAC_WORKFLOW | scan parametes,cardiac | cardiac workflow | revolution ct | explain cardiac workflow. | 1. Prepare the patient. 2. Attach ECG electrodes to the patient. 3. Acquire a scout scan. 4. Acquire a heart localizer scan (recommended, but optional). 5. Acquire a manual bolus timing scan, or utilize the SmartPrep feature in the next step. 6. Acquire a contrast enhanced cardiac scan. 7. Create new images from cardiac scan data. 8. Review reconstructed images.
GCTREV174 | PREPARE_THE_PATIENT | scan parameters,cardiac | prepare the patient | revolution ct | How to prepare the pateint ? | Prepare the patient Follow these steps to prepare the patient for the cardiac exam. Exam prep instructions Advise the patient of the following: • Fast for four hours prior to the exam. • No caffeine 12 hours prior to the exam. • No cardiovascular exercise prior to the exam. • Thoroughly explain all phases of the scan before you start the exam to keep the patient as calm as possible.
GCTREV175 | BREATHING_INSTRUCTIONS | scan parameters,cardiac | breathing instructions | revolution ct | What are the breathing instructions to be followed? | Breathing instructions Diagnostic cardiac images are dependent on the patient’s heart rate during the data acquisition; a calm patient keeps the heart rate steady with a repeatable rhythm. Therefore, breathing instructions are critical for a successful cardiac scan. Follow these breathing instruction guidelines: • Deliver the same instructions for each series starting with the scout scan. • Program cardiac breathing instructions with AutoVoice to deliver consistent breathing instructions.○ The voice instructions should be 10 seconds or longer so that the patient is breathing slowly. ○ After you say, "take a breath in and hold it," allow 3–5 seconds of silence before you click the [Stop] recording button. This gives the patient enough time to hold his breath before the scan starts and for the heart rate to stabilize before the scan. This time prevents the patient from breathing in during data acquisition, which results in image motion. ○ Use the Preset Delay feature in AutoVoice to program an additional 1–7 seconds prior to X-ray On, as needed for each patient. • The scan time for the contrast enhanced cardiac gated acquisition is approximately 1 second for a single scan location that covers the entire heart. The heart rate stabilizes 2–7 seconds after a breath hold, therefore, have the patient start the breath hold 3–5 seconds prior to X-ray exposure. Watch the ECG waveform display so you can determine when the patient's heart rate stabilizes within the breath hold. Adjust your breath hold instructions to achieve as stable of a heart rate as possible during the contrast enhanced cardiac gated scan. • According to your physician and site's policy, deliver 2–4 liters of oxygen via nasal cannula if the patient has difficulty holding his breath. This can help lower the heart rate. • Practice the breathing instructions with the patient.
GCTREV176 | IV_SETUP_AND_CONTRAST | scan parameters,cardiac | iv setup and contrast | revolution ct | Explain iv setup and contrast. | • To conserve CT table time, set up the patient IV line outside the scan room. • Explain the contrast injection and potential effects, even if the patient has had a previous contrast injection. Reassure the patient that what he feels during the injection is normal and to remain relaxed during the scan. An informed patient is a calm, less anxious patient. • Have the injector loaded with contrast by a trained person authorized by your site. The total volume and contrast strength is site dependent. ○ For single-barrel injectors: load with contrast. ○ For dual-barrel injectors: load one syringe with contrast media and one syringe with saline.
GCTREV177 | ECG_ELECTRODES | scan parameters,cardiac | ecg electodes | revolution ct | How to attach ecg electrodes to patient? | Use this procedure to attach electrodes to the patient for an ECG gated scan. 1. Bring the patient into the scan room. 2. Verify that you have a signed consent form, if needed. 3. Transfer the patient to the CT table. ○ Position the patient supine, feet first. ○ Landmark the patient at the sternal notch. 4. Review the ECG trigger device status on the gantry display.Make sure the selected ECG trigger device is connected and active.5. Place the patient’s arms above their head and elevate with pillows or sponges. Do not place the arms flat on the table or gantry. This position can limit vessel flow and cause arm and shoulder motion. As the table moves into the gantry during the scan, it tends to cause the patient to arch his back, raising the chest and causing incorrect registration artifacts. 6. With the patient's arms above his head, place the RA and LA electrodes over the clavicles. Do not place any electrodes over muscle, scar tissue, or hair for the following reasons: ○ Muscle: The ECG can register muscle electrical activity due to the above head arm position. Help the patient relax his arms and reduce shoulder muscle fatigue. ○ Scar tissue: This is denser than normal tissue, which makes it difficult to get a good signal. If the patient has scar tissue in the shoulder area, place the electrode out onto the patient's arm. In the chest area, place the electrode where there is no scar tissue. ○ Hair: Electrodes placed over a very hairy area do not allow good skin contact. If necessary, shave a 4 inch square area for electrode placement. 7. To remove any lotion or oils, gently scrub the electrode location with cotton gauze until the skin is a healthy pink. Do not use alcohol to prepare the electrode area. 8. With the patient in the scan position (arms above the head), apply the ECG electrodes and leads to the patient no more than 5–10 minutes before the scan. Follow these guidelines: ○ Use GE recommended electrodes: Dyna/Trace1500 by ConMed. Electrode ordering information is found on the top of the IVY monitor or in the Integrated Cardiac Module (ICM) Operator Manual. Using other electrodes can result in gating errors that can lead to non-diagnostic images. ○ Use electrodes made for short-term monitoring (for example: exercise monitoring). Do not use electrodes used for long-term monitoring. ○ Always use new ECG electrodes (radiotranslucent with silver/silver chloride) with a fresh gel pad. Dry or expired electrodes can result in poor signal conduction that can cause intermittent triggering.○ Keep the electrodes as far away as possible from the scan field of view. ○ Place the two upper electrodes directly on the middle portion of the patient's clavicle. Move the electrodes towards the arms to produce a better signal, if needed. ○ Place the two lower electrodes at the bottom of the rib cage.9. Review the ECG waveform. Ensure the patient’s heart rate is within allowable system limits for scanning: 30–200 beats per minute. a. If the internal ECG trigger device is selected and active, you can display the ECG status area by selecting the ECG waveform icon using the back, forward and enter keys on the gantry. b. If the external ECG trigger device is selected and active, the ECG waveform will display on the external device only. It will not display on the gantry display. 10. If the ECG waveform looks noisy and/or improper R-peak triggering occurs at any time prior to the diagnostic scan series, use the impedance measurement utility on the ECG trigger device to ensure proper electrode connection. Reposition any electrodes with poor electrode connection and recheck the impedance value. After repositioning any electrodes, initiate another impedance measurement. Repeat this process until all electrodes indicate successful impedance measurements. 11. Check the waveform. The more abnormal the waveform, the less chance for a successful cardiac exam. A calm patient keeps the heart rate steady with a repeatable rhythm. A significantly varying heart rate can result in non-diagnostic images. A low and stable heart rate typically allows for a high-quality study with low radiation dose.12. Complete these activities to improve the ECG trace, if needed. ○ Check the cable connections. ○ If the external ECG trigger device is selected, position it away from the table and gantry to minimize sensitivity to noise. ○ Consider an alternative electrode placement: remove electrodes, prep the patient again, apply fresh electrodes, and repeat the impedance check. ○ Check default electrode selection on the active ECG trigger device.
GCTREV178 | SCOUT_SCAN | scan parametes,cardiac | scout scan | revolution ct | How to aquire a scout scan? | Use this procedure to acquire a scout scan series for the cardiac protocol. 1. From the scan monitor, open the Patient Schedule drawer. 2. Select the patient to be scanned and double left-click or click [Select Patient]. 3. Enter the patient's information. 4. Start or connect the IV line to the contrast tubing. 5. Review the contrast effects with the patient. 6. Attach the electrodes to the patient, if you have not already done so. NOTE: Do not scan until you have an acceptable cardiac waveform and heart rate. 7. Secure both IV tubing and ECG cables. Make sure that the cables do not dangle from the table. ECG lead wire movement during the scan can result in a degraded signal. 8. Review breathing instructions with the patient. 9. From the Protocol Selection window, select a cardiac protocol and click [Accept]. 10. From the Scan Settings window, review the current settings and click [Confirm Settings]. 11. Press Move to Scan. 12. Deliver the breathing and table movement instructions. 13. Press Start Scan. 14. Repeat steps 11–13 to acquire an orthogonal scout scan. 15. Proceed to Section 4.
GCTREV179 | HEART_LOCALISER_OR_CALCIUM_SCORE_SCAN | scan parameters,cardiac | heart localiser or calcium score scan | revolution ct | How to aquire a heart localiser or calcium score scan? | Use this procedure to acquire Axial images that can be used to determine the start and end locations for the heart scan, and to determine the location for the contrast bolus tracking scan. 1. From the Scan Settings window, click [Continue]. 2. Choose one of the following options: ○ Option 1: Gated localizer for calcium scoring 1. Make sure the scan type is Cardiac. 2. For calcium scoring recommended scan settings, 3. Adjust the position of the Graphic Rx region to fully cover the heart. ○ Option 2: Ungated localizer 1. Maker sure the scan type is Helical. 2. Recommended settings: 120 kV, 0.5 seconds, 60 mA (100 mA for large patients), 1.375:1 pitch, 55 table speed, 40 mm detector coverage. 3. Position the localizer lines to cover just the heart, approximately from the carina to 2 cm below the apex of the heart. 3. Evaluate the ECG waveform for heart rate stability. 4. Click [Confirm Settings].5. Press Move to Scan. 6. Deliver breathing and table movement instructions. Give the patient the same breathing instruction that they have practiced. 7. Press Start Scan. 8. Continue on the Scan Settings window to one of the following procedures: ○ Acquire a manual bolus timing scan and acquire a contrast enhanced cardiac scan. ○ Acquire a contrast enhanced cardiac scan using SmartPrep
GCTREV180 | MANUAL_BOLUS_TIMING_SCAN | scan parameters,cardiac | manual bolus timing scan | revolution ct | How to aquire a manual bolus timing scan? | Use this procedure to calculate the bolus transit timing to determine when to start the contrast injection to capture the arterial phase. Omit this step if SmartPrep will be used. 1. From the Scan Settings window, click [Continue]. Recommended settings: 120 kV, 1 rev/s, 50 mA, 5 s prep delay, 1 s ISD, 5 mm/1i 2. Place the localizer line one centimeter below the carina and just above the base of the heart, the optimal location to find the ascending aorta for a timed contrast injection. 3. Evaluate the ECG waveform for heart rate stability. From the Scan Settings window, check that the ECG is reported out in the ECG and Gating collection, or open the ECG and Gating collection and turn on ECG Gating. 4. Click [Confirm Settings]. 5. Deliver breath hold instructions identical to the instructions you will use for the cardiac gated series. See breathing instructions for more details. 6. Press Move to Scan. 7. Start the injector. ○ If using the Integrated Injector option: The injector is started automatically when Start Scan is pressed. ○ If not using the Integrated Injector option: Simultaneously press Start the Injector and Start Scan. 8. Observe the heart rate during the scan. 9. Click the viewport with the reconstructed test bolus images. Review the images and identify the image with contrast in the ascending aorta. 10. From the image monitor, open the Measure / Annotate collection and select MIROI. 11. From the MIROI Analysis window, select [Ellipse ROI]. 12. Click and drag the ROI to the ascending aorta. 13. Size the ROI to fit completely inside the aorta. 14. Click [OK]. 15. Use the graph to calculate the bolus timing a. Count each tick mark to the peak of the curve. b. Multiply the number of tick marks by two. c. Add 5 seconds for prep group delay. ■ This time represents the time it takes for the contrast once injected to reach the aortic root where the coronary arteries branch off (Time to Peak enhancement). ■ If you use the GE reference protocol, image number one is at 5 seconds and the tick marks are 2 seconds apart, as shown in the graph above. d. Add 3–4 seconds to allow for filling of the distal coronary vessels. The minimum total time should be set to 24 seconds or greater. In the example below, total time = 24 seconds. The value will be entered in the Prep Group Delay fields in the Timing collection on the Scan Settings window, for the contrast enhanced cardiac scan. ■ (7 tick marks × 2 = 14 seconds) ■ (+5 for prep group delay = 19 seconds for peak enhancement time) ■ (+4 for distal filling of vessels = 23 seconds) ■ (+1 additional delay to make total delay 24 seconds or greater = 24 seconds) NOTE: If total time >= 24 seconds, no additional delay is required. 16. Click [Continue] to proceed to the next procedure
GCTREV181 | CONTRAST_ENHANCED_CARDIAC_SCAN | scan parameters,cardiac | contrast enhanced cardiac scan | revolution ct | How to aquire contrast enhanced cardiac scan? | Use this procedure to acquire the contrast enhanced cardiac scan. This is the final scan before transferring the images to the Advantage Windows Workstation. 1. Confirm the ECG trace is strong and stable. 2. Select the SFOV 3. Position the scan lines to cover the complete heart from approximately 1 cm superior to the left main coronary artery ostium to 1 cm inferior to the apex of the heart. 4. Confirm that a cardiac scan mode is selected. 5. Review the heart rate measurements made during the breath hold recording, and the ECG and Gating settings that were adjusted based on those measurements. If you have not recorded a breath hold recording, do so now. 6. Review the ECG and Gating settings. ○ The data acquisition windows/parts indicate the portion or portions of the cardiac cycle that will be imaged. ○ When using Auto Gating, the acquisition windows/parts will change depending on the patient's heart rate during the most recent ECG recording during a breath hold. 7. Arrhythmia Management ○ When using Auto Gating and an irregular heart rhythm is detected either during the last breath hold or during the previous few minutes, the system may scan the patient two times, to increase the likelihood for a clinically successful scan. The Auto Gating Profile determines if this setting is enabled or disabled. The user has the ability to change this setting by disabling Auto Gating and changing the Arrhythmia Retrigger setting. ○ This is indicated by a value of two scans. ○ This value can be changed after setting Auto Gating to “Off”. Change this value if it is not appropriate. 8. Enter the prep delay that you determined from the MIROI timing bolus scan. The typical prep delay time is Peak + 9 seconds, assuming the time bolus scan had a prep delay of 5 seconds. 9. The primary reconstruction for the scan series and may influence scan parameters. If needed, prescribe additional secondary reconstructions. 10. From the Contrast collection, enter the contrast description and amount for the IV Contrast setting. Alternatively, you may choose one of any preset items from the list. 11. Program the contrast injector. ○ The protocol to use varies depending on the contrast type, patient size, and injector capability. ○ Consider adjusting the volume of contrast on the second phase to ensure a 6 second range from the end of the contrast injection to the start of X-ray on. This delay time ensures that if using saline, you are getting the advantage of the saline bolus in the right chamber. If you start the scan too soon, you will not see enough saline flush through the SVC. If you start the scan too late, you will completely flush out the right chamber. ○ The time delay from the end of the contrast injection (3) to the start of the scan (5) should not exceed 6 seconds. ○ Skip the next step if SmartPrep is not selected 12. . Typical SmartPrep parameters for cardiac scans: ○ mA = 40–60 ○ Monitoring Delay = 10.0 seconds ○ Monitoring Inter-Scan Delay (ISD) = 1.7 seconds produces an image every 2.0 seconds ○ Enhancement Threshold = 150 ○ Diagnostic Delay = Auto Minimum or desired delay time 13. Arm the injector. 14. Click [Confirm Settings]. 15. Give the patient the same breathing instructions as the previous series. 16. Press Move to Scan. 17. Start the injector. ○ If using the Integrated Injector option: The injector is started automatically when Start Scan is pressed. ○ If not using the Integrated Injector option: Simultaneously press Start the Injector and Start Scan
GCTREV182 | ECG_AND_GATING_SETTINGS | scan parametrs,cardiac | ecg and gating settings | revolution ct | How to set the ecg and gating settings? | Set the ECG and Gating settings The ECG and Gating settings control the timing for when the scan data will be acquired. When you select ECG and Gating from the Scan Settings window, the window displayed will depend on the scan type selected. • The heart icon on the gantry display indicates that the internal monitor is connected to the gantry. • The instantaneous BPM heart rate is shown on the user console and gantry display. • If the heart rate does not display and a lead off message is displayed, follow these guidelines: ○ Check Gating box to view the ECG trace on the scan monitor. ○ Check electrode placement. Consider an alternative placement to improve the cardiac signal.
GCTREV183 | RECORDING_OF_HEART_RATE_DURING_BREATH_HOLD_AUTOMATICALLY | scan parameters,cardiac | recording of heart rate during breath hold automatically | revolution ct | How to automatically record the heart rate during breath hold? | 1. From the image monitor, click 2. Click System Preferences. 3. From the System Preferences window, click the ECG tab. 4. In the Breath Record Length area, adjust the time. Valid recording lengths are 5–15 seconds. The system will apply the Breath Hold Recording Length time and the AutoVoice message prescribed in AutoVoice. ○ If no message is selected, an AutoVoice message will not be played and you will need to provide manual breathing instructions. ○ The breathing commands should be the same as those given for the contrast enhanced scan 5. Click [Save]
GCTREV184 | RECORDING_OF_HEART_RATE_DURING_BREATH_HOLD_MANUALLY | scan parameters,cardiac | recording of heart rate during breath hold manually | revolution ct | How to manually record the heart rate during breath hold? | 1. From the image monitor, click 2. Click System Preferences. 3. From the System Preferences window, click the ECG tab. 4. In the Breath Record Length area, adjust the time. Valid recording lengths are 5–15 seconds. The system will apply the Breath Hold Recording Length time and the AutoVoice message prescribed in AutoVoice. ○ If no message is selected, an AutoVoice message will not be played and you will need to provide manual breathing instructions. ○ The breathing commands should be the same as those given for the contrast enhanced scan. NOTE: The patient’s heart rate is automatically recorded during the Scout, Timing Bolus, Calcium Scoring (Smart Score), or other CT scan series as long as the scan duration is at least as long as the breath hold duration. 5. Click [Apply] or [Save]. 6. Click [Record New HR] from the Acquisition Window and Heart Rate widget with Auto Gating “on”. ○ The selected AutoVoice message plays, giving pre-message breathing instructions. ○ When recording is completed, AutoVoice plays the post-message breathing instructions. 7. A jump point is added to the viewer. Click the jump point to review the recorded trace.
GCTREV185 | CARDIAC_SECONDARY_RECONS | scan parametersw,cardiac | cardiac secondary recons | revolution ct | How to prescribe basic cardiac seconadry recons? | Use this procedure to prescribe cardiac secondary reconstructions pre-scan or post-scan. • Cardiac Axial acquisitions allow reconstruction during a portion of the R-to-R interval. The amount of data acquired is controlled by cardiac phases prescribed in the acquisition. • See Create new images from scan data for retrospective management details. 1. From the image monitor, in the Reconstruction and Imaging Processing area, select the secondary recon task and click [Add Recon]. 2. The window displays all of the recon settings available for the secondary reconstruction. 3. Optional: Type the Start and End locations that fall within the range of the selected group. Add an S (+) or I (-) prefix to the location. 4. Optional: Click Thickness and enter a thickness that populates all groups. ○ If you only need to change the thickness for one group, or only one of multiple groups, select the thickness area for that group. ○ The original detector configuration at which the patient was scanned affects how you can change the slice thickness. The thickness in black is available for reconstruction. 5. Optional: Click DFOV and type a DFOV value. The system defaults to the DFOV used for the scan acquisition. 6. Optional: Click R/L Center and select an individual or all groups. Type an offset value. Use one of the following methods to determine the offset value: ○ From the Recon tab, click [N] under the Graphic Retro column to display the Graphic Retro control panel on the image monitor. The cursor can be used to determine DFOV, R/L and A/P centers. ○ Place the real-time cursor in an image and view the upper-left corner for the RAS coordinates. ○ Click [List/Select] and select the exam and series from which you want to use the R/L values. The series values are displayed in the browser menu. 7. Optional: Click A/P Center and enter an offset value. 8. Optional: Click Recon Type and select the algorithm. 9. From the ECG and Gating settings collection, click the Recon mode for the phase prescription method. This allows you to select the phases of the cardiac cycle from which images are created. ○ Single Phase (post-scan only): User specified single % or ms value and interval to be used for reconstruction. ○ Multi-phase (post-scan only): User specified range of % or ms values to be used for reconstruction. ○ Center Phase (1, All): A reconstruction will occur at the center of the acquisition window(s). ○ Earliest to Latest: A multi-phase reconstruction based on the scan prescription with a user-defined increment. ○ Earliest to Latest (All): A multi-phase reconstruction adapted to the available data acquired with a user-defined increment. The manual options, Single Phase and Multi-phase, for defining the reconstruction timing relative to the cardiac cycle, are only available after the scan has been acquired. A typical Multi-phase prescription is 70–80% in 5% intervals, which is applied to all available heart cycles during X-ray ON. 10. In ECG and Gating settings, select the mode (Single Phase, Multi-phase, et cetera) and enter the desired phase(s) for reconstruction. ○ %R displays standard phase percent reconstruction ○ +ms displays absolute millisecond reconstruction ○ -ms displays reverse absolute millisecond reconstruction 11. For Multi-phase, enter the Start phase, End phase, and Interval to reconstruct. ○ The recommended reconstruction phases are 70–80% with a 5% increment .○ If you are acquiring a functional imaging exam for ejection fraction and wall motion, prescribe phases from 5–95% in increments of 10% and consider reconstructing the data with a 2.5 mm thickness. 12. Click [Save] to close the Phase Selection dialog box
GCTREV186 | ADJUST_RECONSTRUCTION_TIMING | scan parametrs,cardiac | adjust reconstruction timing | revolution ct | How to graphically adjust scan parameters? | Use this procedure to move or reposition image reconstruction timing within a single heart cycle. Method 1 Place the cursor over the recon location and click, drag, and drop to the new location. The annotation in the reconstruction location indicates the % or ms where images will be reconstructed. Method 2 1. Place the cursor over the recon window you wish to move, and right-click Adjust Recon Timing. 2. Type and enter the location as + or - milliseconds. The recon timing for selected heart cycle moves to the new location. 3. Click [OK]. Method 1 and 2 For all cardiac scan types, the recon window cannot be moved outside of its current X-ray on interval
GCTREV187 | INSERT_DELETE_OR_MOVE_AN_R_PEAK_TRIGGER | scan parameters,cardiac | insert delete or move a r peak trigger | revolution ct | How to insert delete or move a r peak trigger? | Use this procedure if gating issues occur during the scan or if the system does not trigger accurately at the R-peak of the ECG trace. • Insert, remove, or move a trigger to normalize a heart cycle when a trigger occurs at an unwanted location. This can be due to abnormal ECG waveform patterns, external noise interference, or a low amplitude trace. • Remove an image reconstruction location from one or more heart cycles to improve image quality when an arrhythmia is present in the scan. 1. Open a prior ECG-gated scan series in the Reconstruction and Image Processing area on the image monitor. 2. To insert a trigger, place the cursor anywhere on the ECG trace where you want to insert it and right-click Insert R-peak Trigger. An additional trigger will display on the ECG trace. 3. To delete a trigger point on the ECG trace and to normalize the heart cycle, place the cursor on the trigger and right-click Delete R-peak Trigger. 4. To move a trigger, place the cursor over the trigger and click and drag it to the new location
GCTREV188 | ARRHYTHMIA_RETRIGGER_TO_RECONSTRUCT | scan parametrs,cardiac,arrhythmia | Arrhythmia Retrigger to Reconstruct | revolution ct | How to change heartbeat from arrhythmia retrigger to reconstruct? | 1. Place the cursor in the ECG Editor, and right-click and select Change Scan to Reconstruct. 2. The reconstruction window moves from one heartbeat to the other heartbeat.
GCTREV189 | IMAGES_IN _FILE_MANAGER | scan parameters,cardiac,image display | images in file manager | revolution ct | How to access images in file manager? | To open File Manager to access browser applications, follow the steps below. File Manager is always accessible on the image monitor (right monitor). 1. On the image monitor, click the File Manager “drawer”. (You do not have to click the File Manager icon – just click the drawer labeled with the [File Manager ] icon 2. From the Exam List, double-click the exam you want to view. The exam opens in a patient session tab, and the File Manager drawer closes. Alternatively, click the Select Patient icon to the left of the patient’s name, or right-click the exam and then click [Open Exam].
GCTREV190 | SELECTING _IMAGES_IN_FILE_MANAGER | scan parameters,cardiac,image display | selecting images in a file manager | revolution ct | How to select images in a file manager? | 1. From the display monitor, open the File Manager drawer. 2. From the Exam list, select the patient name. 3. The Series list shows the series for the selected patient. Select the series you want to view. 4. The series loads into the Image viewport on the bottom right side of the right monitor. 5. Hover your cursor over the right of the viewport to reveal the scroll bar at the right of the viewport. By default, the first image in the series is displayed. To scroll through the images in the series, click and drag the slider on the scroll bar
GCTREV191 | CINE_LOOP | scan parameter,cardiac,iamge display | cine loop | revolution ct | How to view images in cine loop? | Use this procedure to view images in a cine loop, which is useful for tracking flow in samelocation scans taken with contrast or with motion such as flexing an elbow. 1. From the Reconstruction and Image Processing window, load the series you want to view into a viewport. 2. From the viewport control panel, click the arrow next to [Export], then click the [Cine] icon 3.In the Cine area, click the [Play] icon , which uses all the default settings. Alternatively, make selections from the Cine area.
GCTREV192 | REFERENCE_IMAGE | scan parameter,cardiac,iamge display | reference image | revolution ct | How to view a reference image? | Use this procedure to show the image from which the primary image was prescribed in small viewport within the image. This is done from the command line. • Type ri to put on single image. • Type ria to put on all images in the series. • Type noria to turn off
GCTREV193 | APPLYING_FILTERS | scan parameters,cardiac,image display | applying filters | revolution ct | How to apply filters? | 1. From the viewport control panel, click the arrow next to Image Display, then the [Filters] icon . 2. Select a filter type. ○ Click [E1], [E2] or [E3] for edge enhancement. ○ Click [LU] for lung enhancement. ○ Click [S1], [S2] or [S3] for smoothing. 3. To turn off the filters, click [None] for the filter type. Alternatively, click the [Display Normal] icon from the Image Display set of tools.
GCTREV194 | GRAY_SCALE_ENHANCEMENT | scan parameters,cardiac parameters,image display | gray scale enhancement | revolution ct | How to apply gray scale enhancement? | 1. From the viewport control panel, click the arrow next to Image Display, then click the [Filters] icon . 2. Under Gray Scale Filter, click [G1], [G2] or [G3]. 3. To turn off GSE, click [None] for the filter type. Alternatively, click the [Display Normal] icon from the Image Display set of tools
GCTREV195 | FLIP_ROTATE_IMAGES | scan parameters,cardiac parameters,image display | flip rotate images | revolution ct | How to flip and rotate images? | 1. From the viewport control panel, click the arrow next to Image Display. 2. Click the icon corresponding to how you want to reorient the image.3. To return the image to its normal orientation, click the [Display Normal] icon from the Image Display set of tools.
GCTREV196 | DICOM_GSPS_OBJECT | scan parameters,cardiac,image display | dicom gsps object | revolution ct | How to create a gsps object? | 1. In a viewport, modify the image using any of the following features: ○ Window Width and Window Level ○ Roam ○ Zoom ○ Flip ○ Rotate ○ Graphics such as ROI, measure distance, and image annotation 2. From the viewport control panel, click the arrow next to Export, then click the [Save State] icon . The GSPS object will be saved as series 10,000 for the current exam.
GCTREV197 | VIEW_GSPS_OBJECT | scan parameters,cardiac,image display | view gsps object | revolution ct | How to view gsps object? | 1. Select the GSPS object in File Manager. The images display in the Image Viewport in File Manager. 2. Use the image tools to modify the image. 3. To view additional GSPS objects in the series, click the Image List icon and select the desired GSPS from the list, then select the image view icon
GCTREV198 | APPLY_A_MATTE | scan parameters,cardiac parameters,image display | apply a matte | revolution ct | How to apply a matte? | Use the matte procedure to eliminate unwanted information on or around an image. There are two types of matte: rectangular and elliptical. The size of the matte is adjustable
GCTREV199 | MOVE_RESIZE_HIDE_OR_DELETE_A_MATTE | scan parameters,cardiac,image display | move,resize,hide or delete a matte | revolution ct | How to move ,resize,hide,or delete a matte? | 1. To resize the matte, click and drag the open red box in the lower right corner of the matte. 2. To move the matte, click and drag anywhere outside the matte. 3. To delete or hide the matte: a. From the viewport control panel, click the arrow next to Measure/Annotate. b. To remove the matte, click [Erase selected graphics] . c. To hide the matte, click [Hide Graphics]
GCTREV200 | DENSITY_READING | scan parameters,cardiac,image display | density reading | revolution ct | How to measure a density reading? | Use these procedures to obtain information on areas of anatomy and pathology. 1. From the viewport control panel, click the arrow next to Measure/Annotate. 2. Select the [Rectangular ROI] , [Ellipse ROI] or [Draw free hand ROI] . NOTE: The Ellipse ROI button is also available from the hover menu bar which appears when you hover your cursor over a viewport. 3. Position the cursor on the image and click to deposit the insertion point. ○ For Draw free hand ROI, move the solid blue box to the start point. Press and hold <Shift>, then click and drag to define the area for the trace. ○ For Rectangular ROI or Ellipse ROI, click and drag on the small box in the upper-right corner of the ROI to resize or click the ROI number and drag to a new location.
GCTREV201 | MIROI_DENSITY_TO_CALCULATE_BOLUS_TIMING | scan parameters,cardiac,image display | miroi density to calculate bolus timing | revolution ct | How touse miroi density to calculate bolus timing? | Use this procedure to calculate the scan delay for a contrast enhanced exam. Scale for HU, there are two selections. • Absolute: has an actual Hounsfield Unit number for each ROI plotted on the graph. • Relative: the ROI measured for the first image is set to zero and the change from zero is plotted for the rest of the images. 1. From the viewport control panel, click the arrow next to Measure/Annotate to display measurement icons. 2. Click the [MIROI] button to open the MIROI controls window. 3. Define the region of interest using either the eillipse ROI button or the rect ROI button. 4. Click and drag the ROI to the anatomy of interest. Size the ROI to fit completely inside the aorta. 5. Click [Plot]. 6. Use the graph to calculate the scan delay. 7. Optionally, screen save the graph. Right click and select Screen Save
GCTREV202 | ANNOTATE_AN_IMAGE | scan parameters,cardiac,image display | annotate an image | revolution ct | How to annotate an image? | 1. Click on the viewport to which you want to add annotation. 2. From the viewport control panel, click the arrow next to Measure/Annotate. 3. Click the [Annotate] button . 4. Move the mouse to the selected viewport and enter the text. ○ To move the annotation, click on the blue box and drag. ○ To move the end of the line, click and drag the arrowhead. ○ To remove the line, click and drag it into the box. 5. Click the [Annotate] icon to add another annotation. 6. Click the[Erase all graphics] icon on the viewport toolbar to remove all annotations
GCTREV203 | EXAM_OR_SERIES_TEXT_PAGE | scan parameters,cardiac,image display | exam or series test page | revolution ct | How to view an exam or series text page? | Use this procedure to view patient and scan parameter information. 1. From the viewport control panel, click the arrow next to Image Display. 2. Click the [Text Pages] . 3. Click either the [Exam Textpage], [Series Textpage] or [ROI Textpage] button. 4. Click the [Save as Secondary Image] icon to save the displayed information to series 98. You can also right click on the image, then click Screen Save 5. Use arrow keys to navigate between multiple pages. 6. Click [Quit]
GCTREV204 | HIDE_SHOW_OR_REMOVE_GRAPHICS | scan parameter,cardiac,iamge display | Hide show or remove graphics | revolution ct | How to hide show or remove graphics? | Use this procedure to temporarily remove added annotation and graphic objects on an image then re-display them. 1. To hide measurements, such as ROIs: a. From the viewport control panel, click the arrow next to Measure/Annotate . b. Click the [Hide Graphics] icon . 2. To remove measurements, select the measurement, then click the [Erase Selected Graphics.] icon or press the [Delete] button on the keyboard. 3. To hide annotations: a. Select the annotation you want to hide. b. Click [Hide Graphics] icon . c. To remove annotations, click the[Erase all graphics] icon on the viewport toolbar
GCTREV205 | SAVE_AN_IMAGE_SCREEN | scan parameters,cardiac,image display | save an image screen | revolution ct | How to save an image screen? | 1. From the viewport control panel, click [Film Tools]. 2. Click the [Save as Secondary Image] icon to save the displayed information to series 98. You can also right click on the image, then click Screen Save.
GCTREV206 | ADD_OR_SUBTRACT_IMAGES | scan parameters,display | add subtract images | revolution ct | How to add and subtract images? | To create a new image set, for example add thin slices together or subtract pre/post contrast series, follow the steps below. Processed images will be annotated with the date. 1. In File Manager, highlight the exam you want to work with, right click and select [Add/Sub] to open the Image Combination window. Alternatively, highlight the series you want to work with, right click and select [Open with], then [Add/Sub]. 2. Select the images. ○ If one set is selected, each operation produces one resulting image. ○ If two sets are selected, images in the two sets are paired according to physical location in the patient’s body. Unpaired images are ignored. 3. On the Image Combination window, click the left [Select Set]. 4. If you are adding/subtracting two sets of images, select the second set of images, and click the right [Select Set]. ○ By default, equal weighting is applied to the pixels in each pair, but you can use the Ratio slider to change the weighting values. ○ Drag the slider to the left to increase the pixel weighting of the images on the left [Select Set] button. Drag the slider to the right to increase the pixel weighting of the images on the right [Select Set] button. 5. Click [+] or [-]. 6. For subtraction, click [Accept Negative Pixels] unless you want the negative pixel values set to 0 in the resulting image. 7. Typically, use the system provided series number. 8. Click [=] (equal sign) to generate the images in the exam defined by the left [Select Set] button. 9. To repeat the add/subtract procedure with a new series number and description, click [New Save Series].
GCTREV207 | BIND_SERIES | scan parameters,display | bind series | revolution ct | How to bind series? | 1. In File Manager, highlight the exam you want to work with, right click and select [Add/Sub] to open the Image Combination window. Alternatively, highlight the series you want to work with, right click and select [Open with], then [Add/Sub]. 2. Select the images. 3. Click the left [Select Set] button. 4. Click the right [Select Set] button. 5. Click [Bind]. Use the system-provided series number, or enter a new series number in the Save Series field and a name in the Description field. 6. Click [=] to generate the images in the exam defined by the left [Select Set] button. 7. To repeat the add/subtract procedure with a new series number and description, click [New Save Series].
GCTREV208 | CREATE_IMAGES_WITH_MIN_AND_MAX_VALUES | scan parameters,display | create images with min and max values | revolution ct | How to create images with min and max values? | 1. In File Manager, highlight the exam you want to work with, right click and select [Add/Sub] to open the Image Combination window. Alternatively, highlight the series you want to work with, right click and select [Open with], then [Add/Sub]. 2. Select the images. 3. Click the left [Select Set] button. 4. Click [Min] (minimum) or [Max] (maximum). 5. Typically, use the system-provided series number. 6. Click [=] (equal). 7. Optional: Click [New Save Series] to repeat the add/subtract procedure with a new series number and description.
GCTREV209 | AUTOMATIC_REFORMAT | scan parameter, display,reformat | automatic reformat | revolution ct | How to automatically reformat? | For an automatic reformat, no images are acquired before planning the reformat. 1. From the Reconstruction and Image Processing Task List, click the down arrow on the series you want to reformat. 2. Select the protocol you want to use from the Reformat Settings panel. 3. Check the Auto Run box in the lower right of the window to enable auto run. DMPR images will be generated automatically.
GCTREV210 | SEMI_AUTOMATIC_REFORMAT | scan parameter, display,reformat | semi automatic reformat | revolution ct | How to semi automatically reformat? | 1. From the Reconstruction and Image Processing Task List, click the down arrow on the series you want to reformat. 2. Select the protocol you want to use from the Reformat Settings panel. Adjust the settings for the protocol. Auto Run is not selected. 3. After the source images have been reconstructed, select the reformat task 4. Adjust the prescription parameters. 5. Click [Run]
GCTREV211 | MANUAL_REFORMAT | scan parameter, display,reformat | manual reformat | revolution ct | How to manually reformat? | For a manual reformat, the recon data is available. 1. From the Reconstruction and Image Processing Task List, click the down arrow on the series you want to reformat. 2. Select the protocol (or generic batch) you want to use from the Reformat Settings panel. Adjust the settings for the protocol. 3. The default reformat views appear in the viewports: left upper viewport-oblique, left lower viewport-sagittal, right upper viewport-axial, right lower viewport-coronal. If a protocol has been selected, batch lines appear on the views. 4. Enter the batch prescription parameters and graphically adjust the location. 5. Click [Run].
GCTREV212 | SELECTING_EXAMS_TO_BE_SPLIT | scan parameter, display | selecting exams to be split | revolution ct | How to select the exams to be split? | At scan time, select all patient records you want available to split in Patient Schedule. 1. From the scan monitor, open [Patient Schedule]. 2. From the Patient Schedule window, select the patient exams to which you want to apply Exam Split once the images are reconstructed. You can select a maximum of 15 procedures.3. Click [Select Patient]. When the patient record opens, only the description of the last exam selected displays. The system stores the others and they will be available during the split.
GCTREV213 | SPLIT_EXAMS_WHEN_SCAN_IS_COMPLETED | scan parameters,display,exam split | split exams when scan is completed | revolution ct | How to spilt exams when scan is completed? | Use this procedure to split a series into separate groups that can then be used for multiple reads and billing. 1. Complete an exam where multiple Requested Procedure ID or Accession Numbers have been selected. All Patient records that you wish to split, must be selected from the Patient Schedule when you selected [New Patient]. You cannot add procedures after the exam is ended. 2. In File Manager, highlight the exam you want to work with, right click and select [Exam Split] to open the Exam Split window. Alternatively, highlight the series you want to work with, right click and select [Open with], then [Exam Split].○ The Exam Split menu appears with the selected images displayed in the upper viewports. The system may display every image or skip some images depending on the total number of images selected. ○ If the exam you selected does not have Multiple Procedures, a dialog message is posted and Exam Split exits. 3. From the Exam Split window, click first image, press <Shift>, and click the last image to set the range of images to be grouped together per procedure code.4. From the Procedures list area, select a procedure that reflects the procedures selected from Patient Schedule. 5. From the Exam Split control panel, click [Add]. Add select images to the procedure. 6. Adjust the W/L as needed in the exam area of the display. To adjust the W/L use either the W/L Presets on the Exam Split control panel or use the mouse. 7. From the Exam Split control panel, click [Host Selection].8. Select the receiving host and click [Save]. 9. From the list of procedures on the Exam Split control panel, select the procedures you wish to send to the host and click [Send]. ○ To select multiple procedures, press and hold <Shift> while selecting procedures. ○ For systems configured for VES, send exams prior to splitting and it sends GSPS objects. ○ For systems configured as HES, images are sent after splitting and sends a new series
GCTREV214 | ANNOTATE_AN_IMAGE | scan parameter,display,reformat | add a unique annotation | revolution ct | How to annotate an image? | Use this procedure to add specific and unique annotation to an image or apply a pre-set annotation that is saved on your system. 1. Open Reformat. 2. Click the Display tab. 3. Click the [Annotate] icon . 4. On the Annotate screen, click the annotation type (simple, linked, or measure). 5. Place the cursor on the image and click to deposit the cursor and default text. 6. Enter new text in the Annotation text screen. Use this procedure to add specific and unique annotation to an image or apply a pre-set annotation that is saved on your system. 1. Open Reformat. 2. Click the Display tab. 3. Click the [Annotate] icon . 4. On the Annotate screen, click the annotation type (simple, linked, or measure). 5. Place the cursor on the image and click to deposit the cursor and default text. 6. Enter new text in the Annotation text screen.
GCTREV215 | SAVE_PRESET_ANNOTATION | scan parameter,display,reformat | save preset annotation | revolution ct | How to save preset annnotation? | Use this procedure to add a custom pre-set annotation to the Preset Annotation list. 1. Open Reformat. 2. From the Display tab, click the [Annotate] icon . 3. On the Annotate screen, type text in the Annotation Text field. The preset name automatically fills in with part of the text from the annotation text field. 4. Change the preset name. 5. Select the text. 6. Press <Backspace>. 7. Type in a new name. 8. Click [Use for other anatomy] if you want the annotation available for other anatomies or [Use for other protocol] if you want the annotation available for other protocols. 9. Click [Save as preset] to add annotation to preset annotation list.
GCTREV216 | MEASURE_DISTANCE | scan parameters,display,reformat | measure distance | revolution ct | How to measure distance? | 1. Click the [Measure Distance] icon . 2. Make one selection for each menu choice. ○ Measure: ■ Projection (2D): to view the length of the projection. When used in 3D images, the measurement will correspond to the projected red distance in the plane of the screen. ■ From Volume (3D): to view the true distance in the 3D volume. Endpoints can be positioned at different depths in the volume. Rotate 3D volumes to check the exact position. ○ Display: ■ On one slice: to view the measurement on only one slice. ■ On all slices: to view the measurement on all slices. ■ On one viewport: to view the measurement on only one viewport. ■ On all viewports: to place a measurement on different slices or volumes of a multiphase scan at the same time. ○ Along: ■ Straight line: to deposit the first and second point. ■ Curve: to measure along a curve by depositing multiple points on the image. Rightclick to validate the curved line. 3. Click an image at the start of measurement. 4. Move cursor to end point of measurement and click. 5. Click and drag the squares to adjust the points. 6. View the measurement results.
GCTREV217 | MEASURE_ANGLE | scan parameters,display,angle,reformat | measure angle | revolution ct | How to measure angle? | 1. Click the [Measure Angle] icon . 2. Click an image three times to deposit three points. 3. Click and drag the squares to adjust the points. 4. View the measurement results.
GCTREV218 | MEASURE_AREA | scan parameters,display,area,reformat | measure area | revolution ct | How to measure area? | 1. Click the [Measure Area] icon . 2. Click an image to deposit points. 3. Click and drag the squares to adjust the points around the region of interest. 4. Right-click to display the measurement results
GCTREV219 | MEASURE_VOLUME | scan parameters,display,volume,reformat | measure volume | revolution ct | How to measure volume? | 1. Click the [3D ROI] icon . 2. Click an image to deposit the 3D ROI on an object of interest. 3. Click any of the corners to size the volume. 4. View the area measurement results
GCTREV220 | MANAGE_COLOURMAPS | scan parameters,display,colourmaps,reformat | manage colourmaps | revolution ct | How to manage colourmaps? | 1. Open Reformat. 2. Click the Display tab. 3. Click the [3D ROI with color] icon . 4. Click the image to deposit the ROI. 5. Click the color ramp displayed on the left side of the image to open the Color Table screen. 6. Click the Statistics tab to view range percentages for each color. 7. Change the values of the ROI on the viewport. a. Click the Configuration tab to change the values of the ROI on the viewport. b. Click a value in the Min Value column, and type a new value. c. Click a value in the Max Value column and type a new value. d. Click a Range name column text box to type a name. 8. To change a color in the color ramp, select a color in the Color column. a. From the Select new color screen, drag the cursor over a color on the color wheel. b. View the Current selected color. c. Click [Apply]. 9. To add or remove a color in the color ramp, click [Add] or [Remove]. 10. Click the Presets tab to load and save Colormap presets. 11. Click [Save] to save setting adjusted in the Configuration tab. 12. To load a color map preset, select the color ramp preset from the menu. 13. Click [OK].
GCTREV221 | ADD_AN_ROI | scan parameters,display,roi | add an roi | revolution ct | How to add an roi? | Use this procedure to add an ROI on an image to obtain information, volumes, areas, and statistics of anatomy or pathology. The ROI allows you to: • measure the pixel intensity value at a specific point on the image • display the area or volume• display the mean, standard deviation, and minimum and maximum pixel values within the ROI 1. Open Reformat. 2. Click the Display tab. 3. Select the desired ROI tool. 3D color ROI is available one color MR or multi-color CT. 4. Define your preferences on the ROI panel. a. Choose an ROI propagation option. ■ Select One Volume to view the ROI on only the viewport on which the ROI is displayed. ■ Select All Volumes to view the ROI and statistics on all viewports. b. Select one or more Statistic options to view any combination of minimum, maximum, average, standard deviation and area (2D) or volume (3D). c. Explicitly adjust the size of the ROI. ■ For 2D ROIs, type a diameter in the Vertical and Horizontal Diameter fields. ■ For 3D VOIs, type a VOI size for the XY-plane and Z-axis depth. ■ The sphere VOI does not have a depth entry. 5. Click the desired anatomy to deposit the ROI. 6. Click and drag an ROI corner to adjust the size. 7. View the ROI measurements. 8. To hide the ROI statistics, click on the edge of the ROI to make it active, then right-click the center of the ROI, and select Hide Statistics.
GCTREV222 | CURVED_VIEW | scan parameters,display,view | curved view | revolution ct | How to create a curved view? | Use this procedure to display a reformat of a curved, complex view of tortuous vessels or organs. The curved view does not need to lie along a single orthogonal or oblique plane but can follow anatomical lines. 1. Open Reformat. 2. In one of the viewports, right-click on the view type active annotation and select [Curved]. 3. On an axial, sagittal, or coronal image, press and hold <Shift> as you deposit points along the anatomy. 4. As you trace, the curved image updates automatically in the curved reformat viewport.
GCTREV223 | X_SECTION_HISTOGRAM_VIEW | scan parameters,display,view | x section histogram view | revolution ct | How to create x section historam view? | A histogram graph shows the percentage of occurrence of each voxel value, either in a userdefined surface area on a reformatted slice (cross-section histogram) or in the entire 3D object (volume histogram). It also determines boundaries around the class of similar voxel intensities. 1. Open Reformat 2. In one of the viewports, right-click on the view type active annotation and select Histogram or X Section. ○ If you select a view type of histogram, you will immediately view a volume histogram which includes the entire 3D model. No other action is required. Place the cursor on the voxel reference line and hold <Shift>. The image displays the range of voxel class boundary lines by highlighting them in green. ○ If you select a view type of X Section, the view will display Undefined histogram until you start to define the trace. ○ Both the volume and cross-section histograms contain the same information with the exception of the name and measurement units. Cross-section histograms show total area, while volume histograms show the total volume for the entire object without cut planes. 3. For a cross-section (X Section) histogram, in the viewport containing the anatomy of interest, press <Shift> and simultaneously click the image to deposit points on the area of interest to create a trace.
GCTREV224 | MPVR_VIEW | scan parameters,display,view | mpr view | revolution ct | How to create mpr view? | Use this procedure to create an Multi-Projection Volume Reformation (MPVR) view. This type of thick slab reformatted image is often used to see vessels in Computed Tomography Angiography (CTA) scans. 1. Open Reformat. 2. Click the red annotation in the upper-left corner of the viewport and select 3D or MIP render mode. The selections available are dependent on the view type. MIP is used most often in CTA models to demonstrate the most intense voxels. 3. Click the [Oblique Mode] icon . 4. Click and drag the yellow line on the image. 5. Place the cursor over the red thickness annotation to do one the following options. ○ Type the desired thickness and press <Enter>. ○ Middle-click and drag to the desired thickness.
GCTREV225 | PROFILE_VIEW | scan parameters,display,view | profile view | revolution ct | How to create a profile view? | A profile graph shows the voxel value along a 3D trace (profile). This allows you to analyze the voxel value distribution of a 3D object in various ways, as an aid in setting up 3D processing (e.g., thresholding). 1. Open Reformat. 2. Select the Profile Layout preset, if available. 3. In one of the viewports, right-click the view type active annotation and select Profile. 4. On an axial, sagittal, or coronal image, press and hold <Shift> as you deposit points along the anatomy. As you trace, the profile view displays the pixel intensity along the trace. 5. On the profile view: ○ The horizontal axis is the position in millimeters along the trace and the vertical axis is the pixel intensity values as a function of that position. ○ Press <Shift> to display the pixel intensity from the location of the 3D cursor. Click and drag the white line to move it, which in turn, moves the cursor on the image. ○ View the displayed statistics.
GCTREV226 | VR_IMAGE | scan parameters,display,view | vr image | revolution ct | How to display vr image? | Use these steps to display a Volume Render image and the VR tab on the Tools screen. 1. Open Reformat. 2. Right-click the View Type active annotation and select VR.3. From the Tools section, click the VR tab
GCTREV227 | ATTACH_OR_DETACH_OBJECT | scan parameters,display,view | attach or detach object | revolution ct | How to attach or detach an object? | 1. Display a VR image. 2. Click the [Advanced VR Settings] icon . Read the Advanced VR Help on the tab for more information. 3. From the VR Controls screen Active Color area, select an active color. 4. From the Ramp area, click the [Attach Mode] icon 5. From the Active Color area, click [Attach]. 6. Click and drag the colored box at the top of the ramp to adjust the opacity of the colored voxels. 7. Repeat the steps to attach additional colors.
GCTREV228 | CUT_PLANES | scan parameters,diplay, view | cut planes | revolution ct | How to use cut planes? | When structures of interest are hidden on a 3D or VR view, use cut planes to display them prominently. 1. Display a VR image. 2. On a reformatted image, move the 3D cursor to the location of where you want the cut plane reference point. 3. Right-click the No Cut active annotation and select a plane to cut the 3D view. 4. Rotate the view to 3D view to display structures of interest in cut area. 5. Select Cut planes off to deactivate.
GCTREV229 | ADD_COLOURS | scan parameters,display,view,colours,reformat | add colours | revolution ct | How to add colours? | 1. Place the 3D cursor on the anatomy of interest. 2. From the Active Color area, click [Add].
GCTREV230 | ASSIGN_A_RANGE_OF_COLOURS | scan parameters,display,view,colours,reformat | assign a range of colours | revolution ct | How to assign a range of colours? | 1. Select Color. 2. From the Transition area, select [Step]. 3. Click and drag the colored boxes to adjust the colors to the desired value
GCTREV231 | CHANGE_A_COLOUR | scan parameters,display,view,colours,reformat | change a colour | revolution ct | How to change a colour? | 1. Select Color. 2. From the Active Color menu, select a color or click More Colors to display the color wheel. 3. Optional: On the color wheel, click and drag the circle to the desired color. 4. Click [Done].
GCTREV232 | VR_OPACITY | scan parameters,display,view,colours,reformat | vr opacity | revolution ct | How to refine vr opacity? | Use this procedure to refine the VR opacity in an image. 1. Display a VR image. 2. From the VR tab, click the [VR Settings] icon 3. Place the 3D cursor on the anatomy of interest in any multi-planar view. 4. From the VR Settings panel, click [AutoFit] or press <A>. 5. Middle-click and drag the mouse up and down to refine the VR opacity. 6. Click [Save New VR Preset] to save the setting as a new VR preset.
GCTREV233 | MULTI_VR_OBJECT | scan parameters,display,reformat | multi vr object | revolution ct | How to create a multi vr object? | You can merge multiple objects into a multiple VR model into a single view or model to perform more complex VR views with multiple object segmentation and visualization. 1. Open Reformat. 2. From the Tools section, click the Segment tab 3. Click the [Auto Select] icon . 4. Select the desired structure. ○ Small Vessels - segments vessels below a 5 mm diameter. ○ Any Structure - segments vessels greater than a 5 mm diameter and soft tissues growing from a seed point. ○ Bones - segments bony structures growing from a seed point. 5. Click [Add]. 6. Click [Yes] to clear the upper-left viewport. 7. From any 2D view, click and hold on the object you want to add. A green filter will fill the object while this object is being reconstructed in the upper-left viewport. 8. In a viewport, right-click and select Mouse Modes > Left Mouse: Selection. 9. Drag the VR viewport and drop it on the viewport of the isolated object where the message, Drop here to reassign views displays. 10. From the Tools section, click the VR tab. 11. From the VR tab, click the [Multi Objects] icon . 12. On the Multi-Object panel, hover over the objects. The corresponding object appears in the left window. 13. Select an object to adjust its transparency. 14. Click the [Eye] icon to hide or show the corresponding object on the active viewport
GCTREV234 | OPACITY_RAMP | scan parameters,display,view,reformat | modify the opacity ramp | revolution ct | How to modify the opacity ramp? | 1. Display a VR image. 2. From the VR tab, click the [Advanced VR Settings] icon Read the Advanced VR Help for more information. 3. From the VR Controls screen, perform any of the following functions in the Transparency area to modify the opacity ramp. ○ Click and drag the white boxes left or right to independently modify the upper or lower range of the ramp. ○ Click and drag the lower white box to the right to remove soft tissue or noise from an image. ○ Click and drag the white box at the top up or down to modify the overall opacity of the rang. This affects the opacity of all visible voxels. ○ Click and drag the blue box left or right to change the upper and lower values applied to the ramp
GCTREV235 | ZOOM_IN_AND_OUT | scan parameters,display,view,reformat | zoom in and out | revolution ct | How to zoom in and out? | 1. Display a VR image. 2. From the VR tab, click the [Advanced VR Settings] icon Read the Advanced VR Help for more information. 3. From the VR Controls screen, click and drag the red box in the lower window (which represents the scale of voxel values displayed) left or right to shift the display. 4. Click a Zoom icon to zoom in and out of the image. ○ Zoom In ○ Zoom Out ○ Zoom Display - Displays a red mask for you to define a range of interest. 5. Right-click to return to the original histogram range.
GCTREV236 | COMBINE_SEGMENTED_OBJECTS | scan parameters,display,view,reformat | combine segmented objects | revolution ct | How to combine segmented objects? | Use this procedure to combine 3D segmented objects together. 1. Open Reformat. 2. Select and view the images to be used for combining objects. All objects must originate from the same master volume. 3. From the Tools section, click the Segment tab. 4. Perform segmentation to generate objects to combine using any segmentation technique (Threshold, Auto Select, Paint, etc.). 5. Display objects to combine in separate views. 6. In a viewport, right-click and select Mouse Modes > Left Mouse: Selection. 7. Double-click to isolate first view of interest. 8. Single-click to isolate second view of interest. ○ Views of interest display red and green borders. Other views should not have color borders. ○ To adjust red and green borders, click the view to set red, the other view will turn green. 9. Select an operation (results of the operation display in the red views). 10. Click the [Advanced Processing] icon. 11. From the Advanced Processing panel, perform one of the following functions. ○ Click the [Subtraction] icon to remove all the voxels in the primary view that also exist in the secondary view. Before subtracting, make sure the borders of the view are displayed as: Red - Green = Red. ○ Click the [Intersection] icon to keep only the voxels that exist in the same location in both objects. ○ Click the [Addition] icon to keep all the voxels that exist in either of the objects.
GCTREV237 | PAINT_ON_SLICES | scan parameters,display,view,reformat | paint on slices | revolution ct | How to paint on slices? | Use this procedure to draw contours of the structure of interest on different slices. The volume to keep will be interpolated based of the defined contours. 1. Open Reformat. 2. From the Tools section, click the Segment tab. 3. From the Segment panel, click the [Paint On Slices] icon . 4. Move the 3D cursor to the edge of the feature of interest. 5. Optional: Click [Edge Attraction] to automatically refine the drawn contours, adjusting them to nearby structure edges. 6. Press <Shift> and click to deposit the cursor. 7. Click and drag the 3D cursor to define the area. When the mouse button is released, the first and last points connect. 8. Move to the next slice on which you want to paint. ○ It is not necessary to define contours on every slice. ○ The contours interpolates to the intermediate slices. 9. Repeat the process until your reach the last slice containing the structure to contour. 10. Click [Apply]
GCTREV238 | ANATOMY_WITH_AUTOSELECT | scan parameters,display,view,reformat | anatomy with autoselect | revolution ct | How to add or remove anatomy with autoselect? | Use this procedure to automatically remove anatomy from reformatted images. 1. Open Reformat. 2. From the Tools section, click the Segment tab. 3. From the Segment panel, click the [Auto Select] icon . 4. Select a tool for segmenting. ○ Click [Small vessel] for vessels less than 5 mm. Click once to fill and track vessel. ○ Click Any Structure for vessels greater than 5 mm or soft tissue. Click and hold until area of interest is filled. ○ Click [Bones] to segment bones. ○ Click [Pick from VR]. 5. Click [Add] or [Remove] depending on the desired outcome. 6. Scroll through the axial images at minimum slice thickness. If any vessels are contoured in green, put the cursor on the missing vessel in the 2D view and click and hold until the area of interest is filled. 7. Repeat steps to complete all missing vessels
GCTREV239 | REMOVE_FLOATERS | scan parameters,display,view,reformat | remove floaters | revolution ct | How to remove floaters? | Filter Floaters removes small residual objects in the 3D model that can appear after thresholding, usually resulting from noise in the original image set. 1. Open Reformat. 2. Perform a segmentation. 3. From the Tools section, click the Segment tab. 4. From the Segment panel, click the [Advanced Processing] icon. 5. Click the view of interest and click [Filter Floaters].6. From the Filter screen, click [Filter Size] and select a filter size. ○ Small ○ Medium ○ Large ○ Custom 7. Click [Apply Filter] to remove objects smaller than the filter size.
GCTREV240 | THRESHOLD_AN_IMAGE | scan parameters,display,view,reformat | threshold an image | revolution ct | How to threshold an image? | Use this procedure to threshold the image and keep only voxels within a specified range of value. 1. Open Reformat. 2. From the Tools section, click the Segment tab. 3.From the Segment tab, click the [Threshold] icon . 4. Select the desired image from which you want to threshold. 5. From the Threshold panel, use one of the following methods to adjust the threshold range of voxel values. ○ Click and drag the Min. and Max. Threshold sliders. ○ Type a minimum and maximum value in the Min. and Max. fields. ○ Click [Bone] or [Air] to apply preset threshold values. 6. Click [Apply Threshold] to display only the part of the 3D volume with voxel values inside the set range. 7. To further refine segmentation, select the object on the viewport and choose to remove or keep the object. ○ Click [Remove Object] to remove all voxels connected to the object. ○ Click [Keep Object] to remove all objects not connected to the object
GCTREV241 | BATCH_OBLIQUE | scan parameters,display,view,reformat | batch oblique | revolution ct | How to setup a batch oblique? | Use Batch Oblique to define a set of oblique planes adjusting settings (range, angle, position) from a grid on a reference image. 1. Open Reformat. 2. From the Tools section, select the Film/Save tab. 3. From the Film/Save tab, click the [Batch] icon. 4. From the Batch panel, click the view to be used as the reference image. 5. Click the Oblique tab to display a grid used to define the oblique planes. 6. Type a value for the Number of Images, Spacing Between Images, Slice Thickness, Mode (Average, MIP, MIN, VR), and FOV. 7. Select the desired output mode from the menu. ○ Choose Print to send generated images to the default printer (setup the Format and the display of a Reference Image). ○ Choose Save to save generated images in a new series in the browser. ○ Choose Print/Save to film the generated images from a default printer and save them in a new series. 8. Select the desired Apply to mode from the menu. ○ Current Volume – Apply to the displayed volume. ○ Phases of this Series or Series of this Exam – Apply to all phases in a series selected or all series in the exam selected. ○ All Volumes of all Exams – Apply to all volumes of all exams selected. ○ Choose Current Volume to apply to the displayed volume. ○ Choose Phases of this Series or Series of this exam to apply to all phases in a series selected or all series in the exam selected. ○ Choose All Volumes of all exams to apply to all volumes of all exams selected. 9. Click [OK]. The Spacing between views and slice thickness can be set independently of each other creating a gap, contiguous or overlapped images
GCTREV242 | BATCH_FILM_IMAGES | scan parameters,view,display,reformat | batch film images | revolution ct | How to batch filfm images? | Use this procedure to set up a batch rotation of images to film/save. 1. Open Reformat. 2. From the Tools section, click the Film/Save tab. 3. From the Film/Save tab, click the [Batch] icon. 4. From the Batch panel, click the [Rotate] tab.
GCTREV243 | 360_DEGREE_3D_IMAGES | scan parameters,view,display,reformat | 360 degree 3d images | revolution ct | How to rotate 360 degree 3d images? | 1. Click the appropriate arrow to indicate rotation direction. 2. Type in the number of images or angle between images in the appropriate field. Modifying one of these parameters updates the other. 3. Type in the FOV. 4. Select the desired output mode from the menu. ○ Choose Print to send generated images to the default printer (setup the Format and the display of a Reference Image). ○ Choose Save to save generated images in a new series in the browser. ○ Choose Print/Save to film the generated images from a default printer and save them in a new series. 5. Click [OK].
GCTREV244 | RADIAL_PLANES_ON_2D_VIEWS | scan parameters,view,display,reformat | radial planes on 2d views | revolution ct | How to place radial planes on 2d view? | 1. From the Protocols menu, select Fan. 2. Adjust the center of radial slices (A), first radial plane location (B), and the number of radial planes (C) on the 2D reference image. 3. Enter the Number of Images, Spacing between Images, Slice Thickness, Rendering Mode, and FOV. 4. Select the desired output mode from the menu. ○ Choose Print to send generated images to the default printer (setup the Format and the display of a Reference Image). ○ Choose Save to save generated images in a new series in the browser. ○ Choose Print/Save to film the generated images from a default printer and save them in a new series. 5. Click [OK].

## Sheet3
TID (unique) | KEYWORD(unique) | TAG(Separate with comma) | CATEGORY | PID(Separate with comma if more than one) | QUESTION | ANSWER | Documentation link | YouTube solution link
GPRXi1 | DAMAGE_OF_THINGS | damages,measures,xray fluroscopy | radiography and fluroscopy | Precision RXi | What are the measure to be taken to avoid damage of things? | • Before commanding any movements, make sure that no obstacle is present within the movement area. The dangerous areas of your equipment have already been indicated as “areas of crash”. • No object or expendable supplies must be placed on the tabletop, the SFD or the bars. As a matter of fact, during the movements of the SFD or the tabletop, these could crush such objects. • Do not place any objects on any other surface; In case of tilting, these objects could fall, thus involving possible damages. • Do not climb up any point of the tabletop cover or the SFD. The covers could get damaged and cause failures or bad functions of the equipment. • Do not seat on the column or the stand-support when the equipment is in the vertical position. This is prohibited since such a weight could determine damages to the material. • In no case you shall put glasses for contrast liquids (or other objects) on the equipment, the cabinet or the control panel. The contrast liquid could penetrate inside some components of the equipment causing damages to the equipment itself or wrong interpretations of the films. We suggest you should immediately eliminate possible liquid stains. Before the examination, execute a check panoramic without object with the largest possible size.
GPRXi2 | MONITOR _CART_TILTING | monitor cart,tilting,measures,radiography and fluoroscopy | radiography and fluoroscopy | Precision RXi | What are the measures to avoid monitor cart tilting? | The monitor cart is provided with two or three revolving wheels and two fixed wheels. These two ones can be blocked by means of levers. To avoid monitor cart tilting it is important not to push it with force when wheels are fixed or in case of obstacles to be overcome on the floor such as boards or steps.
GPRXi3 | PROTECTION_AGAINST_RADIATION | Radiation protection, recommendations,radiography and fluoroscopy | radiography and fluoroscopy | Precision RXi | What are the reccomendations for protection against radiation? | • You shall collimate the radiated field to the minimum possible value. • You shall provide the patient with the best possible protection by using specific screens in leaded rubber (in compliance with the requirements established by IEC 61331-3), if exposures are to be executed next to the patient's genital organs. • If it is necessary the presence of a doctor or auxiliary personal near the patient during the X- ray exams, they must wear the specific protective garments (lead aprons, gloves/mittens, collar etc.) in compliance with the requirements established by IEC 61331-3. • If it is necessary the presence of a doctor or auxiliary personal near the patient during the X- ray exams, they must occupy only the Significant zone of occupancy shown in Figure 1-2-3 In figures 1 2 and 3 are indicated the X ray radiations value measured in the Significant zone of occupancy and related to Fluoroscopy modality. These values measured on some tables as a sample, are useful to determine the individual protection that must be used, against radiations. • You shall wear your personal control badge and, if necessary, a pen dosimeter. • You shall keep a proper distance compared with the X-ray beam: the exposure should be commanded from the control panel which is normally placed behind the anti-X protection wall; in this way the necessary distance to ensure an adequate protection is automatically respected.
GPRXi4 | CHECK_AUTOMATIC_COLLIMATION | automatic collimation,procedures,checking,radiography and fluoroscopy | radiography and fluoroscopy | Precision RXi | How to check for automatic collimation? | • Insert a 35x43 cm cassette in the spot film device; • Move the focal distance to 115 cm; • Select a size division and verify that the collimator automatically sets to the selected value; • Vary the focal distance and verify that the selected size keeps to the right value; • Replace the cassette in the spot film device by a 24x30 cm cassette and repeat the test at a focal distance of 115 cm for at least two size subdivisions; • Check that for each selected division the size is kept unchanged even if you change the focal distance.
GPRXi5 | SELECTION_OF_TOMOGRAPHIC_LAYER_HEIGHT. | tomographic layer height,selection,radiography and fluoroscopy | radiography and fluoroscopy | Precision RXi | What are the procedures for the selection of tomographic layer height? | These two push-buttons (5 and 23) allow you to choose the tomographic layer height, with respect to the virtual fulcrum (maximum value) around which the rotation of the “X- ray tube/Image receptor” group is executed for the X-ray incidences or oblique radioscopy incidences. • The layer height is displayed • The upper push-button (5) moves the layer height to the higher values. • The lower push-button (23) moves the layer height to the lower values. The layer height selection ranges from 0 to 330 millimeters
GPRXi6 | REVIEW _IMAGE | image memorisation function,review | radiography and fluoroscopy | Precision RXi | How to review or image memorise? | ush button (22) functions: • During fluoroscopy, push button (22) starts/stops image sequence memorization at the speed previous selected by push button (33) • During review function, the push button (22) starts/stops sequence image display at the speed previous selected by push button (33).
GPRXi7 | IMAGE_INVERSION_ON_THE_MONITOR | image inversion,monitor | radiography and fluoroscopy | Precision RXi | How to invert an image on monitor? | • Right-to-left image inversion (34) • Up-to-down image inversion (35)
GPRXi8 | FOCAL_DISTANCE_CONTROL | focal distance,ctrl | radiography and fluoroscopy | Precision RXi | How to ctrl focal distance? | • These push-buttons (36, 37 and 38) allow the possibility to move the focal distance toward the value indicated on the selected push-button. • When the focal distance is at a value between those indicated on the buttons, the three LEDs will blink. • When the focal distance is at exact value indicated on a button, the LED on that button will be lit. • The scale on the compressor column always indicates the focal distance.
GPRXi9 | SIZE_DIVISION_CHOICE | size division,choice | radiography and fluoroscopy | Precision RXi | What are options for size division? | The SFD transversal division options allow the division of any size in the transversal direction; • The SFD crossed division options allow the crossed subdivision in 4 and 6 for the 18x24, 24x24, 24x30 cassette sizes. > Button (43) on the left is always activated when the equipment is switched on an
GPRXi10 | IMAGE_INTENSIFIER_ELEVATION | image intensifier,elevation | radiography and fluoroscopy | Precision RXi | What are procedures for elevating image intensifier? | • The normal condition is the best possible condition for geometrical magnification and resolution. • The efficiency condition minimizes the transition time from fluoroscopy to radiography.
GPRXi11 | COLLIMATOR_MODE | modes,collimator | radiography and fluoroscopy | Precision RXi | What modes can be opted for collimator mode? | • The normal condition is the best possible condition for geometrical magnification and resolution. • The efficiency condition minimizes the transition time from fluoroscopy to radiography.
GPRXi12 | LONGITUDE TABLE _TOP_DISPLACEMENT | displacement,table top,longitudinal | radiography and fluoroscopy | Precision RXi | How to displace table top longitudinally? | • The lateral push button (57) moves the table top towards left side. • The lateral push button (59) moves the table top towards right side. • The central push-button (58) is used to place the tabletop in its original centered position. The LED will be lit when the table reaches the centered position. • If the tilting angle is 0° the maximum tabletop head side travel is 86 cm. • If the tilting angle is different from 0° the maximum tabletop head side travel is 47.5 cm. • The maximum tabletop feet side travel is 15 cm.
GPRXi13 | COMPRESSOR_OPERATION | compressor pressure,compressor,operation | radiography and fluoroscopy | Precision RXi | How to operate compressor and ctrl pressure? | • Push the button (62a) to lower and activate the compressor • Push the button (62b) to raise the compressor As soon as the compressor reaches the bottom position the following movements will be inhibited: • tabletop transversal displacement • tabletop longitudinal displacement • Stand and spot film device (mobile area) longitudinal displacement • Stand inclination or centring • Focal distance decrease
GPRXi14 | XRAY _TUBE_STAND_ANGULATION | angulation,tube stand,xray | radiography and fluoroscopy | Precision RXi | How to angulate xray tube stand? | • Push (64a) to move the X-ray tube stand to the left with subsequent rotation of the tube for the execution of left oblique incidences. • Push (64b) to move the X-ray tube stand to the right with subsequent rotation of the tube for the execution of right oblique incidences
GPRXi15. | CONTROL_OF_PATIENT_SCANNING_MOVEMENTS | patient scanning, movements,controlling | radiography and fluoroscopy | Precision RXi | How to ctrl patient scanning movements? | • Tipping the joystick to the right will move the tabletop the internal side, corresponding to right side of the image on the monitor. • Tipping the joystick to the left will move the tabletop the external side, corresponding to left side of the image on the monitor. • Tipping the joystick up will move the I.I.-Tube Stand group to the left, corresponding to the top of the image on the monitor. • Tipping the joystick down will move the I.I.-Tube Stand” group to the right, corresponding to the bottom of the image on the monitor.
GPRXi16 | OPTIONAL_ON_BOARD_CONSOLE_WITH_JOYSTICK | board console joystick ctrl,optional | radiography and fluoroscopy | Precision RXi | How to ctrl board console using joytick option? | • Tip the lever to the right and the tabletop will move towards the internal side, corresponding to the right side of the image on the monitor. • Tip the lever to the left and the tabletop will move towards the external side, corresponding to the left side the image on the monitor. • Tip the lever up and the I.I.-Tube Stand group will move left along the longitudinal axis, corresponding to the top of the image on the monitor. • Tip the lever down and the I.I.-Tube Stand group will move right along the longitudinal axis, corresponding to the bottom of the image on the monitor
GPRXi17 | LIMIT_XRAY_BEAM_FIELD | beam field,xray,limiting | radiography and fluoroscopy | Precision RXi | How to limit xray beam field? | • The knob on the left (L) commands the diaphragms in the width direction (parallel diaphragm); turn clockwise to close them and counter-clockwise to open them. • The knob on the right (R) commands the diaphragms in the height direction (perpendicular diaphragms): turn clockwise to close them and counter-clockwise to open them.
GPRXi18 | XRAY_CASSETTE_TECHNIQUE | xray Cassette techniques | radiography and fluoroscopy | Precision RXi | How to use xray cassettes? | • Around each adjustment knob (B) there is a graduated scale (C) used to indicate the field size setting. • The table on the front panel (D) shows the settings for various focal distances. Find the row containing relevant focal distance in the far left (cm) or right (inch) column and the column containing the desired field size in the top (cm) or bottom (inch) row; the intersecting value is the corresponding collimator field size setting. • Move the knob to the position on the scale for the field size setting found.
GPRXi19 | POINTER_DEVICE | graphic user interface, xray,ctrl,pointer devices | radiography and fluoroscopy | Precision RXi | How to use pointer device for controlling graphic user interface? | • Cursor Movement (cursors are detailed in part 3.1.19.6 of this section) • Single-click selection on menus and screens • Double-click selection on menus and screens • Select and Drag items on screens • Right click selection for alternate menus and screens (see section 3.1.19.2).
GPRXi20 | CURSOR | cursor ,selection,xray | radiography and fluoroscopy | Precision RXi | How to select cursor? | • When selecting the cursor, do not drag the cursor, just click and move the mouse. • Once over the image, click and hold-down the left mouse button – this is called dragging – in the direction needed. • See chart below for full descriptions.
GPRXi21 | EXPOSURE _HOLD_MODE | modes,exposure hold | radiography and fluoroscopy | Precision RXi | How to set exposure hold mode? | in cases of incorrect operation conditions. • When the system is in the automatic mode and the shutter motors are operating to reach the selected position. • If the system fails to reach the selected position. • The LED is lit and the output contact is activated. • The exposure contact is open
GPRXi22 | SWITCHING _ON_THE_SYSTEM | switchON,power,start | radiography and fluoroscopy | Precision RXi | How to start the system? | 1 Turn ON the power for the positioner. 2 Turn ON the power for the x-ray generator. 3 Turn ON the power for each of the system’s Image Monitors and other peripherals. 4 Turn ON the power for the system PC (switch is on the front panel). The PC will begin to run its power-up self-testing routine. 5 Once the system has completed the Login sequence, the Main Image screen will be displayed.
GPRXi23 | SWITCHING _ON_THE_DIGITAL_SYSTEM | digital system,switch on,get started | radiography and fluoroscopy | Precision RXi | How to work with digital system? | • Turn ON the power for the system PC (switch is on the front panel). The PC will begin to run its power-up self-testing routine. • Turn ON the power for each of the system’s image monitors and other peripherals if required • If auto-login does not occur and the login screen appears, type “operator” in the user field and press “Enter” (no password needed) to complete the log-in procedure. • Once the system has completed the Login sequence, the Main Image screen will be displayed.
GPRXi24 | CLOSE_SYSTEM | stoppung system | radiography and fluoroscopy | Precision RXi | How to stop the system? | Page 119 4.2. Switching off the system 4.2.1. Close the Precision RXi program Normal Daily Operation / System Maintenance: 4 The Precision RXi system will safely and automatically close the Clinical Application software program. The computer is still considered to be ON.
GPRXi25 | SWITCH_OFF_THE_DIGITAL_SYTEM | digital system,switching off | radiography and fluoroscopy | Precision RXi | How to switch off the digital system? | • Close the imaging system’s program. If necessary, follow the instructions in par 4.2.1 to close the program, and then return to these instructions. • Select the button (lower left corner of screen). A submenu opens. • Select the option. A new screen • Select “Shut down” from the pull-down list of options. • Select the OK button from this screen. • The PC will automatically power down.s.
GPRXi26 | SWITCH _OFF_GENERATOR. | generator,switching off | radiography and fluoroscopy | Precision RXi | How to switch off generator? | Exit the current screen(s) to return to the main console menu. Press to switch the X-ray generator OFF. The touch screen will power off immediately
GPRXi27 | SWITCH_OFF_THE_POSITIONER | positioner,switching off | radiography and fluoroscopy | Precision RXi | How to switch off the positioner? | Use the large red push-button placed on the right upper corner of the console to de-energize all the power circuits and reset the microprocessor’s circuits.
GPRXi28 | DAILY_XRAY_WARMUP_PROCEDURES | procedures warmup,xray,daily | radiography and fluoroscopy | Precision RXi | How to warm up xray daily? | 1 Choose the tube to be warmed-up by selecting an image receptor that is programmed to select the desired tube. 2 Close the collimator blades. 3 Select the following: • 80 kV. • Large focal spot. 4 For a 200 to 300 kHU tube use approximately 160 mAs per exposure. For a 300 to 400 kHU tube use approximately 200 mAs per exposure. Depending on the X-ray tube power rating,
GPRXi29 | SEMI _AUTOMATIC _XRAY_ WARMUP | semiautomatic, warmup,daily | radiography and fluoroscopy | Precision RXi | How to warmup xray semiautomatically ? | 1 From the main console menu, press TUBE WARMUP. 2 Close the collimator blades. 3 Choose the tube to be warmed-up by selecting an image receptor that is programmed to select the desired tube. 4 Make the first exposure. 5 The console will time a 30 second interval, and then allow the next exposure. 6 Perform the second and third exposures when permitted. 7 Warm-up the second tube if desired by repeating steps 2 to 5. 8 Select MAIN MENU to return to the main console menu. The system is now ready for use.
GPRXi30 | REMOVE_COMPRESSOR_MANUALLY | manual,removing,compressor | radiography and fluoroscopy | Precision RXi | How to remove the compressor manually? | 1. You shall take the release lever [B] (Figure 1 and Figure 3) and place it so as the Velcro results to be on the cable-support turret [H] (Figure 1) 2. You shall insert the hexagon [M] of the lever (Figure 3 and Figure 2) inside the screw [L] protruding from the back cover (Figure 3 and Figure 2) Figure 1 Figure 2 Figure 3 3. You shall lift the compressor by means of the lever [B] (Figure 3) as shown in Figure 4 and Figure 5: you shall turn the lever counterclockwise if you are standing behind the back side of the equipment.
GPRXi31 | INSERT_XRAY_CASSETTE | insertion,xray cassette,cassetre holder ,table. | radiography and fluoroscopy | Precision RXi | How to insert xray cassette in cassette table holder? | 1) You shall press the push-button controlling the opening of the SFD cassette-holder: this push-button is placed on the console or next to the cassette-holder itself. 2) You can now insert the X-ray cassette by pushing the two external 3) You shall therefore settle gently the cassette on the cassette-tray till the cassette itself goes beyond the external guide. 4) You shall now fit the X-ray cassette into the guide paying attention to your fingers position. 5) You shall adjust the “Right” and “Left” indicator if necessary..
GPRXi32 | ANATOMICAL_PROGRAMMING_SELECTOR | select,anatomical ,programming | radiography and fluoroscopy | Precision RXi | How to use anatomical programming detector? | 1 Select the region of interest by pressing the appropriate area on the anatomical figure on the touch screen display (i.e. ABDOMEN). The selected area will change to yellow and a pop-up window will appear on the display. • The caption on the pop-up window indicates the selected anatomical region (i.e. ABDOMEN). • The pop-up window will display one or more menu, procedural, or APR items. A menu item will have an M beside the name, a procedural item will be identified with a P beside the name, and an APR item will have an I beside the name. * A menu item (i.e. ABDOMEN) has one or more submenus. The submenu(s) may include additional menu items, procedural items, and / or APR items. * A procedural item (i.e. FOOT) has one submenu. That submenu will contain only APR items. * An APR item (i.e. FOOT AP) has no submenus. Selecting the desired APR item will set predefined radiographic or fluoroscopic protocols in preparation for an exposure. 2 If the appropriate APR item (I beside the item name) is shown on the selected menu, select that item. The selection will be highlighted, predefined radiographic and / or fluoroscopic parameters and techniques will be set in preparation for an exposure, and the patient size will default to medium. • To change the patient size, press the corresponding button. * = pediatric. * = small. * = medium. * = large. • If the assigned parameters and techniques are appropriate, an X-ray exposure may be taken. • The default settings may be changed, if desired, before making an exposure. Please review the sections RADIOGRAPHY CONTROLS AND DISPLAY and FLUOROSCOPY CONTROLS AND DISPLAY before making any changes. Changes made are not saved to memory. To make and then save APR changes, refer to section SYSTEM UTILITY – APR EDITOR3 If the appropriate procedural item (P beside the item) is shown on the selected menu, select that item. A submenu will appear showing one or more APR items within that procedure. APR items within a procedural menu will be automatically selected in sequence. For example, assume that the procedural item FOOT contains the three APR items FOOT AP, FOOT LAT, and FOOT OBL. • The first item on the list (i.e. FOOT AP) is automatically selected. The selection will be highlighted, and predefined parameters and techniques will be assigned for that study. • An X-ray exposure may be made as per step 2. • When the exposure is finished, the APR will advance to the next APR item (i.e. FOOT LAT). That selection will be highlighted, and predefined parameters and techniques will be assigned, readying the generator for that study. • Another X-ray exposure may be made. • If an exposure needs to be repeated (i.e. FOOT LAT), reselect that item, and repeat the exposure. • The generator will advance to the last item (i.e. FOOT OBL). The last X-ray exposure may be made. 4 If the appropriate menu item, i.e. SKULL, (M beside the item) is shown on the selected menu, select that item. A submenu will appear showing one or more menu, procedural, or APR items within that category. • Menu items may be nested. It may be necessary to drill down through several submenus to find the desired procedural or APR item. • The appropriate procedural item or APR item may be selected as per the previous steps. • Menu, procedural, and APR items may be programmed to suit your requirements. Refer to section 4.5.4 for details. 5 Press the BACK button to back out of the APR menu structure. A new region of interest may be selected at any time by pressing the appropriate area on the anatomical figure on the touch screen display as in step 1. It is not necessary to back out of an existing menu first.
GPRXi33 | TESTING_DAP_DEVICE | testing,DAP device | radiography and fluoroscopy | Precision RXi | How to select DAP device to test? | 1 In order to select the DAP device to test, select a receptor that is programmed for the desired tube / DAP device (tube 1 or tube 2). This will also select the DAP to reset if programmed for individual dose-area product readings for each tube. • If the two DAP readings are summed, the reset function may be performed on either the tube 1 or tube 2 selection. The DAP display is common to both tubes, therefore the DAP display may be reset at either the tube 1 or tube 2 tube position. • If the generator is programmed for individual DAP measurements for each tube, and you are intending to reset the DAP display, confirm that the correct tube / DAP device is selected before continuing. 2 Press the DAP button at the bottom of the touch screen display. This will present the DAP reset / test / print menu. • Press RESET to reset the DAP display to zero. The display will reset, and the reset / test / print menu will close. • Press TEST to test the DAP. The message DAP Test [PASSED] will be displayed if the DAP has passed its test. Press OK to close the test status window. • REPEAT THE TEST OR RESET FUNCTION FOR THE OTHER TUBE IF DESIRED. • Press PRINT to open the DAP printer interface window. The DAP printer function is described in detail on the next page. • Making a selection from the reset / test / print window will automatically close the window.
GPRXi34 | PRINTER_INFORMATION | information entry ,before print | radiography and fluoroscopy | Precision RXi | How to enter information before printing? | • Enter the patient’s gender by pressing Male or Female. • Press Name to enter the patient’s name. An English keyboard will pop up at the bottom of the screen. A partial keyboard with international symbols corresponding to the installer-selected language may be displayed at the top of the screen. • When finished entering the name, press ENTER on the bottom keyboard. The touch screen will return to the printer interface window, and the name will appear on the graphic of the label to be printed. • Follow the preceding procedure to enter the ID, Tech Line 1, and Tech line 2. Tech line 1 and tech line 2 are typically used to enter the techniques as appropriate for the record. A maximum of 30 characters may be entered per line. • Enter the patient’s date of birth by pressing the up or down arrows to the right of the Month, Day, and Year. • Press Clear to clear the label, if desired.
GPRXi35 | CHECK _BEFORE_PRINTING | information, check,before printing | radiography and fluoroscopy | Precision RXi | How to check informations before printing? | • Enter the patient’s gender by pressing Male or Female. • Press Name to enter the patient’s name. An English keyboard will pop up at the bottom of the screen. A partial keyboard with international symbols corresponding to the installer-selected language may be displayed at the top of the screen. • When finished entering the name, press ENTER on the bottom keyboard. The touch screen will return to the printer interface window, and the name will appear on the graphic of the label to be printed. • Follow the preceding procedure to enter the ID, Tech Line 1, and Tech line 2. Tech line 1 and tech line 2 are typically used to enter the techniques as appropriate for the record. A maximum of 30 characters may be entered per line. • Enter the patient’s date of birth by pressing the up or down arrows to the right of the Month, Day, and Year. • Press Clear to clear the label, if desired.
GPRXi36 | SYSTEM _UTILITIES | system utilities | radiography and fluoroscopy | Precision RXi | What are the advantages of system utilities? | • Press Clear to cancel an incorrect password. • Press Cancel to return to the main console menu. • Press Accept to access the system utilities menu. After a brief delay, the system utilities screen will be displayed. The message Access Denied indicates that an incorrect password was used.
GPRXi37 | APR_EDITOR | APR editor,uses | radiography and fluoroscopy | Precision RXi | How to use APR editor? | 1 From the system utilities menu, press APR Editor. A screen similar in appearance to the normal operating screen will be displayed. However, as a reminder that you are in APR editor mode, the word APR EDITOR will be displayed in the APR window. 2 Refer to the applicable subsections (following) for the procedures to change parameters and technique for existing APR items, and to edit, add, or delete APR, procedural, or menu items.
GPRXi38 | APR_EDITING_PARAMETERS | changing parameters,existing APR item, APR editor | radiography and fluoroscopy | Precision RXi | How to change parameters for existing APR item using APR editor? | 1 Select the APR item to be changed. If the APR item is in a submenu of a menu or procedural item, select the parent menu item or procedural item, then press Open Sub-Menu. It may be necessary to drill down through several submenus to find the desired APR item. 2 When the selected APR item is highlighted, select the patient size. The desired parameters / technique may be changed for that patient size (rad / fluoro kV and mA, mAs, ms, density, AEC / mA/ms / mAs, focal spot, film screen, AEC fields, image receptor, AEC lock, fluoro mode (continuous / pulsed fluoro / HLF), I.I. mode, ABS). Repeat for all patient sizes for that APR item as required. The highlight will change to red when the programmed APR has been altered. 3 Repeat steps 1 and 2 for other APR items within that menu or submenu, if applicable. 4 Press BACK when finished changing APR items in that menu / submenu. This may need to be done more than once to return to the top APR menu for the selected region of interest. A pop-up window will display asking if you wish to save the changes. Press YES to save the changes; CANCEL cancels the changes.
GPRXi39 | EDIT_ITEM_NAME | edit,APR item name | radiography and fluoroscopy | Precision RXi | How to edit name for an APR item? | 1 Select the APR item as per step 1 under “To change parameters or technique for an existing APR item”. The selection will be highlighted. 2 Press Edit. Keyboard(s) will display on the screen as described above, with the current name of the APR item highlighted. • Press CANCEL on the lower keyboard to exit without changing the name. • Press DEL to delete the highlighted name on the keyboard. • Type in the new name for that APR item. Use the BACKSPACE key to back space if corrections are needed. • Add / edit the ID code if desired (if displayed). • Press ENTER when finished.
GPRXi40 | DELETE _APR_ITEM | delete,APR item | radiography and fluoroscopy | Precision RXi | How to delete an APR item? | 1 Select the APR item to be deleted. The selection will be highlighted. 2 Press Delete. A pop-up window will display asking if you are sure you want to delete this item. Press YES to delete the item; NO cancels the deletion.
GPRXi41 | ADD_APR_ITEM | add,APR item | radiography and fluoroscopy | Precision RXi | How to add an APR item? | 1 Select the appropriate location to add the APR item. An APR item may be added directly to a main APR menu, or may be added to a submenu of another menu or procedural item. If the APR item is to be added to a submenu of a menu or procedural item, select the parent menu item or procedural item, then press Open Sub-Menu. It may be necessary to drill down through several submenus to find the desired location for the new APR item. 2 Press Add. A pop-up window will display allowing you to select three item types to be added. Select APR Item (this is the default selection, and the only available selection if adding to a procedural menu). 3 Press OK to continue. Cancel will cancel this action. 4 Keyboard(s) will display on the screen as described previously. • Type in the name of the new APR item. • Add the ID code if desired (if displayed). • Press ENTER when finished. • The new APR item will appear on the selected menu or submenu. • Change the parameters and technique as per the subsection “To change parameters or technique for an existing APR item”.
GPRXi42 | ADD_EDIT_DELETE_PROCEDURAL_ITEM | add,edit,delete, procedural item,APR editing | radiography and fluoroscopy | Precision RXi | How to add,delete or edit procedural item? | 1 To delete a procedural item, follow the steps in “To delete an APR item”. Doing so will also delete the submenu associated with that procedural item. 2 To edit a procedural item, follow the steps in “To edit (change the name of) an APR item”. 3 To add a procedural item, follow steps 1 to 3 in “To add an APR item”, except select APR Procedure in step 2. 4 Keyboard(s) will display on the screen as described previously. • Type in the name of the new procedural item. • Add / edit the ID code if desired (if displayed). • Press ENTER when finished. • The new procedural item will appear on the selected menu or submenu. • Select the newly added procedural item. The selection will be highlighted. • Press Open Sub-Menu. A “generic” APR item named FIRST ITEM has been automatically inserted in that submenu. • You may now edit the name and change the parameters and technique for that APR item, and add additional APR items for that procedural item as per previous steps
GPRXi43 | ADD_EDIT_DELETE_MENU_ITEM | add,edit,delete, menu item,APR editing | radiography and fluoroscopy | Precision RXi | How to add,delete or edit menu item? | 1 To delete a menu item, follow the steps in “To delete an APR item”. Doing so will also delete the submenu associated with that menu item. 2 To edit a menu item, follow the steps in “To edit (change the name of) an APR item”. 3 To add a menu item, follow steps 1 to 3 in “To add an APR item”, except select APR Menu in step 2. 4 Keyboard(s) will display on the screen as described previously. • Type in the name of the new menu item. • Add / edit the ID code if desired (if displayed). • Press ENTER when finished. • The new menu item will appear on the selected menu or submenu. • Select the newly added menu item. The selection will be highlighted. • Press Open Sub-Menu. A “generic” APR item named FIRST ITEM has been automatically inserted in that submenu. • You may now edit the name and change the parameters and technique for that APR item, and add additional APR items for that menu item as per previous steps.
GPRXi44 | APR_BACKUP_RESTORE | backup,restore,APR | radiography and fluoroscopy | Precision RXi | How to backup or restore APR? | 1 From the system utilities menu, press APR Backup / Restore. A pop-up window will display showing the available backup files and the factory-default APR files in various languages. Pressing Close will exit the backup and restore menu. 2 Refer to the applicable subsections (following) for the procedures to back-up the current APR data, and to restore saved APR data. The entire APR (parameters and techniques, APR text, menu structures, etc) is backed up.
GPRXi45 | BACKUP _CURRENT_APR | backup,current APR | radiography and fluoroscopy | Precision RXi | How to backup current APR? | 1 Press BACKUP. 2 Keyboard(s) will display on the screen as described previously. • Type in the name of the new backup file. • Press Enter when finished. The console will return to the APR backup / restore menu. • A pop-up window will display indicating that the backup was successful. Press OK to continue.
GPRXi46 | BACKUP _RESTORED_APR_DATA | backup restored APR data | radiography and fluoroscopy | Precision RXi | How to restore backuped APR data? | 1 Select the APR data file to be restored. The selection will be highlighted. 2 Press Restore. 3 After a brief delay, a pop-up window will display indicating that restore was successful. Press OK.
GPRXi47 | DELETE _APR_DATA_FILE | delete,APR data file | radiography and fluoroscopy | Precision RXi | How to delete APR data file? | 1 Select the APR data file to be deleted. The selection will be highlighted. 2 Press DELETE. A pop-up window will display asking if you are sure you want to delete this file. Press YES to delete the file; NO cancels the deletion.
GPRXi48 | RESTORE _FACTORY_DEFAULT_APRDATA | restore factory default,APR data | radiography and fluoroscopy | Precision RXi | How to restore factory default APR data? | 1 Press Factory Defaults. 2 After a brief delay, a pop-up window will display indicating that restore was successful. Press OK.
GPRXi49 | XRAY _EXPOSURE | xray exposure, radiography | radiography and fluoroscopy | Precision RXi | How to expose xray properly in radiography? | To take a radiography: 1) Centre the beam toward the bucky; 2) Insert the cassette into the bucky; 3) Centre and adjust the field on the patient; 4) Set the exposure values and technique on the generator; 5) Press the preparation button and, when the generator is ready for the exposure, press the exposure button. The grid in the bucky starts working and authorizes the exposure execution.
GPRXi50 | TUBE _WALLSTAND | use of tube, with wallstand | radiography and fluoroscopy | Precision RXi | How to use tube with a wall stand? | With the vertical travel of the tube and wallstand it is possible to take not only the thorax radiography, but also the lower limbs. To take a radiography with a wallstand: 1) Put the stand in the position indicated by the centering 2) Rotate the tube 90° on its support arm in order to align and centre the beam toward the wallstand. Centre and place the patient, insert the cassette in the wallstand, set the exposure technique on the generator and take the radiography.;
GPRXi51 | LOADING_CASSETTE | cassette loading | radiography and fluoroscopy | Precision RXi | How to load cassette prorperly? | To release the cassette tray push the small lever at the top of the cassette tray slot; the cassette tray can now be removed from the bucky and loaded. To load the cassette: 1) Using the proper handle pull out the cassette tray out of the bucky until the stop is reached; 2) Rotate clockwise the lever that unlocks the jaws that block the cassette; 3) Open the jaws enough to insert the cassette in use; 4) Position the device or the cassette vertical stop in the housing, in correspondence with the vertical side dimension of the cassette in use; 5) Insert the cassette in the jaws, leaning it on the vertical stop; 6) Rest the jaws against the cassette; 7) Rotate clockwise the lever to lock the jaws; 8) Insert again the cassette in the bucky, using the proper handle.
GPRXi52 | PERFORMING_XRAYS | xray performing, operations | radiography and fluoroscopy | Precision RXi | How to operate a xray to perform properly? | The operations to do to perform an x-ray are the following: 1) Center the x-ray source to the bucky 2) Center and collimate on the patient the area that is to be radiographed 3) Insert the cassette in the bucky 4) Set the values and the exposure technique on the generator user interface 5) Check to make sure everything is ready for exposure. 6) Instruct the patient: please take a breath and hold it! Trigger exposure. 7) Push the preparation button and, when the generator is ready for the exposure, push the exposure button. 8) The grid in the bucky starts moving and performs the procedure.
GPRXi53 | EXTERNAL _CASSETTE_HOLDERS | cassette holders,external,uses | radiography and fluoroscopy | Precision RXi | How to use external cassette holder? | The external cassette holder allows the user to perform x-rays without using the grid with cassettes placed outside the bucky. The cassette holder is directly fasten to the upper part of the frame that supports the bucky and leaned to its front panel. According to the maximum cassette size that they can get, they have different dimensions and weights that change according to the bucky vertical movement balancing conditions. So, it is necessary to always lead the bucky vertical movement when the brakes are un-locked. To insert the cassette in the external cassette holder do as follows: 1) Low the lower wing that holds the cassette by acting on the levers that un-lock its movement. 2) Insert the cassette in the upper wing of containing and lean it to the cassette holder. 3) Get close again the lower wing of containing to the lower side of the cassette. The lower wing stopping device in position releases automatically.
GPRXi54 | DELETE _SEQUENCE _OF _IMAGES | delete, image sequence | radiography and fluoroscopy | Precision RXi | How to delete an image sequence? | • A “Sequence of Images” is a series of images within a study. • Each sequence of images is shown as a Thumbnail on the Main Image screen. • Each sequence of images is numbered automatically as it is acquired. • When a sequence is deleted from a study, the original numbers for the remaining sequences will be retained.
GPRXi55 | SEQUENCE_NUMBERING | select sequence number | radiography and fluoroscopy | Precision RXi | How to select sequence numbering? | 1 Open the preferred patient file and select the Thumbnail button if not already active to display the sequence thumbnails. 2 Position cursor over the thumbnail representing the sequence to be deleted, and right click. 3 From the Right-click menu, select the Delete Sequence option. 4 A “D” will mark the sequence for deletion in the upper right corner of the Thumbnail image, and the right-click menu will close. 5 Repeat these steps to mark other sequences for deletion. 6 When the patient file is closed, a warning will appear detailing the number of images marked for deletion. • Select Yes to delete the marked sequences and close the patient file. • Select No to close the patient file without deleting any • Select CANCEL to cancel request to close the patient file (no images will be deleted)..
GPRXi56 | SORTING_THE_PATIENT_LIST | sorting,patient list,xray | radiography and fluoroscopy | Precision RXi | How to sort the patient list? | 1 In the Patient Information Area, the Patient List allows sorting by Patient Name, ID, DOB, Physician of Record or Study Date. To sort the database by a specific field, click on the Column’s Title. 2 In the Study Information Area, the Patient List allows sorting by Study ID, Study Date/Time, Referring Physician and Description. To sort the database by specific field, click on the Column’s Title.
GPRXi57 | OPEN_THE_EXISTING_PATIENT_FILE | opening,existing patient file | radiography and fluoroscopy | Precision RXi | How to open the existing patient file? | "1 From the Main Image screen, click on Patient List button. 2 The Patient List screen will appear (see below). Position cursor over any part of the preferred patient file and click the left mouse button. 3 Highlight the Study you desire to Open. 4 Click the OPEN button. The system will open the patient file and selected study." 5 To acquire images for this patient file select the button from the Main Image screen. 6 To Review images refer to Chapter 7 of this manual for more information on the review of function. 7 A prompt will appear - use the instructions that follow to either add images to the current study or start a new study.
GPRXi58 | ACQUIRE_ADD_NEW_IMAGES_TO_CURRENT_NEW_STUDY | acquire or add,new images ,current new study | radiography and fluoroscopy | Precision RXi | How to Acquire/Add New Images to Current Study? | 1 Select NO at the prompt. The screen will return to Main Image screen. 2 Select preferred acquisition rate from the Rate. The next acquired image will be added in the next available thumbnail slot for this study. 3 See Chapter 6 of this manual for more information on the Acquire functions and features.
GPRXi59 | ACQUIRE_ADD_NEW_IMAGES_TO_CREATE_NEW_STUDY | acquire or add,new images ,create new study | radiography and fluoroscopy | Precision RXi | How to Acquire/Add New Images to Create Study? | 1 Select YES at the prompt. 2 The Patient Information screen will open. Click on <New study>. 3 Fill in the new information for the study (right side of the screen). 4 Click Open. 5 See Chapter 6 of this manual for more information on the Acquire functions and features
GPRXi60 | PATIENT_ DATA _DICOM _FEATURE | import ,patient data,DICOM ,worklist, feature | radiography and fluoroscopy | Precision RXi | How to add Import Patient Data Using DICOM WorkList Feature | "Use the WorkList feature to search a database (usually a centralized database connected via a DICOM network) for patients based on specific criteria. Look for groups of patients, or a specific patient file. This feature is used to search by Patient Data only - not images" Accessing Worklist 1 From the Main Image screen, click on the Patient List button. 2 Select the Worklist tab. The Worklist screen will be displayed.
GPRXi61 | SETUP _AUTOMATIC _QUERY _SRRCH _PATIENT_ FILE | setup automatic query, search,patient ,file, | radiography and fluoroscopy | Precision RXi | How to Search for Patient File automatically | Auto Query can be used to search for Patient files from either the local system or any system that is linked to a DICOM network. The Worklist will contain all patients scheduled for the time period specified in the search. From the Worklist screen, select the button. 2 The Auto Worklist screen will be displayed.
GPRXi62 | EDITING_ EXISTING _QUERY _AUTO_QUERY _LIST | editing , existing ,query ,auto ,query, patient list | radiography and fluoroscopy | Precision RXi | How to Edit an Existing Query on the Auto Query List | Select the button from the Patient List screen – Worklist Tab to display the Auto Worklist screen. Click the Query button from the Auto Worklist screen Select a previously saved Query from the Auto Query List that requires editing. Then, complete all the steps in the Automatic Query Setup Procedure (Steps 1-8 on the Auto Worklist screen).5 Set the Time Parameters for running Automatic Queries.
GPRXi63 | ADD_NEW_PHYSICIANS | add,new physicians | radiography and fluoroscopy | Precision RXi | How to add new physicians? | 1 Select the Patient List button. 2 Select a new or existing patient file. 3 Click on Down arrow for either Performing or Referring Physician The resulting menu will show all physicians currently established as well as a <New Physician> option. 4 Select the <New Physician> option. 5 The Physician Name Editor screen will open.6 Enter Physician’s name, using drop down menu boxes for Prefix and Postfix if required. 7 Click on “Add New Physician” button. The doctor’s name is added to the list. 8 Repeat these steps to add other physicians to the list.
GPRXi64 | DELETE _A_PHYSICIAN | delete,physician | radiography and fluoroscopy | Precision RXi | How to delete any physician?. | 1 Select the Patient List button. 2 Select a new or existing patient file. 3 Click on Down arrow for either Performing or Referring Physician. The resulting menu will show all physicians currently established as well as a <New Physician> option. Select the <New Physician> option. The Physician Name Editor will be displayed. 4 Highlight the Doctor’s name from the list of those available. 5 Select the Delete physician button. 6 You will be prompted to confirm this selection. Click Yes to permanently remove the specified physician or No to cancel the transaction.
GPRXi65 | SET_DOCTOR-PREFERENCES | setting,doctor preferences | radiography and fluoroscopy | Precision RXi | How to set doctor preferences? | 1 Select the Patient List button. 2 Select a new or existing patient file. 3 Click on Down arrow for Performing Physician. The resulting menu will show all physicians currently established as well as a <New Physician> option. Select the <New Physician> option. The Physician Name Editor will be displayed. 4 Highlight the Doctor’s name from the list of those available.5 Select the Edit physician preferences button. 6 The Doctor Preferences screen will be displayed. Doctor Preferences Edit Screen 7 Select each of the buttons on the left to select the preferred settings for each mode: Acquisition, Review, Print, Archive, & Reference (each setting is described in detail in Section 06 Acquiring Images). 8 When settings are complete, select the OK button. 9 The Doctor Preference screen will close and return to the Physician Name Editor screen. 10 Repeat steps 1-9 to set the preferences for a different Doctor and when the process is complete, select the OK button.
GPRXi66 | FLURO_ROAD_MAPPING | activate/deactivate, fluoro road mapping | radiography and fluoroscopy | Precision RXi | How to activate/deactivate fluro road mapping? | 1 Left click on the Mask button. 2 The text within the dialog box will “light up” (change from gray to white). There are three ways of selecting the image to be used as the mask image. • LIH (Using the Last Image Hold on the screen) • Max Opacification • Minimum Opacification
GPRXi67 | MASK_ACQUISTION | last,image hold,mask acquisition | radiography and fluoroscopy | Precision RXi | How to hold the last image? | 1 While producing fluoro radiation, inject the target vessel with a contrast agent 2 When the target area is full of contrast, release fluoro to obtain an LIH image 3 The resulting LIH image will be the mask 4 Activate fluoro again to acquire the roadmap image
GPRXi68 | DSA_PROGRAMMABLE_SEQUENCES | programmable,dsa sequence | radiography and fluoroscopy | Precision RXi | How to program DSA sequences? | 1 Right-click while cursor is over the image display area. The Right Click Menu will be displayed. 2 Position cursor on button and a second Sub-menu will be displayed. 3 Left-Click on 4 The Programmable Sequences screen will be displayed. 5 Left-Click on the New button.6 Enter the “Sequence Title” in text box. 7 Click on the Phase to be edited. 8 Double-click on “0 fps” in Rate column - in line with the Phase selected. 9 The FPS (frames per second) drop-down menu will be displayed (fps = frames/second). Click on the desired acquisition rate. 10 Click on the 0 in the Time column – in line with the Phase selected. 11 Enter the number of seconds to acquire at the specified rate12 Repeat steps 4-11 to configure as many Phases as needed. 13 Click on the Save button to save the Sequence settings. Optionally, click on the Cancel button at any time to clear the current settings. 14 After the Save button is clicked, the Programmable Sequences window displays with the new Sequence added to the list..
GPRXi69 | EXISTING_ PROGRAMMABLE _SEQUENCE | existing,programmable sequence, editing | radiography and fluoroscopy | Precision RXi | How to edit exiting programmable sequence? | 1 Right-click with cursor over the image display area. The Right-Click Menu will be displayed. 2 Position cursor on button and a second sub-menu will be displayed. 3 Left-Click on . 4 The Programmable Sequences screen will be displayed. 5 Left click on the sequence to be edited. 6 Left click on the Edit button. The Edit Programmable Sequence window is enabled7 Edit the Sequence Title if necessary. 8 Double left click on the Rate drop-down list to change the acquisition rate for any phase. 9 Enter the number of seconds to acquire in the Time box10 Click on the Save button to save the sequence settings. 11 The Edit Programmable Sequence menu is closed, and the Programmable Sequences window is updated with the new Sequence added to the list..
GPRXi70 | STORE_ FLURO_ LIH _TO_ LOCAL _DATABASE | storage,fluro LIH | radiography and fluoroscopy | Precision RXi | How to store fluro LIH to local database? | This feature provides the capability to selectively store LIH images to the system’s database. A thumbnail will be created for each image stored. 1 Click the Store Image button in the Standard Control Box. 2 The LIH image is transferred to the local database. 3 A thumbnail for this image is added to the end of the list for the current study. Fluoro images can be automatically stored by enabling the Auto Save LIH to disk option in Fluoro Doctor Preferences dialog.
GPRXi71 | AUTO_ACQUISITION_REPLAY | change default bahavior,auto acquisition replay | radiography and fluoroscopy | Precision RXi | How to change default behavior? | 1 Right-click in the image. Menu will appear. 2 Position the cursor on Configure, and then click on Current Doctor Prefs 3 Make selections from Doctor Preference set-up screen. 4 This function allows the temporary modification of any of the Doctor Preference fields for the current patient/doctor.
GPRXi72 | SHORT_SAVE_CURSOR | short save,cursor,folder | radiography and fluoroscopy | Precision RXi | How to use a short save cursor? | When a sequence is opened for Review in 4on1 or 16on1, the ShotSave™ cursor will appear in the Cursor Tool bar. 1. Single click on the cursor. Cursor becomes “Active” 2. Move ShotSave cursor over image representing the sequence that to be saved to the ShotSave™ folder and double click. 3. The image will be saved to the ShotSave™ folder and the number on the ShotSave Folder button will increase to reflect the addition of the new image. The image will be marked with an “S” in the lower left corner of the image on the main image screen while in the MID mode. To view the ShotSave folder, click down arrow under <Study> and left click on ShotSave Folder.
GPRXi73 | MARK_FOR_DELETE | mark,for delete | radiography and fluoroscopy | Precision RXi | How to mark for delete? | This feature allows an entire sequence to be marked for deletion. In order to delete a sequence, the pointer cursor must be active and the sequence thumbnails displayed. 1. Place the cursor over a thumbnail and right-click. 2. The Right Click menu will be displayed 3. Select “Delete Sequence” 4. A Delete “icon” will appear as a “D” in the upper right corner of the thumbnail as well as in the lower right corner of the image in the main image display area. 5. When the patient file is closed, a warning window will be displayed asking for confirmation of the delete sequence operation 6. Read the descriptions for each button and then make selection.
GPRXi74 | PAN | panning,function | radiography and fluoroscopy | Precision RXi | How to operate pan function? | 1. Position the Cursor on the image. 2. Click and hold the left mouse button. 3. Drag the Cursor in the direction you want to move.
GPRXi75 | ANNOTATION_SUBMENU | text annotation,submenu | radiography and fluoroscopy | Precision RXi | How to annotate using submenu? | 1. Select the Annotation Cursor from the Cursor Tool Bar to enter Annotation mode. 2. Move Cursor into the Main Image area. 3. Cursor will become an “A” indicating Annotations has been activated. 4. Double click within Main Image area to display the Annotations Sub-menu. 5. Click on Predefined from Annotations Sub-menu. 6. Click on Configure to display the Annotation Name screen. Annotation Name Screen 7. Type in the Preferred Text. This Text will also be used to form the list of Predefined Annotations.8. Click . Annotation Name Screen 9. When adding Annotations is complete, select the Close button to save the list. Selecting the “X” in the upper right corner of the window will Cancel/Erase any Annotation changes made in the previous steps.
GPRXi76 | ANNOTATION_RIGHTCLICK MENU | text annoation,using right click menu | radiography and fluoroscopy | Precision RXi | How to do predefined annotation using right click menu? | Predefined Annotations can also be created from the Right Click menu accessible from the Main Image screen. 1. Right click within Main Image screen to display sub-menu. 2. Place cursor over Configure. The Right Click sub-menu will appear. 3. Click on Annotation Presets4 The Annotation Name screen will be displayed 5 Type in the desired text. This text will also be used to form the list of Predefined Annotations. 6 Click. . 7 When adding Annotations is complete, select the Close button to save the list. Selecting the “X” in the upper corner will Cancel and Erase any Annotation changes made in the previous steps.
GPRXi77 | DELETE _AN_ANNOTATION | deletion,predefined annotation | radiography and fluoroscopy | Precision RXi | How to delete a predefined annotation? | 1 From the list of Predefined Annotations on the Annotation Name screen, highlight the Annotation to be deleted. 2 Click the Delete button to instantly delete the Annotation from the list.
GPRXi78 | DELETE _TEXT_POINTER_ANNOTATIONS | delete, text or pointer annotation | radiography and fluoroscopy | Precision RXi | How to delete text or pointer annotations? | 1. Position Cursor over the Delete menu item in either the Text or the Pointer Sub-menu. 2. A window will appear asking for Confirmation of Delete Request; make appropriate selection. Both the Pointer and the Text will be deleted from the display.
GPRXi79 | EDIT _SEQUENCE _TITLE _PATIENT _LIST | edit,sequence title,patient list | radiography and fluoroscopy | Precision RXi | How to edit sequence title from patient list? | 1. Right click with cursor over the image in Main Image screen. The Right Click menu will be displayed. 2. Click on Select Patient: 3. The Patient List screen will be displayed: Patient List Screen 5. Select the Edit button . The Patient Information Screen will be displayed: Patient Information Screen 6. Left click on Sequence in the Studies/Series window. 7. Screen changes to allow sequence title to be changed (see figure that follows). Sequence Title Screen 8. Enter New Title in Description Box and than click Save to save the changes.
GPRXi80 | CREATE_PREDEFINED_SEQUENCE_TITLE | create,sequence title,predefine | radiography and fluoroscopy | Precision RXi | How to create predefined sequence title? | 1. From the Sequence Information window, click on <Configure Sequence Titles>. 2. The Predefined Sequence Titles screen will be displayed (see next figure):3. Enter the name for the sequence, then click the button. 4. Repeat step #3 to add other Sequence Titles. 5. Click to store the new sequence titles and exit screen.
GPRXi81 | EDIT_SEQUENCE_TITLE | edit,by thumbnail,right click menu | radiography and fluoroscopy | Precision RXi | How to edit sequence title using thumbnail right click menu? | A sequence can quickly be renamed directly from the thumbnail of the image. 1. Right Click while over the Thumbnail image to display sub-menu 3. Enter revised sequence title, then click OK. 4. The sequence name will be changed immediately.
GPRXi82 | SEQUENCE _TITLE_LIBRARY | selection,sequence title library | radiography and fluoroscopy | Precision RXi | How to select from sequence title library? | To select a title from the Sequence Title Library: 1. Click down arrow to view available titles next to the Description box.
GPRXi83 | MASK_SUBTRACTION | deactivate subtraction | radiography and fluoroscopy | Precision RXi | How to deactivate subtraction? | 1. Single click the “Mask” button. Mask Button 2. The current image will be displayed unsubtracted, using previous image’s Processing settings. 3. The mask button will return to its previous state.
GPRXi84 | DISPLAY_LUT | look up tables | radiography and fluoroscopy | Precision RXi | How to display LUT? | 1. Right click over the Main Image screen when there is an image displayed. The right-click menu will display. 2. Position the cursor over “Configure, “ bringing up the sub-menu. NOTE: Display LUTs will not be available for images acquired with AIO ON, test patterns, or any unsaved Fluoro image. 3. Click on “Display LUTs”
GPRXi85 | PRINT_STUDY | print,images in a study | radiography and fluoroscopy | Precision RXi | How to print images in a study? | 1 To print all the images in a study, click the arrow on the Print button. 2 The Print Layout sub-menu will be displayed: 3 Click on Print Study 4 The Configure Print Settings window will appear 5 Make appropriate selections and then click OK The images for the current study will be sent to the printer.
GPRXi86 | PRINT_SHOTWAVE_IMAGES | print,shotwave image files | radiography and fluoroscopy | Precision RXi | How to print shotwave image files? | 1 To print the Shotsave file images, click the arrow on the Print button. 2 The Print Layout sub-menu will be displayed: 3 Click on Print Shotsave 4 The Configure Print Settings window will appear. 5. Make appropriate selections and then click button. 6. All images in the Shotsave file will be printed.
GPRXi87 | CONFIGURE_AUTO_PRINT_FEATURES | configure,autoprint,features | radiography and fluoroscopy | Precision RXi | How to configure auto print features? | When the AutoPrint feature is ON, the system sends each image (for the current patient file only) to the printer as it is acquired with PACE parameters applied (if enabled). Use these instructions to set the parameters that will be used for the images printed using the AutoPrint feature. (Refer to section 8.7 for information on turning the AutoPrint feature on and off). 1 To access this feature, click the Arrow button. 2 The Print Layout sub-menu will be displayed: At this window, the processing defaults used for printed images can be changed. 3. Click on Configure AutoPrint 4. The Configure Print Settings window will appear.
GPRXi88 | AUTOPRINT | features | radiography and fluoroscopy | Precision RXi | How to autoprint? | When the AutoPrint feature is ON, the system sends each image (for the current patient file only) to the printer with PACE parameters applied as the image is acquired. 1 To turn ON the AutoPrint feature click the arrow next to the Print button. 2 The Print Layout sub-menu will be displayed: NOTE: This feature will remain active for the current patient file until it is turned off. The feature can be turned OFF by reopening window and clicking on . 3. Click on AutoPrint. The sub-menu will disappear and AutoPrint will be enabled. 4. When AutoPrint is enabled, a checkmark (√) will appear in the box. To check whether or not AutoPrint is enabled, reopen the Print Layout sub-menu and look for the checkmark.
GPRXi89 | IMPORT_FROM _CD_DVD | importing,cd or dvd | radiography and fluoroscopy | Precision RXi | How to import from CD or DVD? | 1 At the Main Image screen, click on the Patient List button to display the Patient List screen. 2 At the Patient List screen, click on the tab. A new menu will open (see Figure below). 3 Insert DICOM Compliant Media containing Patient data to be Imported. 4 Select the CD/DVD Media Device name5 A list of available files/studies/series will be displayed in the corresponding pane of the Media Import screen. 6 Highlight the Patient, Study, or Series to be imported. 7 Select one of three available import buttons (see chart for descriptions).8 The System will import the patient information and images from the CD. 9 Select the Patient List tab. The Patient List will now show the recently imported patients.
GPRXi90 | IMPORT_USING_DICOM_RECEIVE | import,dicom receive | radiography and fluoroscopy | Precision RXi | How to import using DICOM receive? | 1 At the Main Image screen, click on the Patient List button to open the Patient List. 2 At the top of the Patient List screen, click on the button. Patient List Screen 3 The system is now in DICOM Receive mode. The patient files/images can now be sent from the other DICOM device. 4 The patient file information and images will be received and added to the Patient List. 5 Check on the status of the process to see when the import function is complete (see instructions in 9.1.4 - Status Screen). 6 Once a patient file has been added to the Patient List, select the Patient List tab and then select the desired patient file for review.
GPRXi91 | IMPORT_USING_DICOM_QUERY/RETRIEVE | import,dicom retrieve | radiography and fluoroscopy | Precision RXi | How to import using DICOM retrieve? | This feature allows the system to search DICOM Databases for patient information and images that can be imported into the system via the DICOM Retrieve function. The system acts as a DICOM Query/Retrieve Service Class User to perform this function. 1 Click on the Patient List button to display the Patient List screen. 2 Click on the Query/Retrieve tab. Query/Retrieve Screen 3 Select the desired Provider from the Provider List Drop-down menu. 4 Click on the Verify button to Verify the Status of the network device. The Status of the device will be shown in the Status box: should be Available or Ready. 5 Enter desired Search criteria – Patient information and/or Study dates. 6 Click on the Query button. The Precision Imaging System will retrieve the Patients from the selected remote device. The Status of this process is shown in the Status box.7 When process is complete, the Status box will say “Query Complete,” and the Patient Information will be updated in the corresponding panes of this screen. 8 Highlight the desired Patient for Retrieval. 9 Click . When the transfer is complete, the newly retrieved file will be listed in the Patient List. Select the Patient List tab to view. Query/Retrieve Results Screen To display Study Information for a selected Patient, double click on the Patient’s Name. Related Studies will be displayed in the Study window. Similarly, to display specific Series, double click on the Study Title. The series related to that Study would display
GPRXi92 | STATUS_SCREEN | status | radiography and fluoroscopy | Precision RXi | How to enable status screen? | A Status screen is available to verify that the operation has been completed successfully or that the Import is still in progress. 1 At the top of the Main Image screen, click on the button. The Status screen will display. The Status tab will be visible. 2 At the Status Dialog screen, click on the Messages tab. The Messages screen will display. Messages Screen 3 This screen will show the Status of the Import and is updated with each change. 4 When finished, select Exit to Close this window. If Import process fails, try again or contact the Site Network Administrator for assistance.
GPRXi93 | EXPORT_TO_DICOM_CD_/DVD | exporting,to dicom,or,cd,or dvd | radiography and fluoroscopy | Precision RXi | How to export to dicom,cd or dvd? | 1 At the Main Image screen, click on the button to open the Patient List. 2 At the Patient List screen, click on the tab to open the Media Export screen 3 Insert Media/Disc into drive. 4 Select Device: [Media-DVD] from the Device drop-down menu. 5 Select a single or multiple Patient(s) for Export. 6 Optionally, select single or multiple Studies to be Exported.7 To have Processing Features (previously applied to the images, such as Edge, Brightness, Contrast adjustments, etc.) burned into the exported images, select the Process Image button. 8 To have Text Annotations or Pointers burned into the exported images, select the Embed Annotations button. (Note: This option is only available if Process Image is active. This EMBEDS the Annotations and Pointers into the images pixels as a permanent part of the image, and cannot be removed after Export;) 9 Select either Export Patient button to export the entire Patient, or select the Export Study button to Export the selected Study. 10 A Transfer Job is created and added to the Queue. See Paragraph 9.3.6 Job Queue Export Processing in this Section for more information.
GPRXi94 | MANUAL_DICOM_SEND | manual,sending,using dicom | radiography and fluoroscopy | Precision RXi | How to manually send using dicom? | 1 At the Main Image screen, click on the DICOM button. 2 A short DICOM sub-menu will be displayed. 3 Select either Send Study or Send Shotsave. 4 The Archive Destination screen will be displayed.5 Click on the desired Destination from list of those available. 6 To include any processing features previously applied to the images, (such as Edge, Brightness, Contrast adjustments, etc.) with the exported images, select the Process Image button. 7 To include any text annotations or pointers with the exported images, select the Embed Annotations button. 8 Click on the OK button to send the images.
GPRXi95 |  |  | radiography and fluoroscopy | Precision RXi
GPRXi96 |  |  | radiography and fluoroscopy | Precision RXi
GPRXi97 |  |  | radiography and fluoroscopy | Precision RXi
GPRXi98 |  |  | radiography and fluoroscopy | Precision RXi
GPRXi99 |  |  | radiography and fluoroscopy | Precision RXi
GPRXi100 |  |  | radiography and fluoroscopy | Precision RXi
GPRXi101 |  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi
|  |  | radiography and fluoroscopy | Precision RXi

## Sheet4
TID (unique) | KEYWORD(unique) | CATEGORY | QUESTION | ANSWER | DOCUMENTATION LINK | YOUTUBE SOLUTION LINK
GACTS1 | ANRT_SCANNING | ANRT,scan applications | How to enable anrt scanning? | Step 1: Set up a scanning plan. Step 2: In the ViewEdit area, press Recon. Step 3: In the Recon Setting area, press Recon Option. Step 4: In the XNR area, press On. Step 5:Continue to proceed with scanning according to the original plan. "XNR" will be displayed on the top left corner of an image | https://drive.google.com/file/d/1Bu2OPkwlq201kwt1w-g1hpYcl7FlrjzC/view?usp=sharing
GACTS2 | MA_REDUCTION_GUIDANCE_SCAN | ANRT,scan applications | How to aquire a scan using ma reduced value? | Step 1: Set up a scanning plan. Step 2: Press mA in the ViewEdit area. Step 3: Press mA Reduction in the mA control area. Step 4: Enter the desired mA value. Step 5: Press OK in mA control. Step 6 Continue scanning according to the original plan
GACTS3 | RECONSTRUCTION_OF_ASIR_IMAGES | ANRT,scan applications | When to reconstruct asir images? | *Reconstruction of ASiR images can be set in advance during Recon1, and can be set in advance during Recon2, Recon3... Recon10, or set during post-reconstruction. * If the post-reconstruction method is adopted, a new series should be selected when applying ANRT reconstruction data
GACTS4 | CONSTRAINTS_DURING_ASIR_AQUISITION | ANRT,scan applications | What are the constraints that apply to the acquisition of data where ASiR will be applied? | * ANRT can only be used for digital tilted images, tilt angle cannot be 0 degree. *When ANRT is being applied, 5 reconstruction calculation methods can be used, [STND], [STD+], [BONE], [EDGE] and [DETL]. *ANRT is not compatible with any real-time interactive image modes, for example the SmartPrep (baseline and monitoring images) and Biopsy modes. * ANRT is not compatible with the AutomA mode. * ANRT cannot be used when the actual scanning current is less than 50mA
GACTS5 | EFFECTS_OF_REDUCED_MA_VALUES | ANRT,scan applications | What are the effects of reduced ma values? | *Applying this program, pressing the Reduce mA value button can make the system's ANRT function automatically effective. | https://drive.google.com/file/d/1CUK1zzCMmtH7Gio20Pa559FxtjELpLJd/view?usp=sharing
GACTS6 | MA_REDUCTION_VALUE_WHILE_SCANNING_A_HEAD | ANRT,scan applications | What is the reduced ma percentage while scanning the head? | *When scanning a head,the default mA value reduction percentage is 20%
GACTS7 | MA_REDUCTION_VALUE_WHILE_SCANNING_LUMBAR_CERVICALF_VERTEBRAE | ANRT,scan applications | What is the reduced ma percentage while scanning the lumbar an vertebrae? | *When scanning the lumbar and cervical vertebrae, the default mA value reduction percentage is 30%
GACTS8 | MA_VALUE_GUIDANCE | ANRT,scan applications | What are the factors to be considered when utilizing reduced mA values for guidance? | * When the desired mA value is less than 50mA, the mA value reduction function will be ineffective. * When scanning the lumbar andcervical vertebrae, and the necessary mA value is between 50mAand 75mA, when doing a headscan and the necessary mA value is between 50mA and 65mA, [mA reduction] is effective, but when you press [Ok], [mA reduction] will not be usable and will return an operational message. * Further, the actual mA value displayed over the mA table is the necessary mA value. *The operational message will be displayed on the Message Board: The mA reduction function cannot be used, because the requested mA value reduction is less than 50mA. * When [mAReduction] is pressed, ifthe actual mA value and thenecessary mAvalue are inconsistent, DR and the actual mA value will be displayed in the mA row. *When [mAReduction] is pressed,the [ANRT] button locatedin the ReconOption areawill be enabled automatically.
GACTS9 | ANRT | ANRT,scan applications | What is anrt used for? | *ANRT is a reconstruction technology used for decreasing the noise of diagnostic images while preserving the structural detail information of these images. ANRT noise-reduction technology can, while maintaining image noise level, lessen the dose of some X-ray testing, or improve image noise level and image quality.
GACTS10 | ASIR | ASIR,scan applications | What is asir used for? | *Adaptive Statistic alIterative Reconstruction (ASiR) is a reconstruction technique designed to reduce noise in diagnostic images while preserving the structure details in the image. *ASiR noise reduction technique may allow forthe reduction of the x-ray dose in some studies while maintaining image noise levels or be applied to improve images with an unacceptable level of noise.
GACTS11 | SNR_CALCULATION | ASIR,scan applications | How to calculate snr? | *The measurement involves calculating a signal delta divided by the standard deviation (noise).
GACTS12 | ASIR_GUIDANCE | ASIR,scan applications | What are the factors to be considered while selecting asir recon mode? | * ASiR is not compatible with Neuro 3D Filters (N1, N2, N3). * ASiR is not compatible with any of the real time interactive image modes such as Smart Prep or Biopsymode.
GACTS13 | IMAGE_RECONSTRUCTION | ASIR,ANRT, scan applications | What is image reconstruction? | *Image with 100% noise reduction. *There are 10 blending levels available which are simply the amount of noise reduction based on the maximum of the 100% image reconstructed with the original image data. *It is expected that confidently selecting the optimal amount of noise reduction and integrating the reduction in dose possible wil lrequire some adapting on how this reconstruction technique works
GACTS14 | BEFORE_USING_ASIR | ASIR,scan applications | What are the things to be done before using asir? | *Before using ASiR, the site physicist, in collaboration with the Radiologist, should conduct image quality evaluations with varying degrees of ASiR and different scan techniques, both at routine dose and with decreased dose. *This should be done using your site's preferred method and phantoms. *Using this information, an appropriate starting point for ASiR level and diagnostic scanning techniques can be incorporated into your site's protocols. *WhenASiR is usedfordose reductionorimage qualityimprovement,parameter settings shouldbe made in steps and then reviewed for diagnostic image quality until the desired clinically appropriate image quality is achieved.
GACTS15 | NOISE_INDUCTION_ADJUSTMENT_FACTORS | ASIR,scan applications | How to calculate noise induction adjustment factor? | *Based on calculating the relative difference in measured noise standard deviation (SD) with Filtered Back Projection (FBP) and ASiR on a 20cm water phantom
GACTS16 | MA_ADJUSTMENT_FACTOR | ASIR,scan applications | How to caluculate ma adjustment factor? | *These factors were calculated based on a simple inversion of the squared NI Adjustment Factors
GACTS17 | ASIR_SLICE_MODE | ASIR,scan applications | What is asir slice mode? | * can be applied to acquisitions of any slice thickness and any acquisition mode where noise reduction is in X and Y (2D). * can be applied to a range of slices.
GACTS18 | SCAN_WITH_ASIR | ASIR,scan applications | How to aquire a scan with asir? | Step 1: Prescribe a scan. Step 2: Click Recon Options and from the Recon Options screen, select an ASiR mode (Slice) and an ASiR % value. As the percentage value increases, the noise reduction increases. * If you want to use ASiR to maintain image quality but reduce dose, from the mA Control screen,reduce mA or increase noise index value. Step 3: Proceed to scan.
GACTS19 | ASIR_PROSPECTIVE_RECONSTRUCTION | ASIR,scan applications | How to perform asir prospective reconstruction? | Step 1: Prescribe a scan. Step 2: From the ViewEdit screen, click Recon Tab. Step 3: From the Recon tab, click any prescribed PMR R2 to R10. Step 4: From the Recon Option screen, select an ASiR mode: Slice and an ASiR % value. Step 5: Click OK. Step 6:Continue with your prescription and proceed to scan.
GACTS20 | ASIR_RETROSPECTIVE_RECONSTRUCTION | ASIR,scan applications | How to perform asir retrospective reconstruction? | Step On the Exam Rx desktop, click Retro Recon. 2. From the Retro Recon screen, select exam and series and click Select Series. 3. From the Retro Recon Setup screen, click New Series Number. 4. ClickReconOptions. 5. From the Recon Option screen, select an ASiR mode: Slice and an ASiR % value. Click OK. 6. Click Confirm.
GACTS21 | EFFECTS_OF_MANUAL_MA | ASIR,scan applications | What are the effects of using manual ma? | *When utilized with manual mA a Reference mA is entered, the Reference mA is the baseline, nondose reduced mA. * When a % Dose Reduction value is entered the system calculates the actual mA that will be usedfortheacquisitionandselects theappropriateASiRleveltomaintainimagequality atthedose-reduced mA. *For example if a Reference mA value of 200, a % Dose Reduction of 50% is entered, the actual mA for the scan acquisition will be 100
GACTS22 | EFFECTS_OF_AUTO_MA | ASIR,scan applications | What are th effects of auto ma? | *When utilized with Auto mA, a % Dose Reduction value is entered. *The system will calculate the mA for the scan taking into account the % Dose reduction value. *The dose-reduced mA will be shown in the mA table and the system selects the appropriate ASiR level to maintain image quality at the dose-reduced mA.
GACTS23 | SCAN_WITH_DOSE_REDUCTION_GUIDANCE | ASIR,scan applications | How to perform scan usin dose reduction guidance? | Step 1: Prescribe a scan. Step 2: From the ViewEdit screen, click mA. Step 3: From the mA, control screen click % Dose Reduction in either manual mA or Auto mA. Step 4: Enterthe desired dose reduction percentage. Step 5: From the mA control screen in Manual mA click Reference mA. Step 6: Enterthe desired mA value. Step 7: From the mA control screen, click OK. Step 8: Continue with your prescription and proceed to scan.
GACTS24 | BIOPSY_MODE | ASIR,scan applications | What is biopsy mode? | *Biopsy Mode improves the efficiency of setting up and acquiring slices during a biopsy. *All biopsy scan parameters are available on a single screen from which you can launch the biopsy scan. * You are able to choose which direction the scanner acquires images from a centering point as well as how many images to acquire. *You can change slice thickness and/or interval. You can enter a specific location for a slice if required. *AutomA is turned off when Biopsy Mode is entered, Review the Manual mA value prescribed. *All scans taken in Biopsy Mode are the Axial Scan Type. | https://drive.google.com/file/d/1gLF2epSNvEvTt0X3_KP0NV9fFU_NAoil/view?usp=sharing
GACTS25 | BIOPSY_REFERENCE | ASIR,scan applications | What is biopsy reference? | The Biopsy Reference tells the system how to accquire images based on the Biopsy Location. * Superior means all images are acquired from the Biopsy Location, superiorly. * Centered means the system acquires images above, at and below the Biopsy Location, If you enter an even numbers of slices, the system places the extra image in the syperior portion. * Inferior means the images are acquired from the Biopsy Location inferiorly
GACTS26 | SCAN_IN_BIOPSY_MODE | ASIR,scan applications | How to aquire a scan in biopsy mode? | Step 1: From the ViewEdit screen orthe Dynaplan screen, click Biopsy Rx. * The images remain in the same series. Step 2: From the Biopsy Rx screen, complete the following selections: a. Select a Biopsy Reference. * Click Superiorto make the biopsy reference location the first scan, with subsequent images acquired superiorly from thatlocation. * Click Centered to make the biopsy reference location the center slice, with the other images acquired above and below that location. Typically, select an odd numberfor Centered to split the images evenly above and below the location. * Click Inferior to make the biopsy reference location the first scan, with subsequent images acquired inferiorly from thatlocation. b. Select a biopsy location. * Type a specific table location in the Biopsy Location field. * In orderto obtain positioning light position, the table must be moved to align the positioning lights, then click Internal or External. c. Type the number of images for each pass in the Number of Images field. d. Select a slice thickness. Axial mode only allows 2.5, 5.0, or 10.00 mm. Step 3:Click Confirm Biopsy Rx to activate the biopsy scan. Step 4: Optional: manually film the biopsy images. Step 5: When the biopsy scan is finished and if additional groups are being scanned in the same series, re- activate AutomA. is turned off with Biopsy Mode
GACTS27 | DMPR_CONSIDERATIONS | DMPR,scan applications | What are the considerations while in dmpr? | *Each group must be contiguous to be able to be combined with the prior group in a DMPR session. *A DMPR Auto Apps session can be programmed for only Recon 1. *DMPR sessions are limited to 2,000 images. *Additional reformat protocols can be added using the Reformat application or Volume Viewer in the General category from ImageWorks. *Any scans acquired after scanning has completed on the original scan group using Add Group or One More scan are not added to the DMPR session. *Remember to include all of the desired coverage area in the original scan prescription. *If the patient orientation is prescribed as Decubitus (right or left), you will observe that the Paging slider for the sagittal image will scroll images in DMPR coronal viewport and vice versa. This is because in Decubitus orientation, patient’s sagittals and coronals are switched. * Hence, the DMPR sagittal viewport contains the coronals from patient’s reference axis, and the DMPR coronal viewport contains the sagittals from the patient’s reference axis
GACTS28 | DMPR_CONSIDERATIONS_WITH_AUTO_BATCH | DMPR,scan applications | What are the consideration of dmpr with auto batch? | * DMPR displays image anatomical orientation within a viewport as follows: A is at the top, P is at the bottom, R is on the left and L is on the right. If you have a data set where the patient was scanned prone, the image display is flipped to this orientation. * DMPR sessions are limited to 2,000 images. *Reformat protocols used in DMPR must be built in Reformat or Volume Viewer in the General category as Single Step protocols. Multiple Step protocols are excluded from the list of available protocols for selection in DMPR. *For AutoBatch, the reformat protocol cannot use the Oblique viewport. * A DMPR Auto Apps session can be programmed for each of the 10 PMR1s. A DMPR session can be a single group or combined groups.
GACTS29 | DMPR_SCAN_CONSIDERATIONS_WITH_AUTO_BATCH | DMPR,scan applications | What are the scan considerations of dmpr with auto batch? | * If DMPR is enabled under the Recon tab, Add Group displays some insensitive fields because it is combined with the current DMPR session. Turn DMPR off for the added group if you no longer want it to be included as part of the DMPR session. This allows changes to any of the acquisition parameters. * The W/L of the axial source images affect the edge detection of the auto batch. If the W/L is too narrow, then anatomy is cut off from the auto batch. If the W/L is too wide, then undesired anatomy is included in the auto batch. Make sure to set the appropriate W/L in Recon Options. * If you prescribed auto batch protocols and then quit DMPR before the auto edge detection starts, the reformatted images are not completed. * To page/scroll through images during DMPR reconstruction, use the Oblique view for paging. Once the reconstruction is complete, the image may change DFOV to accommodate all the data in 3D. * If a scan stops before all DMPR images for a session are acquired, it may not be possible to activate the interactive mode because all the images are not yet available. Quit the DMPR session to return to Exam Rx display and generate any reformatted images you want in Reformat from ImageWorks desktop. * DFOV for AutoBatch axial images defaults to the DFOV of the scan range instead of the DFOV in the protocol. * When multiple sessions of MPR are stacked, you cannot get to the next one until you finish the previous session. * Click Start Direct MPR Review to get to the Interactive display screen and select Quit. *If you do not want to quit the current DMPR session, you can still see the reconstructed axial images. * Click Auto Link and view the images in the viewport. * You are not able switch to another AutoView layout until reconstruction for the DMPR session is complete. To review images for another exam while images are being reconstructed, go to the Image Works desktop and use the Viewer. * If you want to leave the DMPR session available once the reformat model is built, click AutoReview Layouts to change the focus from DMPR. This allows review of images on Exam Rx display without ending the DMPR session. * Do not quit until auto batch is started, as indicated by the auto batch in the reference image or until auto edge detection is done, which is indicated by the black line directly under the Start Direct MPR Review button.
GACTS30 | DMPR_SCAN_WITH_AUTO_BATCH | DMPR,scan applications | How to aquire a dmpr scan with auto batch? | Step 1: Prescribe a scan. Step 2: Click the Recon tab. Step 3: From the Recon tab, click Off for group under the Auto Apps column. Step 4: From the Session Setup screen, click Direct MPR and then click OK. Step 5: From the Session Setup screen in Session Selection, click Start New to setup a DMPR session. *parameters must be the same: slice thickness, interval, SFOV, DFOV, scan type, rotation speed, image center, and algorithm. The groups must be contiguous. If you want to change parameters between groups, click Start New instead of Combine Current. *If you want to change parameters between groups, click Start New instead of Combine Current. Step 6: To select reformat protocols to be used in the DMPR session, click Unused under Batch Protocol to open the list of available protocols. * Reformat Protocols built in Reformat or the General category in Volume Viewer are displayed in the Batch Protocol List. *Select up to five protocols from the Batch Protocol List, and click OK. *Use the Next and Prior arrows to move through the list of GE reference reformat protocols and user define reformat protocols. *The order you selectthe protocols is theordertheyare listed forBatch Protocolandthe order Auto Batch applies them. Step 7:Prescribe Filming, Auto Batch, Auto Store, and Auto Transfer as needed for each protocol. a. Click Auto Batch Off to turn on Batch Reformats to be created automatically once all the images for the DMPR series have been reconstructed. *If Auto Batch is enabled, an automatic edge detection algorithmdetects the outer edges ofthe object and sets the range for Batch Reformats. *Auto Batch is not available if the reformat protocol was created in the Oblique viewport. b. Click Auto Store Off to automatically have Batch Reformats archived to designated archive device. Step 8:Click Auto Transfer Off to select up to 4 host to automatically transferthe DMPR Batch Reformats by series. Step 9:Click OK to close DMPR Session Setup screen. Step 10:Click Confirm. Step 11:press Move to Scan 12. Press Start Scan. *The Start Direct MPR Review button displays the Auto Batch progress in the Exam Rx control panel. *After four images have been reconstructed, the currently selected AutoView layout is replaced with the DMPR AutoView layout. In this layout, the following image types are displayed in specific viewports: *upper left = oblique (or AutoFilm accessed by clicking on the page turn icon) *upper right = axial *lower left = sagittal *lower right = coronal Step 13 Click Start Direct MPR Review to open the Direct MPR Interactive Review screen. *thi step is onlfy necessary if you want to review the data Step14:From the upper-left viewport, use the crosshair cursors, click and drag, or use the review controllerto interact with the displayed images. Step 15: Once all tasks are finished, including all Batch Reformats, click Quit and then, click Yes to end the DMPR session. Step 16:For details on manipulating the DMPR image, see Manipulate DMPR images.
GACTS31 | DMPR_SCAN_WITHOUT_AUTO_BATCH | DMPR,scan applications | How to aquire a dmpr scan without auto batch? | Step 1: Prescribe a scan. Step 2: Click the Recon tab. Step 3: From the Recon tab, click Off for group under the Auto Apps column. Step 4: From the Session Setup screen, click Direct MPR and then click OK. Step 5: From the Session Setup screen, click Start New to setup a DMPR session. *You can combine groups or set each group as a new session. To combine groups the following parameters must be the same: slice thickness, interval, SFOV, DFOV, Scan Type, Rotation Speed, image center, algorithm, and ASiR Level. The groups must be contiguous. * If you want to change parameters between groups, click Start New instead of Combine Current. Step 6: To select reformat protocols to be used in the DMPR session, click Unused under Batch Protocol to open the list of available protocols. * Reformat protocols built in Reformat or the General category in Volume Viewer are displayed in the Batch Protocol List. * Multi-step reformat protocols do not appear in the list. Step 7: Select up to five protocols from the Batch Protocol List, and click OK. * Use the Next and Prior arrows to move through the list of GE reference reformat protocols and user define reformat protocols. *The order you select the protocols is the order they are listed for Batch Protocol. Step 8: Prescribe Filming, Auto Store, and Auto Transfer as needed for each protocol. Step 9: When generation of the DMPR session is complete, the Direct MPR BatchRx screen automatically displays. Step 10: From the Direct MPR BatchRx screen, select a protocol from the list of protocols. *If you use a protocol, it propagates the batch prescription to the viewportwith the Slice Thickness, Spacing Between Views, FOV, and Number of Slices programmed in the protocol. * Confirm that the Mode is Oblique. Step 11: Set the parameters for Thickness, Spacing Between Views, FOV, and Number of Views. Step 12: Prescribe the Reformat range. For details on manipulating the DMPR image, see Manipulate DMPR images. Step 13: Click Preview and then Apply to review the reformat prescription. *Click Pause to pause the review. *Click Resume to continue the review. *Click Stop to end the review. Step 14:Click Save and then Apply to save the prescribed DMPR reformat series. * The following actions can be applied together with Save: * Click Film to film the DMPR Reformat series based on camera selection and setup in Session Setup *Click Store to automatically archive the reformat series to designated archive device. * Click Show to view the reformat images as they are being created in the review viewport in the lower-right corner of the viewport. * Click Transfer to network reformat images to Host selected in Session Setup. Step 15: Repeat steps 6 to 14 to create additional batch reformats. Step 16: Click Back to exit Direct MPR BatchRx. Step 17: Click Quit and then, click Yes to close the DMPR session or click Back to continue to Direct MPR Review at a later time.
GACTS32 | EXAM_SPLIT | exam split,scan applications | What is exam split? | *Exam Split gives you the capability to "split" a series of patient images into separate groups. These new smaller image groups can be networked to desired reading stations for multiple "reads" and multiple billings on select patient exams. *Using Exam Split allows for split images from a single acquisition and assign them to a Requested Procedure ID or accession number retrospectively. *On the image Works desktop using Exam Split, all the images of the scan will be loaded. You can also use the mouse to select a range of images to be sent to a specific exam procedure. *At scan time, all patient records that you wish to have available to split to must be selected from the Patient Schedule when selecting New Patient. Your system is configured in one of two modes for Exam Split. * The mode configured is dependent on the capabilities of the system you are sending images to review
GACTS33 | SPLIT_EXAMS_WITH_CONNECT_PRO | exam split,scan applications | How to splfit exams with connect pro? | Step 1: From the Browser, click the Exam Rx icon. Step 2: Click the Patient Schedule icon. Step 3: From the Patient Schedule screen, select the desired patient exams to which youwantto apply Exam Split once the images are reconstructed. *A maximum of 15 procedures can be selected. Step 4: Click Select Patient.
GACTS34 | NEURO_3D_FILTER | neuro 3d filter,scan applications | What is neuro 3d filter? | *Neuro 3D Filter can be used to reduce pixel noise standard deviation or reduce dose while maintaining image quality. * It may be possible to reduce dose depending on the level of filtering applied.
GACTS35 | NEURO_3D_FILTER_CONSIDERATIONS | neuro 3d filter,scan applications | What are the considerations whilfe applying a neurro 3d filter? | * Nuero 3D Filter is not compatible with series that have duplicate slice locations. * To complete groups the following parameters must be the same: Slice Thickness, Interval, Tilt, SFOV, DFOV, Scan Type, Rotation Speed, image center, and algorithm. The groups must be contiguous. *Two series are reconstructed: the original and a series with Neuro 3D Filter applied.\ * Maximum number of images forNeuro 3D Filter is 500. * Images are annotated with the applied filter: N1, N2 or N3. * Images processed with Low N1- series plus 30. * Images processed with Med N2- series plus 40. *Images processed with High N3- series plus 50. * Images can be set to automatically transfer to up to four hosts. * Neuro 3D Filter is not compatible with ASiR.
GACTS36 | APPLY_NEURO_3D_FILTER | neuro 3d filter,scan applications | How to apply neuro 3d filter? | Step 1: Prescribe a scan for a neuro application. Step 2: Click the Recon icon , in the Auto Apps column for the group. Step 3:Click off to display the Session Selection screen. Step 4: From the Session Selection screen, under Filters, click Neuro 3D. Step 5: Set the Filter Strength. a. Click Low N1, Med N2, or High N3. b. ClickOK. Step 6: From the Filter Auto Transfer by Series screen, set Auto Transfer destination for up to four network hosts. a. ClickHost1. b. Click the network host in the list of availale destinations. c. ClickOK. Step 7: ClickConfirm. Step 8: Press Move to Scan. Step 9: Press Start Scan.
GACTS37 | SMART_PREP_SCAN | smart prep scan,scan applications | How to perform smart prep scan? | Step 1: From the ViewEdit screen, click SmartPrepRx. Step 2: Click On, if SmartPrep is Off. Step 3: To automatic transition from Monitor Phase to Scan Phase when Transition RIOI reaches threshold, click Dynamic Transition. Step 4: From the scout localizer on the display monitor, click and drag the red X on the slice line to define the Monitor location. Alternatively, type in the S/I location in the Monitor Location text field, for example, S110. Step 5: From the SmartPrep screen, enter the following values: * mA (range 10 to 100 in 10 mA increments), typically enter 40 mA for most studies. * MonitoringDelay (0to 60 seconds in1 secondincrements) sets the time before monitoring scans begin. * Monitoring ISD(1 to 60 seconds in1 secondincrements) sets the time between monitoring scans. *Enhancement Threshold (0 to 1,000 in steps of 1) sets the bar on the graph to monitor threshold enhancement. With Dynamic Transition is On, the Scan Phase starts automatically when the HU value of the Transition ROI reaches the Enhancement Threshold. * Diagnostic Delay, (range 3 to 60 in 0.1 second increments) typically set to the minimum value (automatic) for arterial studies and 10 seconds or longer for venous studies. * Allow enough time to deliver breathing instructions if Auto Voice is not being used. * The minimum time may vary when Auto Voice is turned on and depending on the monitor location relative to the start location. * If Scan Location is changed after the SmartPrep screen is closed, to accommodate the new table location, the Diagnostic Delay time is automatically updated to the minimum time possible when Auto Minimum Delay is on. Respond to the confirmation prompts that indicate a Diagnostic Delay change. * The Diagnostic Delay time is displayed in PrepGroup for Group1 onthe TimingTab,for example SP11.2. * AutoVoice Pre-message is available if Auto Voice is prescribed in the Voice Lights Timer inTiming parameters. * Auto Minimum Delay sets the Diagnostic Delay to the most minimum value possible, taking into consideration location of the monitor phase, start phase location, and Auto Voice Pre-message length. Step 6: From the SmartPrep scren, click Accept. Step 7: Proceed to Acquire a SmartPrep scan
GACTS38 | BASELINE_PHASE | smart prep scan,scan applications | How to aquire base line phase scan? | TheBaselinephase acquires one non contrast scan and allows you to establishan area to monitor contrast enhancement. Step 1: From the ViewEdit screen click the Confirm icon to proceed to scan. Step 2: Review and respond to the warning message. Step 3: Review the values on the Dyna Plan screen for all three phases: baseline, monitor, and phase. Step 4:Press Move to Scan to advance the table to the monitoring location. Step 5: Press Start Scan. Step 6: From the reconstructed image on the display monitor, click Ellipse ROI and place a maximum of three ROIs over areas of interest. For example, if scanning a liver, place an ROI in liver parenchyma away from vessels, if evaluating a vessel place the ROI over the vessel. * Use,asneeded,thezoom,roam,displaynormal,hide/showgraphics,erase,orexplicitmag(factor range: 0.5 to 2.0) from the SmartPrep display control panel.7. From SmartPrep Display control panel, select the ROI to be used for the Transition ROI. The last ROI placed on the image will be designated as the Transition ROI or set the desire ROI using Transition ROI selection in the SmartPrep Display control panel. * TheTransitionROIwillbedisplayedinyellowwheninactiveandinbluewhenactive,theptherROIs will be displayed in white.
GACTS39 | BASELINE_PHASE_FAILS | smart prep scan,scan applications | What to do if base line scan fails? | During Smart Prep, if the ROIs set on the baseline image fail to display during the monitoring phase, use the ROI statistics to monitor elapsed time and visually determine transition point to Scan Phase. If ROIs fail to display, they will not be on the screen saved SmartPrep image. Step 1:From the Scan Progress screen, click Monitor Phase. Step 2: Simultaneously press Start Scan and start the IV contrastinjection. A maximum of 40 monitoring images are available. * If all 40 monitoring scans have been used, you must press Move To Scan to move the table to the start location and press Start Scan to initiate the scan prescription plan. * The system waits the time set in the Monitoring Delay area and then begins acquiring images at the time set for the ISD1 * The total SmartPrep monitor scan time and Interscan Delay time between monitor images is displayed on the Real Time Information screen 3. View the display monitor and track the IV contrast graph. * Four viewports display from top leftto bottom right: baseline image,realtime images, graph,time stamp for each real time image. * A horizontal line indicating the enhancement threshold only appears on the graph if an ROI is placed on the baseline image. * The tick marks on the SmartPrep graph maynot show equal spacing,which is due to time between acquisition rounding issues. The SmartPrep scans are acquired at the correct timing interval, as prescribed
GACTS40 | NEW_IMAGES_FROM_SCANNED_DATA | retro reconstruction | How to create new images from scan data? | Step 1: Click Retro Recon. Step 2: From the Retro Recon List Select screen, select the appropriate patient, exam, and series. Step 3: Click Select Series. Step 4: On the Retro Recon screen, type a new series description in the Series Description area if desired. Step 5: Click New Series Numberto reconstruct images in a system-assigned new series number between 301 and 399. * If more than 399 series are already reconstructed, series 399 is repeated with a unique UID. * If not selected, the new series number will be original series number +100 (e.g., original series is 6, recon series is 106). Step 6: Click Retro to select all groups and click Yes or click Y for an individual group. * If you select Retro, a confirmation message,"Retro this group?" displays. Click Yes to proceed. *must determine which groups have data you wish to reconstruct. Step 7: Type in Retro Start and Retro End locations. * Cine images do not need a S or I designation. All others do. Step 8: Click Thick and select an image thickness. Click 0.625Z to generate overlapped 0.625 mm thickness axial images. Step 9: Click Time and type the Start and End times. * For cine, specify a time period. Step 10: Click Interval and type in an image interval. Step 11: Click DFOV and type in desired value. Step 12: Click R/L Center and/or A/P Center to set image offset. * Place the real time cursor in image to display the values in upper left corner of the image. You may also click List/Select to highlight the correct exam and series or use Graphic Retro recon. Step 13: Click Recon Option and select one of the options on the Recon Option screen. * For details, see the Set the ReconOptions procedure. * Set the W/L, ASiR level, Flip/Rotate if enabled. * If Flip/Rotate is applied, New Series should be selected. * Flip/Rotate and non Flip/Rotate images cannot be mixed in the same series. Step 14:Click Recon Type and select one of the available algorithms. * A window displays with all of the available algorithms from which you can choose the appropriate algorithm. * Soft and Standard are only available if Segment recon mode is selected. * Soft,Standard,Detail,Lung,Bone,BonePlusandEdgeareavailableif Fullreconmode is selected. Step 15: Click Graphic Retro Y to graphically prescribe the reconstructions, one group at a time. * Click and drag the blue circle on the reference image to the desired location. The SFOV is indicated by a red circle. Step 16 Click Confirm to generate images.
GACTS41 | GRAPHICAL_RETRO_SCAN | retro reconstruction | How to graphically prescribe retro scan? |  | https://drive.google.com/file/d/1SB9I6IWf5Crl8RDQSL4d8GrOvbIQxeAm/view?usp=sharing
GACTS42 | UNRECONSTRUCTED_IMAGES | retro reconstruction | How to find unreconstructed images? | Use this procedure to find the images that have not been reconstructed. Step 1: Click Recon Status to display the Recon Management screen. Step 2: Onthe Recon Management screen, clickDelete SuspendedEntries if youdo notwantto reconstruct the images orUnsuspend Entries. * Unsuspendentries retries the reconstructionandplaces images in the correct exam and series ifreconstructionwas successful. *selectedentry. Step 3: Click Update List to see images that may have been deleted or added. Step 4: Click Quit to exit.
GACTS43 | PAUSE_THE_QUEUE | retro reconstruction | How to pause reconstruction queue? | Images being prospectively reconstructed cannot be deleted from the queue. Images being produced retrospectively can be removed from the queue. Step 1: Click Recon Status to display the Recon Management screen. Step 2: On the Recon Management screen, click Pause Queue. * The recon queue can always be paused or restarted. * Only the entries listed in the Retro queue can be deleted. Step 3: Click Quitto exit Recon Management screen.
GACTS44 | RESUME_THE_QUEUE | retro reconstruction | How to resume reconstruction queue? | The reconstruction queue is not automatically started, you mustresume the queue forreconstruction to restart. Step 1: Click Recon Status to display the Recon Management screen. Step 2: On the Recon Management screen, click Restart Queue. Step 3: Click Quitto exit Recon Management screen.
GACTS45 | REMOVE_DATA_FROM_THE_QUEUE | retro reconstruction | How to remove data from the queue? | Use this procedure to remove data from the Retro Reconstruction queue. Step 1; Click Recon Status to display the Recon Management screen. Step 2: On the Recon Management screen, click Pause Queue. Step 3: ClickDelete Retro Recon Entries. Step 4: On the Delete Retrospective Entries screen, click on each job you want to remove and click Delete Selected. * Click Delete All to delete all data, independent of what you have selected. * If an entry is in reconstruction process, it may not delete. Restart the queue and try again. * Do not delete queued retros while scanning is active, scanning could stop. Step 5:ClickRestartQueue. Step 6: Click Quit to exit.
GACTS46 | RESERVE_OR_RELEASE_AQUIRED_DATA | retro reconstruction | How to reserve or release aquired scan data? | Use this procedure to prevent or allow the deletion of raw scan data. These instructions coverreserving the data. To release data, follow the same steps, except use the release buttons. Step 1:Click Recon Status to display the Recon Management screen. Step 2:On the Recon Management screen, click Reserve Scan data. Step 3: To reserve exams, click Reserve All Scanfiles, Select All Exams Listed or select individual exams. * If you Reserve All Scanfiles you cannot scan if the scan disk is full. Step 4: To reserve series, click Select All Series Listed or select individual series. Step 5: To reserve scans, click Select All Scanfiles Listed or select individual scans. Step 6: ClickReserve SelectedScanfiles. Step 7: Click Quit to exit.
GACTS47 | ANONYMOUS_PATIENT_SCAN_DATA | retro reconstruction | How to save anonymous patient scan data? | 1. Place a DVD-RAM media into the DVD-RAM drive or connect a USB3 device into the USB port on the front of the media tower. 2. Click Recon Status to display the Recon Management screen. 3. From the Recon Management screen, select USB orDVD. 4. Click Save ScanData. 5. Selectthe exam/series/scans youwish to save anonymously. 6. Click Save Selected Annon.Data. 7. Click Quit to exit. 8. Do not eject the device until the light is out on the DVD-RAM drive. 9. Once the save is complete to the USB device, exit Recon Management and disconnect the USB device. Then, scanning can resume.
GACTS48 | SAVE/RESTORE_SCAN_DATA_TO/FROM_A_USB/DVD-RAM | retro reconstruction | How to Save/restore scan data to/from aUSB/DVD-RAM? | Make sure the selected scan files do not exceed the space available on the DVD-RAM. * Scan data can only be restored on the same system model that it was acquired on. Step 1: Verify that you have enough time to complete the save uninterrupted. You cannot scan while saving scanfiles. * Saving a 60-second helical acquisition could take up to 1.0 to 1.5 hours on a DVD-RAM. Canceling will only stop when the current scan save is finished. Step 2: Connect the external USB device to the USB port or insert the DVD-RAM. Step 3: Click Recon Status to display the Recon Management screen. Step 4: Select the USB or DVD. Step 5: Click Save ScanData. Step 6: Select the exam you wish to save in the exam screen. Step 7: Select the individual scan files forthe series in the series screen. Step 8: Select individual scans or click the scans you want to save by clicking Select All Scanfiles Listed. Step 9: Optional, click Save Selected Anon. Data to save the data with no patient information in the image header. Step 10: Click Quit when the save is complete. * If cancel a Save or Restore of scan data, the cancel activates afterthe current scan file is saved orrestored. * If you are using an external USB hard drive, disconnect the USB device 1 minute after exiting Recon Management and the software indicates that scan data save is complete. Do not leave the USB drive connected once the transfer of scan data is complete. * Eject only after all the lights on the media are no longer illuminated. Scan File Save does not finish even though the Saved File message is posted. The problem is due to the time takes to it unmounttheDVD. It may take 5 to 7 minutes orlongerfortheDVDto unmount. Please note:the save is complete when the light on the DVD drive has turned green and then, the DVD is unmounted.
GACTS49 | REQUIREMENTS_FOR_REFORMAT | reformat | What are the requirements to be met before reformat? | Step 1:A valid image set for reformat must have the same:  Matrix size  Displaycenter  Orientation Step 2: Four or more images must be selected before selecting Reformat. Step 3:You cannot have two images with identical locations. Step 4:Spacing must be less than 10 mm. Step 5: The image set can only include axial, sagittal, or coronal images. Step 6: Different Display Field Of Views and Digital Tilts can be loaded, however you have to select between different sources of images to view them. The source annotation is located in the upper-left corner of the image. You can change between sources by clicking on the red annotation.
GACTS50 | REFORMAT_DO_NOT_MATCH_W/L_USER_PREFERNCES | reformat | What is to be done if the W/L1 values selected from the pull down menu in reformat do not match your W/LUsers Prefs? | Step 1: OpenImageWorks. Step 2: ClickUserPreferences. Step 3: Review saved values and correct them if necessary. See Set W/L preset preferences. Step 4: If this does notresolve the issue contact your service representative.
GACTS51 | OPEN_REFORMAT | reformat | How to open reformat? | Step 1: Click the ImageWorks icon. Step 2: From the browser, select the desired exam/series/images to reformat. * To select a subset of images, click on the firstimage and simultaneously press and hold the Shift, select the last image. * To select specific images click on the image numbers and simultaneously press and hold Ctrl. Step 3 :From the browser list of applications, click Reformat. * An oblique, axial, sagittal, and coronal are displayed in the four viewports. *My Tools tab can be customized to display any display icons. Therefore, the instructions in this chapter are for accessing a particular icon from the tab on which it originated. On your system, the icon may be on your My Tools tab. Step 4: Change the view type by selecting a view type from the red annotation.
GACTS52 | ROTATE_FREEHAND_OR_PAGE_IMAGES | reformat | How to rotate freehand or page images? | Step 1: Click the Rotation/Paging icon to activate a freehand rotation orto page through images. Step 2: Click and drag to rotate 3D/Oblique views and page axial/sagittal/coronal views. * No rotation handle is displayed on 3D/oblique views when this mode is selected.
GACTS53 | ZOOM_OR_MAGNIFY_IMAGES | reformat | How to zoom or magnify images? | Step 1: Click the Zoom icon to activate zoom mode. Step 2:Click and drag to zoom the image in or out.
GACTS54 | PAN_OR_ROM_IMAGES | reformat | How to pan or roam images? | Step 1: Click the Pan icon to activate a roaming mode. Step 2: Click and drag the image to move it up and down and left and right (only applicable if the image has been zoomed in).
GACTS55 | SELECT _ OBJECTS | reformat | How to select objects? | Step 1: Click the Selection icon to activate the selection mode
GACTS56 | W/L_IMAGES | reformat | How to activate w/l mode? | Step 1: Click the W/L icon to activate the W/L mode. Step 2:Click and drag over an image to adjust the window width (left-right motion) or the window level (updownmotion)
GACTS57 | CHANGE_THE_IMAGE_ORIENTATION | reformat | How to change image orientation? | Step 1:Click an Image Orientation icon to change the plane or a 3D or oblique reformat image. S = Superior I = Interior A = Anterior L = Left P = Posterior R = Right
GACTS58 | ACTIVATE_THE_OBLIQUE_MODE | reformat | How to activate oblique mode? | TheMultiple Oblique Mode button left) displays three oblique planes defined by three adjustable color axis (orange,green, blue). Step 1:Adjust any axis to update the two other oblique planes. Step 2: Click againto de-activate theMultipleOblique mode,which keeps the orientations defined in oblique viewports. TheSingleObliqueMode button(right) displaysa line cursorin Reformat, whichis used to define a newplane. Step 1: Set the function of one view port to Oblique. Step 2: Make another viewport primary and then click the Single Oblique Mode button. * A solid yellow line appears, which represents the plane of the Oblique reformat. Step 3: Place the cursor on the solid yellowline and click and drag it to tilt the yellow line to the desired plane.
GACTS59 | USING_RIGHT_CLICK_FUNCTIONS | reformat | How to use the right click functions? | Step 1:Place the cursor over the viewport and right-click to display the on-view menu. Step 2: Move the cursor down the menu until your choice is selected. Not all choices are available under all conditions.
GACTS60 | TEXT_ANNOTATION | reformat | What are text annotations used for? | *You can add notes and comments directly on the views, and use a marker with the annotation to point out anatomical features. *User text annotations can be edited, moved or deleted as necessary. *When images are saved, any annotations shown on the images are saved with them.
GACTS61 | PLANE_OR_VIEW_TYPE | reformat,red annotation | How to change plane or view type? | Step 1: Click and select from the drop down menu to change the plane or select a 3D mode
GACTS62 | IMAGE_LOCATION | reformat,red annotation | How to change image location? | Step 1: Press the middle mouse button and drag horizontally to scroll through the 2D images or left-click to increment and right-click to decrement one image at a time.
GACTS63 | DFOV | reformat,red annotation | How to use the magnifying factor? | Step 1: Press the middle mouse button and drag horizontally to magnify the image in real-time or left-click to increment and right-click to decrement the magnification factor.
GACTS64 | IMAGE_ROAM | reformat,red annotation | How to move the image within view port? | Step1:Press the left mouse button and move the image within the viewport.
GACTS65 | SLICE_THICKNESS | reformat,red annotation | How to change slice thickness? | Step 1:Middle-click and drag horizontally to change the slice thickness or in realtime or left click to increment and right-click to decrement the slice thickness (2Dimages).
GACTS66 | W/L_W/W | reformat,red annotation | How to change window width or window length? | Step 1: Press the middle mouse button and drag horizontally to change the W/L in real-time orleft click to increment and right-click to decrement the W/L. Step 2:Left click to select predefined WW/WL setting.
GACTS67 | PATIENT_NAME | reformat,red annotation | How to select patient name? | Step 1: Click and select from the menu to show or hide the patient name.
GACTS68 | REMOVE_A_TOOL | tools,reformat | How to remove a tool? | Step 1 : To remove a tool from My Tools palette, drag and drop the tool outside the palette and right-click the tool and select Remove from My Tools.
GACTS69 | ADD_A_TOOL | tools,reformat | How to add a tool? | Step 1: To add a tool to My Tools palette, go to the tab where to tool to add is, then drag and drop it to the My Tools palette or right-click the tool and select Add to My Tools (for example to add annotations preset tool go to Displaytab)
GACTS70 | MOVE_OR_HIDE_PALETTE | tools,reformat | How to move or hide the palette? | Step 1: The My Tools tab can be floated anywhere in the interface. Step 2:Click and drag the top part of the tab to move it. Step 3:Press the space barto show or hide at a mouse location.
GACTS71 | ANNOTATE_AN_IMAGE | reformat,annotate | How to apply preset annotation or unique or specific annotation? | Step 1:OpenReformat. Step 2: Click theDisplay tab. Step 3: Click the Annotation icon . Step 4:On the Annotation screen, click the annotation type (simple, linked or measure). Step 5:Place the cursor on the image and click to deposit the cursor and default text. Step 6: Enter new text in the Annotation text screen.
GACTS72 | PRESET_ANNOTATION | reformat,annotate | How to add custome preset annotatiion to preset annotation list? | Step 1:open reformat Step 2: From the Display tab, click the Annotation icon . Step 3: On the Annotate screen, type text in the Annotation Text field. * The preset name automatically fills in with part of the text from the annotation text field. Step 4:Change the preset name. Step 5: Select the text. Step 6: PressBackspace. Step 7: Type in a new name. Step 8: Click Use for other anatomy if you want the annotation available for other anatomies or Use for other protocol if you want the annotation available for other protocols. Step 9: Click Save as preset to add annotation to preset annotation list.
GACTS73 | MEASURE_DISTANCE | reformat,measure | How to measure distance? | Step 1:Open Reformat. Step 2:Click the Display tab Step 3: Click the Measure Distance icon. Step 4: Make one selection for each menu choice. * Measure: * From Volume (3D): to view the true distance in the 3D volume. Endpoints can be positioned at different depths in the volume. Rotate 3D volumes to check the exact position. * Projection (2D): to view the length of the projection. When used in 3D images, the measurement will correspond to the projected red distance in the plane of the screen. * Display: *On one Viewport: to view the measurement on only one viewport. * OnallViewports:to place a measurement ondifferent slicesor volumesof a multi-phase scan at the same time. * On one slice: to view the measurement on only one slice. * On all slices: to view the measurement on all slices. * Along: * Straightline: to deposit the first and second point. * Curve: to measure along a curve by depositing multiple points on the image. Right-click to validate the curved line. Step 35: Position cursorfor start of measurement and move cursor to end point of measurement and click.
GACTS74 | MEASURE_ANGLE | reformat,measure | How to measure angle? | Step 1:Open Reformat. Step 2:Click the Display tab Step 3:Click the Measure Angle icon. Step 4: Click and drag the squares to adjust the three points.
GACTS75 | MEASURE_AREA | reformat,measure | How to measure area? | Step 1:Open Reformat. Step 2:Click the Display tab Step3:Click the Measure Area icon. Step 4: Click and drag the squares to adjust the points around the region of interest. Step 5: Right-click to display results.
GACTS78 | MEASURE_VOLUME | reformat,measure | How to measure volume? | Step 1:Open Reformat. Step 2:Click the Display tab Step 3 Click the Measure Volume icon. Step 4: Segment an object of interest. Step 4: Click the viewport of the segmented object. Step 5: Click any of the corners to size the volume.
GACTS79 | COLOUR_MAPS | reformat,colour maps | How to manage colour maps? | Step 1:open reformat Step2 : Click theDisplay tab. Step 3: Click the colored ROI icon. Step 4: Click the image to deposit the ROI. Step 5: Click the color ramp displayed on the left side of the image to open the Color Table screen. Step 6: Click the Statistics tab to view range percentages for each color. Step 7: Change the values of the ROI on the viewport. a. Click the Configuration tab to change the values of the ROI on the viewport. b. Click a value in the Min Value column, and enter a new value. c. Do the same for the Max Value and Range name. Step 8: Select a color. a. Click a color in the Color column to view the Select new color screen. b. Drag and drop the cursor on the color wheel. c. ClickApply. Step 9:Click Add orRemove to change the viewed colors. Step 10: Click Presets tab to load and save Color map presets. Step 11: Click Save to save setting adjusted in the Configuration tab. Step 12: To load a color map preset, select the color ramp from the list
GACTS80 | ADD_ROI | reformat,display | How to add an roi? | Step 1: open reformat Step 2: Click the Display tab to view the ROI selections Step 3: Click the desired ROI tool. * 3D color ROI is available 1 color MR or multi-color CT. Step 4: Set the ROI statistics, default size, and display location on the ROI preferences screen. Step 5: Click the desired anatomy to deposit an ROI and view the measurements atthe bottom ofthe image. Step 6:Adjust the size by clicking and dragging on an ROI corner
GACTS90 | ROI_PREFERENCES | reformat,display | How to to set the type of information displayed in the viewports for an ROI? | Step 1: open reformat Step2: From the Main Control screen, click More Tools. Step 3:Click Display to view the measurement tools. Step 4: Select the desired view on which you want to place the ROI. Step 5: Click the ROI icon on the Review Controller to view the ROI preferences screen. Step 6: For a 3D ROI, one of the following options. * Click One Volume to view the VOI on only the viewport on which the VOI is displayed. * Click All Volumes to view the VOI and statistics on all viewports. Step 7: Click one or more Statistics option buttons to view any combination of minimum, maximum,average, standard deviation and area (2D) or volume (3D). Step 8: Type in the Vertical and Horizontal direction fields: (2D) o rROI height,width and depth(3D). The sphere VOI does not have a depth entry. Step 9: To hide the ROI statistics, click on the edge of the ROI to make it active, then at the center of the ROI, right-click on the center dot, and select Hide/Show statistics.
GACTS91 | CURVED_VIEW | reformat,display | How to display areformatof a curved, complex view of tortuous vesselsor organs? | Step 1: open reformat Step 2: In one of the viewports, right-click on the view type active annotation and select Curved. Step 3: On an axial, sagittal, or coronal image, press and hold Shift as you deposit points along the anatomy. Step 4: As you trace, the curved image updates automatically in the curved reformat viewport.
GACTS92 | X_SECTION_HISTOGRAM_VIEW | reformat ,view types | How to create an x section histogram view? | Step 1: open reformat Step 2: In one of the viewports, right-click on the view type active annotation and select Histo or X Section. * If you select a view type of histogram, you will immediately view a volume histogram which includes the entire 3D model. No other action is required. Place the cursor on the voxelreference line and hold Shift. The image displays the range of voxel class boundary lines by highlighting them in green. * If you select a view type of X Section, the view will display Undefined histogram until you start to define the trace. * Both the volume and cross-section histograms contain the same information with the exception of the name and measurement units. Cross-section histograms show total area, while volume histograms show the total volume forthe entire object without cut planes. Step 3: For a cross-section (X Section) histogram, in the viewport containing the anatomy of interest, press Shift and simultaneously click the image to deposit points on the area of interest to create a trace. Step 4: Click Exit. Step 5:Click OK.
GACTS93 | MPVR_VIEW | reformat,view types | How to create an mpvr view? | Step 1: OpenReformat. Step 2: Click the red annotation in the upper-left corner of the viewport and select 3D or MIP render mode. The selections available are dependent on the view type. * MIP is used most often in CTA models to demonstrate the most intense voxels. Step 3: Click the Oblique Mode icon. Step 4 :Click and drag the yellow line on the image. Step 5: Place the cursor over the red thickness annotation do one the following options. * Type the desired thickness and press Enter. *Middle-click and drag to the desired thickness. Step 6: Click Exit and OK to finish
GACTS94 | PROFILE_VIEW | reformat ,view types | How to create a profile view? | Step 1:OpenReformat. Step 2:Select the Profile Layout preset, if available. Step 3: In one of the viewports, right-click the view type active annotation and select Profile. Step 4: Onanaxial, sagittal, or coronal image, press andholdShift as youdepositpoints alongthe anatomy. Step:5 As you trace, the profile view displays the pixel intensity along the trace. Step 6: On the profile view: * The horizontal axis is the position in millimeters along the trace and the vertical axis is the pixel intensity values as a function of that position. * Press Shiftto display the pixel intensity from the location of the 3D cursor. Click and drag the white line to move it, which in turn, moves the cursor on the image. * The mean and standard deviation are displayed on the bottom. Step 7: Click Exit and OK to return to the browser.
| DATA_SET | reformat,view types,volume render | What are the effects of volfume rendering on data set? | It makes highly opaque objects more visible and at the same time it makes less opaque objects more transparent.
| ATTACH_OR_DETACH_OBJECTS | reformat,view types | How to attach or detach objects? | Step 1:OpenReformat. Step 2:With a VR image displayed, click the VR tab. Step 3: Click Advanced VR Settings. * Read the text on the tab for more information. Step 4:. Click a color to activate. Step 5: Click the Attach Mode icon. step 6: Click Attach at the bottom, center of the VR Controls screen. Step 7: Adjustthe opacity ofthe colored voxels by clicking on the colored box atthe topof the ramp and drag up or down. Step 8:. Attach another color by repeating the steps above.
| CUT_PLANES | reformat,view types, volume render | How to use cut planes? | Step 1:OpenReformat. step 2: Move the 3D cursor to the location of where you want the cut plane reference point. step 3: On a reformatted image, right-click on the No cut red annotation and select a plane cutting the 3D view. Step 4: Rotate the view to 3D view to display structures of interest in cut area. step 5: Select Cut planes off to deactivate.
vi
ew

## Sheet5
TID (unique) | KEYWORD(unique) | TAG(Separate with comma) | CATEGORY | PID(Separate with comma if more than one) | QUESTION | ANSWER | DOCUMENTATION LINK | YOUTUBE SOLUTION LINK
Chapter 12 | OPEN_IMAGE_WORKS |  | ACT imaging | ACT | What are the procedures to open image works? | Use this procedure to open ImageWorks from which you can access the browser program. 1. Click the ImageWorks icon to open the ImageWorks workspace 2. From the browser patient list, select the desired images. 3. From the browser list of applications, click Viewerto open the Viewer workspace. 4. Click the Display tab to open the Display control panel. You can also double click the series or thumbnail to start the viewer. If the number of images in a series is more than 20,000, then the viewer can not startthis series. Select an image subset from the patient schedule, then click Viewer to view these images. When finished viewing images, close Viewer and select the remaining images, then click Viewer. Further, you may also open two Viewers, every Viewer has one image subset.
Chapter 12 | ADD_ANNOTATIONS |  | ACT imaging | ACT | What are the steps to add annotations to images as a commentary? | Using the following steps can add Annotation to yourimages, allowing you to make commentary when landmarking, or draw attention to a certain area of interest. 1. OpenViewer. 2. Click the viewport for which you want to compose messages. 3. In the Display tab, click the Annotation icon. 4. In the Annotation text box, edit annotations as necessary.  Click and drag the empty red box to specify the size of the text box.  Move the cursor to a text box and enter a message. This tool automatically surrounds the text and expands textbox size as you type.  Move the cursor to insert a string of text and enter information.  Select existing text, you can also enter over it or press Delete. 5. If you must add or remove an arrow in the text box, click Select Object inthe Annotation screen orremove the selection Include Arrow.  Click and drag the arrow tip to position the arrow.  Clicking and dragging arrow lines will move arrows and boxes as a single unit.  Click and drag the solid red box to move the text box.  Click and drag the empty red box to change the size of the text box.
Chapter 12 | CHANGE_FONT_SIZE |  | ACT imaging | ACT | How to change the font size of the patient data? | Using the following steps can change the font size for the patient's name, identification number, registration number and scan date as displayed on the image or film,. When the view image format is larger than 1 X 1, these changes are most noticible. Figure 12-2 Model example of large fonts for patient name and scan date 1. OpenViewer. 2. From the Display tab or Film/Save tab, click the Tools icon. Select the appropriate tag in accordance with the fonts needed for changing film (Film/Save tab) or view (View tab). 3. In the menu, select Large Font. 4. Check the option box with the fields requiring the largerfont. 5. ClickApply. Clicking Save as Default can save annotation font size status after restarting.
Chapter 12 | MODIFY_ANNOTATION_LEVEL_DISPLAY_IMAGE |  | ACT imaging | ACT | How to modify annotation level of images displayed? | 1. OpenViewer. 2. On the Display tab, click the Display Preference Settings icon. 3. On the Set Display Preference screen, from the menu select Screen Annotation Settings. 4. Selectthe desired annotation option.  No Annotation deletes all annotations and graphics. It does not delete Window Width and WindowLevel. Select Customize to display item lists of what can be selected for display. Select the necessary annotation option, then click OK.  Viewthe Annotation screen to view the imageswith confirmed customized annotation area. 5. Click Save as Default to save annotation settings, in order to view the sessions in the future.
Chapter 12 | MODIFY_ANNOTATION_LEVEL_ON_FILM |  | ACT imaging | ACT | How to change annotation levels on the film? | 1. OpenViewer. 2. From the Film/Save tab, click the Flim Annotation Settings icon. 3. Selectthe desired annotation option.  No Annotation deletes all annotations and graphics. It does not delete Window Width and WindowLevel.4. Select Customize to display item lists of what can be selected for display. Select the necessary annotation option, then click OK.  Viewthe Annotation screen to view the imageswith confirmedcustomized annotation area. 5. Click Save as Default to save annotation settings, in order to view the sessions in the future.
Chapter 12 | EXPAND_ANNOTATION |  | ACT imaging | ACT | How to expand or copy annotations?. | 1. OpenViewer. Addannotations. 2. In the Display tab, click the Selected Object icon. 3. Enactthe necessary selection.  Select All Images in series to apply annotation changes to all images in a series.  SelectingThisImage Forwardonlyappliesannotationchanges to images inthe series following this serialnumber.  Select Range to designate a layer plate range. 4. ClickApply.
Chapter 12 | REMOVE_ANNOTATION |  | ACT imaging | ACT | How to remove annotations from the messages? | Using the steps below you can remove Annotations from images. 1. OpenViewer. Addannotations. 2. Ensure annotations are active (yellow). If not, please click annotations to put them into an active state. 3. In Erase Hide Menu, click the Erase icon. To removeall annotationmessages from all images in the current series, click the EraseAll icon.
Chapter 12 | MATTE_TOOL |  | ACT imaging | ACT | How to remove unwanted images/artifacts/messages using matte tool? | Using the Matte tool will remove displayed images or images on film or unwanted messages or artifacts surroundingthem. 1. OpenViewer. 2. Click the viewport for which you want to use the matte on the images. 3. From the Display tab, click the arrow nextto theMatte icon and select Rectangular Matte icon or Elliptical Matte icon. 4. Click and drag the solid red box to move the matte. Figure 12-9 Move Matte 5. Click and move the empty red box to change matte size and shape.6. Click on the outside of images (on the matte) to display the red box allowing you to adjust their size or move them. Click on the image (inside the matte) to fix the matte.
Chapter 12 | EXPAND_MATTES |  | ACT imaging | ACT | How to expand or generate copied mattesfor all images in a series? | You can use the following steps to generate copied mattes for all images in a series. 1. OpenViewer. Activate a matte. 2. Click on the outside of images (on the matte) to display the red box, indicating the matte is activated. 3. In the Display tab, click the Selected Object icon. 4. Carry out expand selection.  Select All Images in series for making changes to one image change all images in that series.  Selecting ThisImage Forward only changes the images in a series following the current series number.  Select Range to establish the layer plate range for matte copying. 5. ClickApply.
Chapter 12 | REMOVE_MATTE |  | ACT imaging | ACT | How to remove matte from images? | Using the steps below can remove Matte from images. 1. OpenViewer. Activate a matte. 2. Click on the outside of images (on the matte) to display the red box, indicating the matte is activated.  To remove an activated matte from a single image, execute one of the following operations. PressDelete. Click the Erase icon.  If you want to remove an activated matte from all images in the current series, click the Erase All icon.
Chapter 12 | ACTIVATE_MEASURE_TOOL |  | ACT imaging | ACT | What are the steps to activate a measurement area/tool to obtain information, volumes,areas and statistics of anatomy or pathology? | You can use the following steps to activate a Measure tool, in orderto obtain information, volumes, areas and statistics of anatomy or pathology. 1. OpenViewer. 2. Click the viewport where you want to place a measurement. 3. In the Display tab, click the Measure arrow and select a Measure ROI icon.  Measurement distance is shown in mm.  Measurement is displayed on the bottom right corner. 4. Adjust the size and shape of Measure ROI and move the Measure ROI.  For straight line and angle measurement: Figure 12-11 Adjust size, shape and move measurement Click and drag on empty to expand the box to adjust measurement shape. Click and drag the line segment to move the measurement.  For a free draw image measurement, place the cursor in the viewport and click to place a startingpoint. Figure 12-12 Free drawing measurement Click and drag on solid to expand the box to create trace lines. Click the borders of trace lines to move the trace lines. Click anyplace on a cyan (unactivated) measurement border, making it turn yellow and activated.
Chapter 12 | EXPAND_MEASUREMENT |  | ACT imaging | ACT | What are the steps to expand/generate copied measurement ROIs for all images ina series? | You can use the steps listed below to generate copied measurement ROIs for all images in a series. 1. OpenViewer. MeasurementArea. 2. Click anyplace on a cyan (unactivated) measurement border, making it turn yellow and activated. 3. In the Display tab, click the Selected Object icon. 4. Carry out expand selection.  Select All Images in series for making changes to one image change all images in that series.  Selecting ThisImage Forward only changes the images in a series following the current series number.  Select Range to specify the range of slice forthe extended measurements. 5. ClickApply.
Chapter 12 | REMOVE _MEASUREMENT |  | ACT imaging | ACT | What are the steps to remove measurement from images? | You can use the steps listed below to remove a measurement from images. 1. OpenViewer. MeasurementArea. 2. Click anyplace on a cyan (unactivated) measurement border, making it turn yellow and activated.  To remove an activated measuerement from a single image, execute one of the following operations. PressDelete. Click the Erase icon.  If you wantto remove multiple images from an active viewport, click the Erase All icon.
Chapter 12 | ACTIVATE _ROI |  | ACT imaging | ACT | What are the steps to activate ROI? | The steps below can be used to activate a ROI, in order to obtain information related to the standard deviation, average value and area. 1. OpenViewer. 2. Click the viewport where you want to place a ROI. 3. In the Display tab, click the ROI arrow and select a ROI icon. 4. Adjust ROI size, shape and move ROI.  For an oval or rectangular ROI: Figure 12-13 Oval and rectangular ROI Click and drag on empty to expand the box to adjust ROI size. Click and drag on the center cross or border to move the ROI. If the ROI box must be rotated, click a border having a cross hair cursor, place the cursor on a cross hair cursor, a white rotating symbol appears. Click and drag the rotation cursor to rotate the rotating frame pointer.  For a polygonal ROI, place the cursor in the viewport and click to place every point. Figure 12-14 Polygon ROI Click and drag a square box to adjust ROI shape. Right button click on a box to remove this box. Click and drag a center cross to move ROI. Place the cursor on a line between two adjacent solid red boxes and right-click to place anotherbox. Click anyplace on a cyan (unactivated) ROI border, making it turn yellow and activated.
Chapter 12 | EXPAND_ROI |  | ACT imaging | ACT | What are the steps to be followed to expand/generate copied measurement ROIs for all images in a series? | You can use the steps listed below to generate copied measurement ROIs for all images in a series. 1. OpenViewer. ActivateROI. 2. Click anyplace on a cyan (unactivated) ROI border, making it turn yellow and activated. 3. In the Display tab, click the Selected Object icon. 4. Carry out expand selection.  Select All Images in series for making changes to one image change all images in that series.  Selecting ThisImage Forward only changes the images in a series following the current series number.  Select Range to specify the range of slices to duplicate ROI.  ClickApply
Chapter 12 | REMOVE _ROI |  | ACT imaging | ACT | How to remove ROI from images? | Using the steps below can remove a ROI from images. 1. OpenViewer. ActivateROI. 2. Click anyplace on a cyan (unactivated) ROI border, making it turn yellow and activated.  To remove an activated ROI from a single image, execute one of the following operations. PressDelete. Click the Erase icon.  If you want to remove multiple ROIs from an active viewport, click the Erase All icon.
Chapter 12 | ADJUST_W/L_VALUES |  | ACT imaging | ACT | What are the methods to change window width and window length values therefore modifying brightness and contrast? | Use one of the fourfollowing methods to adjust Window Width/Window Level, in orderto control image brightnessandcontrast. Requirements OpenViewer. Method 1: Middle mouse button By using the steps below you can make use of the middle mouse button to adjust Window Width/Window Level. 1. Placethe cursoronanyimage,click themiddlemousebuttonanddragupordowntochangewindow level. 2. Placethe cursoron anyimage, click the middle mousebuttonanddrag leftorrightto changewindow width. Method2: Left mouse button Youcanuse thesteps listedbelowtomakeuseoftheleft mousebuttonforadjustingWindowWidth/Window Level. 1. Click the W/L (Window Width/Window Level) icon on the Image Controls Panel, change the left mouse button into the Window Width/Window Level control keys. Figure 12-15 Image controls panel 2. Placethecursoronanyimage,clickthe leftmousebuttonanddragupordowntochangewindowlevel. 3. Placethecursoronanyimage,clicktheleftmousebuttonanddragleftorrighttochangewindowwidth. Method 3: Presets You can use the steps listed below to make use of the presets for adjusting Window Width/Window Level. 1. Place the cursor in the lower right corner of the viewport's orange window width and window level position. 2. Click the right button and select a preset window width and window level.Method 4: Control panel By using the steps below you can make use of the Keyboard to adjust Window Width/Window Level. 1. Click the W/L (Window Width/Window Level) icon on the Viewer Control Panel. 2. Enter a specific value in the window width and window level text box. Method 5:Keyboard By using the steps below you can make use of the Keyboard to adjust Window Width/Window Level. 1. Place the cursor on any viewport. 2. Press and hold or quickly press and release the up/down arrow keys to change Window Level. 3. Press and hold or quickly press and release the left/right arrow keys to change Window Width. Figure 12-17 Arrow keys 4. Press the shift key and one of the following keys to edit preset W/L: F6, F7, F8, F9, F10, F11. 5. Press one of the following keys to activate the preset window: F6, F7, F8, F9, F10, F11.
Chapter 12 | PREDEFINE_W/L_VALUES |  | ACT imaging | ACT | How to self define or predefine preset settings like for window width /WindowLevel? | You can use the steps below to self-define W/L (Window Width/Window Level) preset settings. 1. OpenViewer. 2. In the Display tab, click the W/L (Window Width/Window Level) icon. 3. In the Windowing area, select a preset value. 4. Click Edit. 5. In the Apply Windowing area, complete the following operations: a. Select the preset from the Presets menu. b. If required, in the Name field, enterthe preset W/L name. c. Please enter a value in the Width field. d. Please enter a value in the Level field. e. Click Save and give a response to any prompts.
Chapter 12 | ADD_LAYER |  | ACT imaging | ACT | How to add a layer / cross reference line in an image? | You can use the steps below to add a Cross Reference line to an image already displaying having a cross reference. 1. OpenViewer. 2. Click the viewport where you want to add additional cross reference scan lines. 3. In the Command Line, enter xra, then enter the series number you want to add and the layer plates orlayer plate range, for example if you want to add layers 10-16 from series 3, enter xra 3 10-16. Press the Enter key.
Chapter 12 | ADD_ALL_LAYER |  | ACT imaging | ACT | How to add a cross reference line in all layers in a series? | You can use the steps below to place a Cross Reference line on all layers in a series. 1. OpenViewer. 2. Click the viewport where you want to paste up cross reference scan lines. 3. In the Command Line, enter the xr command, follow by adding the series number, for example: xr 2. Press the Enter key.
Chapter 12 | ADD_FIRST_OR_LAST_LAYER |  | ACT imaging | ACT | How to add cross reference line to first or last layer images? | You can use the steps below to apply a first or last layer Cross Reference line to images. 1. OpenViewer. 2. Click the viewport where you want to paste up cross reference scan lines. 3. In the Command Line, enter the xr command, follow by adding the series number and extrema, for example: xr 2 extrema. Press the Enter key.
Chapter 12 |  |  | ACT imaging | ACT | How to add cross reference line to an interval of layers like even/add? | You can use the steps below to place a Cross Reference line on the intervals of a series or even/odd layer positions. 1. OpenViewer. 2. Click the viewport where you want to paste up cross reference scan lines. 3. In the Command Line, enter the xr command, follow by adding the series number, the colon symbol (:) and interval. For example, if you want to display every third layer in series 2, please enter xr 2:3. Press the Enter key.
Chapter 12 |  |  | ACT imaging | ACT | How to add cross reference line to a range of layers? | You can use the steps below to display the cross reference line series of images in a common series. 1. OpenViewer. 2. Click the viewport where you want to paste up cross reference scan lines. 3. In the Command Line, enter the xr command, follow by entering the first image-last image, for example: xr 2 1-6. Press the Enter key.  If necessary, use the % symbol to mark the last image in a series.
Chapter 12 |  |  | ACT imaging | ACT | How to remove cross reference line from any layer? | You can use the steps below to remove a Cross Reference line placed on an image. 1. OpenViewer. 2. Cross reference lines displayed on an image. 3. In the Command Line, enter noxr. Press the Enter key. All reference lines have been removed.
Chapter 12 |  |  | ACT imaging | ACT | How to turn pages of an image like watching video using cine tool/cine play program? | Use the Cine tool to turn pages in an image, like you're currently watching a video. 1. OpenViewer. 2. Select the series you want to view in Cine mode. 3. In the Display tab, click the Movie icon. 4. In the Cine area, click the Play icon,this will use all default settings. Additionally, you can conduct selections from the Cine area.  The system will automatically change the format to 1-on-1.  Define layerrange Click Play All to place all images in a series into the Cine. Click Play Range and enter a From and To image number in the text.  Define Image Interval. This is usually maintained at the default value 1. When displaying a cine from a multi-phase series, within the image range enter a number representing the number of phases.  Select an image interval, this interval will allow you to skip some images in the series when the images are being displayed in cine mode. For example, if your series is a 4 phase acquisition, you may want to view one phase one time, therefore you will enter a 4 interval.  Select a view mode. When displaying a Cine, the Loop icon will go from start to end, then again from start to end. For example, if there are 20 images in a Cine, the image display method is 1-20, 1-20, 1-20, etc. When displaying a Cine, the Rock icon will go from start to end, then again from end to start. For example, if there are 20 images in a Cine, the image display method is 1-20, 20-1, 1-20, 20-1 etc.  Select a Frame Rate, it indicates the number of images displayed in a second.  Click the Play icon to play backwards or forwards. From the cine stop or pause point start playing the cine.  Click the Pause icon to stop the movie from playing.  Clicking the End icon can take the last image in a cine and place it in the active viewport.  ClickingtheBeginning iconcantake thefirstimage ina cineandplaceitintheactiveviewport.  Clicking the Step Forward icon can increase a step from forwards to backwards to an image. Using this icon can increase a step from the image to a certain start point of the first image not within the cine layer range.  Clicking the Step Back icon can increase a step from backwards to forwards to an image. Using this icon can increase a step from the image to a certain start point of the first image not within the cine layer range.
Chapter 12 |  |  | ACT imaging | ACT | How to perform command line programming? | 1. OpenViewer. 2. At the bottom of the Control Panel, in the Command Line enter ? and press the Enter key.  The Command Line screen opens, listing all hotkey commands. 3. In the Command Line window, select the desired command. 4. ClickAccept. 5. Place the cursor at the end of the command character string in the Command Line, and press the Enter key.  This command is then applied to all displayed viewports.  As next and priorimages are selected,the command will be applied to allof these images as well.
Chapter 12 |  |  | ACT imaging | ACT | How to use color mapping tool to change color of all viewports? | Use the Color Map tool for scheduled acquisitions, DWI ADC images, histograms and dynamic enhanced viewing, in order to change the color of all viewports. 1. OpenViewer. 2. In the Display tab, click the Color Map icon and select an option from the menu. To remove a color map, from the Color Map menu, select none or gray.
Chapter 12 |  |  | ACT imaging | ACT | How to open multiple images for comparisom using image compare program? | By using the steps below, you can open multiple grouped images in the Viewer for comparison. Comparison mode can be started through the three following kinds of workflows: From the Data Apps list start and compare 1. Navigate to the ImageWorks desktop. 2. Press and hold the Ctrl key, and at the same time in the Series List click 2, 3 or 4 series.  The selected series will be highlighted and displayed in the Series List. 3. In the Data App list, click the Viewer, in order to display the selected series in a single viewport. View Compared Images 1. Pressing Page Up and Page Down can at the same time go up the page or down the page in these series. 2. The Window Width/Window Level of every viewport can be controlled separately. 3. The Maximize button in View Controller can not be used in the comparison mode. 4. All available options in the Viewer Control Panel can be controlled in every viewport. 5. Page turning can only be used in two series. Undo Link/Re-link viewport 1. Click the viewport to be viewed alone. 2. In the View Controller, click the Ports icon to undo a series link. Figure 12-22 View Controller linking  Pressing the Page Up button and Page Down button can scroll through the browser series. 3. If you need to re-link a viewport, please click on the Ports icon.
Chapter 12 |  |  | ACT imaging | ACT | How to restore an image to its original size using display normal procedure? | You can use the steps below to restore an image to its original state. 1. OpenViewer. 2. In the Viewer Display tab, clicking the Display Normal icon can return the image to its original image state and erase the display functions below:  Flip/rotate  Zoom in or zoom out  Imageenhancement(filter)  Scroll or pan  Color DisplayNormal has no effect on the Rectangular Matte, Annotation or Graphic operations. Using the Erase All button can remove these elements from images.
Chapter 12 |  |  | ACT imaging | ACT | How to erase/hide extra object programs in viewports? | Using this program when one or more graphical objects are placed in any or all viewports, can Erase/Hide theobjects. 1. OpenViewer. 2. Display an image object,for example an Annotation information, Measurements, Mask, ROI string, etc. 3. Click the arrow next to the Erase/Hide icon to select the icon that matches the desired operation.  Clicking the Hide All icon can temporarily hide annotations orimage objects.  Clicking the Show All icon can restore annotations orimage objects, making them visible.  ClickingtheErase Single iconcanerase thecurrently selectedimage object.Usingthiserase option can erase one graphics element once.  Clicking the Erase All icon can remove all graphical objects from all viewports. Clicking Ctrl and X at the same time can erase the selected graphical objects.
Chapter 12 |  |  | ACT imaging | ACT | How to smooth images using filter image program? | Usingthe steps belowcan use one ofthe Filters in the Viewerinorderto enhance margin orpulmonary area orrenderimages smooth. 1. OpenViewer. 2. Click the viewports where you want to change the filter. 3. In the Display tab, click the Filter icon. 4. Select a filterfrom the menu.
Chapter 12 |  |  | ACT imaging | ACT | How to change direction of displayed images using flip/rotate image program? | Use the Flip/Rotate tool to change the direction of the displayed images. 1. OpenViewer. 2. In the Display tab, click the Flip/Rotate icon. 3. In the Flip and Rotate Images areas, click the icon matching the desired operation.
Chapter 12 |  |  | ACT imaging | ACT | How to change viewport display format? | Using the steps below can change the viewport display format. 1. OpenViewer. 2. In the Display tab, click the Format icon. 3. On the Format display, scroll the cursorto view the available formats. 4. Click to select the desired format.
Chapter 12 |  |  | ACT imaging | ACT | How to enhance contrast of an image by giving it a black and white appearance using gray scale enhancement contrast program? | Use the Gray Scale tool to enhance the contrast of an image in order to give the image more of a black and whiteappearance. 1. OpenViewer. 2. Click the viewports where you want to change the gray scale. 3. In the Display tab, click the Gray Scale icon. 4. In the Filters area, select a Gray Scale filter.
Chapter 12 |  |  | ACT imaging | ACT | How to place a grid on the main image using grid display procedures? | Usingthe stepsbelowcanplaceagrid(matrix)onthemainimage, inordertomeasuretheanatomicalpoints or pathological area on an image. 1. OpenViewer. 2. Click the viewport in which you want to display the grid. 3. In the Display tab, click the Grid icon to open the grid, and then click again to close it. Figure 12-24 Grid on Remove grid Using the steps below can move the grid in the viewport. 1. Click a grid, making it active status (blue). 2. Place the cursor on a grid cross reference point and drag it to a new position.
Chapter 12 |  |  | ACT imaging | ACT | How to adjust grid preferences? | Using the steps below can set grid preferences. 1. In the Display tab, click the Tools icon. 2. From the pull=down menu, select Grid Preferences. 3. From the Grid Preferences screen area, select the options you need.  Matrix Lines: Open or close the lines for the four-corner grid pattern formed by the central horizontaland vertical lines.  Line Style: Turn gird lines into solid lines or virtual lines.  Grid Spacing: In units of millimeters, changes the spacing between grid lines.  Tick Spacing: In units of millimeters, changes the spacing between tick lines.  Tick length: In units of millimeters, changes the length of tick lines. 4. Clicking Save as defaults can permanently save the content of the selections.  You can also click Apply to apply the selected options. This is a temporary selection, only applicable to the currently displayed image.
Chapter 12 |  |  | ACT imaging | ACT | How to change control of left mouse buttpn using image control program? | Using the steps below you can change the control of the left mouse button. 1. OpenViewer. 2. Select the Image Controls icon on the Control Panel. Figure 12-26 Image Controls  Clicking the Slice number control icon can use the left mouse button to scroll through the images in the current series. Click and drag up/down or left/right in order to advance to higher or lower series numbers.  Clicking the Magnifying Glass icon can use the left mouse button as a magnifying function. Click and drag up to increase the magnification factor, click and drag down to decrease the magnificationfactor. Magnification factoris annotated on the upperright area of an image. Zoom is applied to the images in the active viewport and the viewport forward.  Clicking the Scroll icon can use the left mouse button to scroll images. Click and drag images in the active viewport to a new location.  Clicking the Selection icon can use the left mouse button for a selection function and disable all other image controls.  Clicking the W/L (Window Width/Window Level) icon can use the left mouse button for Window Width and Window Level imgaes. Click and drag the cursor in the directions below to changeWindowWidth/Window Level: Move up to make image brighter(reduce Window Level value) Move down to make image darker(increase Window Level value) Move right to make Window Width wider Move left to make Window Level more narrow.  ClickingtheMinimize/Maximize iconcan change the viewport area to a 1-on-1 display, then resetformat, for example, change to a 4-on-1 display.
Chapter 12 |  |  | ACT imaging | ACT | How to select images displayed in the viewer using image select program? | Use one of the programs below to select the images displayed in the Viewer. Method1:Active image annotations 1. OpenViewer. 2. Place the cursor on an orange Image Number annotations on the upper left of a viewport. Figure 12-29 Orange image number annotations 3. Click to progressively add current images. 4. Right button singel click, decrease current images. Method 2: Command Line 1. Open Viewer. 2. Place the cursor on the Command Line. 3. EnterI XX and press Enter, here XX is the desired image number. This image number will appear in the left side of an active viewport. Method 3:Keyboard 1. OpenViewer. 2. Press Page Up to add the current image. 3. Press PageDown to decrease the current image. 4. Press Home to display the first image. 5. Press End to display the last image. Method 4:Image controls panel 1. OpenViewer. 2. Clicking the Image Page icon (located in the Image Controls Panel) to change the left mouse button into the layer number control. 3. Place the cursor in the viewport, then drag up orright/down or right/left, to advance to a higher or lowerimage serial number.
Chapter 12 |  |  | ACT imaging | ACT | How to invert image display using image invert program? | You can use the steps below to invert image display. Method 1: Window Width and Window Level icon 1. OpenViewer. 2. On the Display tab, click the Window Width and Level icon. 3. In the Windowing area, click the Invert icon.  The Invert icon is a toggle icon; click it again to reset it. Method 2: Command Line 1. OpenViewer. 2. In the Command Line, enterinvert.  Clicking invert again can reset the icon and return it to its normal state.
Chapter 12 |  |  | ACT imaging | ACT | How to change display format to 1 on 1 and return format to previously selected format using maximize/minimize tool? | UsingtheMinimize/Maximize tool canchange the formatto 1-on-1display (Maximize),the returnthe format to the previously selected value. 1. OpenViewer. 2. In the Display tab, click the Minimize/Maximize icon to change the viewport area to a 1-on-1 display. Figure 12-30 1-on-1 display 3. Click the Minimize/Maximize icon again, restore format. In this example,restore a 4-on-1 display.
Chapter 12 |  |  | ACT imaging | ACT | How to copy/expand objects from viewports using copy object program? | Mask,Measurements, ROIs andAnnotation program functions all have copying functions. Using the following steps can use the Command Line to expand objects in all or selected viewports. 1. OpenViewer. 2. Generate a graphic:Annotation,Mask,Measurements, ROI,Grid. 3. Click thegraphic youwantto copy (annotationbox, matte, measurement, ROIs, grid) and make itinto an active state (yellow). 4. In the Display tab, click the Selected objects icon and select the propagation option.  Select All Images in series to apply annotation changes to all images in a series.  SelectingThisImage Forwardonlyappliesannotationchanges to images inthe series following this serialnumber.  Select Range to designate a layer plate range. 5. ClickApply.
Chapter 12 |  |  | ACT imaging | ACT | How to backup command lines? | 1. Place cursor in the Command Line. 2. Enter one of the following commands:  Enter prop a and press Enter, copying graphics in all viewports. To erase copied objects, in the Erase Hide Menu, click the Erase All icon.
Chapter 12 |  |  | ACT imaging | ACT | How to display,define an image in a small viewport according to image using reference image activation program? | Use the Reference Image tool to display an image, define the original image within the small viewport according to the image. 1. OpenViewer. Select the image that you want to view. 2. Clicking Ref Image > All On to display reference images on all images in a series.  Select Selected On to only display a reference image on the primary viewport.  Reference images must be on a different level from the main image.  These two images (main image and reference images) must come from the same exam, and must have the same horizontal boundary line, the same patient location and the same patient entry method.  Viewportimages and reference images cannot be Screen Save, 3D rendered images or members of an image group.  All Reference Image viewports display the same image. You can not place a different image in a single or different viewing areas. If an image in a Reference Image viewport is changed, all other Reference Image viewports will update to the same image.  Inverse Video at the same time effects the main image and reference images.  Click the Reference image to change focal point and adjust window width/window level, magnification, panning, flip/rotate and display normal. When you initiate these changes on a reference image viewport, all reference images will update.  Series Binding cannot be applied to reference images. If you want to erase reference images from a series, please click Ref Image > All Off.
Chapter 12 |  |  | ACT imaging | ACT | What are the steps to display acursor,reportinganatomicalposition and CT valueon the upperleft corner of an image using report cursor procedure? | Youcan use the steps belowto display acursor,reportinganatomicalposition and CT valueon the upperleft corner of an image. Figure 12-32 Cursor Annotations 1. OpenViewer. 2. In the Display tab, click the Tools icon. 3. From the menu, select Guides Settings. 4. SelectReport Cursor. 5. Click and drag the cursor to view updated location and CT values. Clicking Save as Default can save the report cursor status during the restart process.
Chapter 12 |  |  | ACT imaging | ACT | What are the steps to change the left mouse button into the scrolling or panning control in order to move images in a viewport using scroll images program? | Using the steps below can change the left mouse button into the scrolling or panning control in order to move images in a viewport. 1. OpenViewer. 2. From the Image Controls Panel, click the Scroll icon. 3. Click and drag images in the active viewport to a new location. Clicking the Display Normal icon can cancel scrolling on the images.
Chapter 12 |  |  | ACT imaging | ACT | What are the steps to define the layout of a viewport, and from which every main viewport can define a format using screen layout program? | Using the steps below can define the layout of a viewport, and from which every main viewport can define a format. 1. OpenViewer. 2. In the Display tab, click the Screen Layout icon. 3. In the Screen Layout area, select one of the available main screen layouts.  1-on-1 2-on-1verticalsegmentation  2-on-1horizontalsegmentation  4-on-1 4. In the Display tab, click the Format icon and select the desired format.  This image format will be applied to the selected main viewport. 5. Select the main viewport where you want to place the screen layout image. This operation will activate the viewport. 6. In Data Selector, select the desired series and click View.  Images in the selected series appear in the selected format in the current viewport. 7. Since each main viewport can use a single series and format, please repeat steps 4-6 for every main viewport. Like the example below selected a 4-on-1 screen layout and every main viewport adopts a single format.
Chapter 12 |  |  | ACT imaging | ACT | How to reset viewports? | 1. Select Format icon and select 1-on-1 format display. 2. Select the Screen Layout icon and select 1-on-1 viewportformat.
Chapter 12 |  |  | ACT imaging | ACT | How to save an image in a series using screen save image program/tool? | Use the Screen Save tool to save a image in a series. Saved images can be viewed from the ImageWorks workspace. Figure 12-34 SSAVE Series 1. OpenViewer. 2. Click the image to be screen saved to make it active. 3. Click the Film/Save tab. 4. Click the Screen Save icon.  On the Screen Save image you can still adjust W/L (Window Width/Window Level) and magnificationvalues.  Once a screen is saved, all random annotations to an image can not be erased, because it is a "snapshot" of the screen and cannot be changed.  Screensavedimageshavepost-processinglimits,forexamplefilteringcannotbeperformedon them.  Screen save images saved in a SSAVE screeen save series, they can perform archiving and networktransferring.
Chapter 12 |  |  | ACT imaging | ACT | How to activate/deactivate series binding? | Using the steps below can turn Series Binding on or off. 1. OpenViewer. 2. In the Display tab, click the Tools icon. 3. From the tools pull-down menu select Series Binding. 4. Select ON or OFF.  When Series binding is open, after the last layer in a series, click Image +(next image) or press Page Down to display the first layer of the next series.  When Series binding is off, after the last layer in a series, click Image + or press Page Down to display the lastlayer of the same series. And it will not automatically advance to the next series. 5. Click Save as defaults to permanently save selected content.
Chapter 12 |  |  | ACT imaging | ACT | How to open the viewer from image works? | Using the steps below can view sessions and open the Viewer. 1. Navigate to the ImageWorks desktop. 2. Select the desired exam, series or image. 3. ClickViewer.
Chapter 12 |  |  | ACT imaging | ACT | How to use Text Page tool to displaya specific patient exam's Exam TextPageor Series TextPage or ROItext page using text page display program? | Usingthe Text Page tool can displaya specific patient exam's Exam TextPageor Series TextPage or ROItext page. 1. OpenViewer. 2. Click the Film/Save tab. 3. Click the Text Page icon. 4. Click the desired Text Page. Exam Text page, Series Text page, ROI text page. 5. Select one option from among those at the bottom of the page.  Clicking the Film iconcanplace the page in the nextavailable film frame in the film desinger.  Clicking the Screen Save icon can add the SSAVE image type to the patient exam.  Use theNext/Prior arrow to navigate the text page.  Click the Quit icon to close the window.
Chapter 12 |  |  | ACT imaging | ACT | What are the steps to use small ruler vertical or horizontal Tick Marks on your image using tick mark display program? | Using the steps below can use small ruler vertical or horizontal Tick Marks on your image. 1. OpenViewer. 2. In the Display tab, click the Tools icon. 3. In the Set Display Preferences area, from the menu selectGuides Settings. 4. Select Vertical Tick Marks and/or Horizontal Tick Marks to open them. Figure 12-35 Tick marks on 5. Clicking Save as Default can save restarted tick mark status. Tick marks is a toggle option. Once they are opened, they will always maintain an On status, until you select this option again and turn it off.
Chapter 12 |  |  | ACT imaging | ACT | How to Zoom the image, in order to best display the targeted area. If the selected view is set as the main view, and axillary views using zoom image program? | Chose one of the following program to Zoom the image, in order to best display the targeted area. If the selected view is set as the main view, and axillary views follow after, then zoom will be applied to all images in the series. If the selection of other views is canceled, then zoom is only applied to the main view. Method 1: Display Tab 1. OpenViewer. 2. Click the viewport for which you want to zoom the images. 3. In the Viewer Display tab, click the Zoom icon. 4. In the Zoom area, enter a factor in the text box or use the slider to determine the zoom factor. Click the DisplayNormal icon to cancel the zoom on the image. Method 2:Image controls panel 1. OpenViewer. 2. Click the viewport for which you want to zoom the images. 3. From the Image Controls Panel, click the Zoom icon. 4. Place the cursor in an active viewport, click and drag to change the zoom factor. Click the Display Normal icon to cancel the zoom on the image. Method3: ScreenAnnotations 1. OpenViewer. 2. In the image viewport, place the cursor on orange Mag and click the right mouse button to display the Zoom slider in the image viewport. 3. Clicking and dragging the slider can change the zoom factor. 4. Click anywhere in the viewport except the slider to close the slider.Method 4:Magnifying Glass 1. OpenViewer. 2. Click the viewport for which you want to use the Magnifying Glass. 3. In the Image Controls Panel, click the Magnifying Glass icon. 4. Place the cursor in an active viewport, click and drag to view the magnified area.
Chapter 12 |  |  | ACT imaging | ACT | How to use Save State tool to save the current window width and window level, image controls and all image annotations and graphics to a GSPS in the Patient List using save state program? | Use the Save State tool to save the current window width and window level, image controls and all image annotations and graphics to a GSPS in the Patient List. MethodOne 1. OpenViewer. Select the image that you want to view. 2. Display images as per the necessary state.  Places annotation text far away from the viewport margin. If the edge of the text box is near the edge of the viewport, then the displayed image on another viewer (for example, a PACS viewer) may not display the text box. 3. Place the cursor in the command line and enter ss, press the Enter key to save image state.  PROSP series and GSPS series can perform network transfers and archiving, so that when they are displayed on remote systems, they can display saved image controls and images. These two types of series must be simultaneously sent to the terminal destination. If only the original image is sent, then the image controls will not be sent, because the image controls are located in the GSPS series.  Archiving and network transfers of original series and GSPS series can be performed both manually and automatically. If you want to perform automatic archiving or network transfer of GSPS series, you must open Auto Archive or Auto Transfer by Exam, then create a GSPS series before the exam is completed. You must manually perform archiving or network transfer on any GSPS series acquired after an exam has been completed. MethodTwo 1. On the Film/Save tab, click the Save State icon. 2. Entertext as necessary in the Presentation Label text box. 3. Selectimage range. 4. Click Save State. View Save State image (GSPS) 1. Openthe ImageWorks desktop. 2. Select a GSPS series. Figure 12-36 Example of single GSPS series during the exam 3. ClickViewer.
Chapter 13 |  |  | ACT imaging | ACT | What are the steps to open add/subtract screen? | Ontheright sideofthebrowser, clickAdd/SubstracttoopentheAdd/Substract(ImageCombination) screen. Figure 13-1 Add/Substract screen Figure 13-2 Image Combination area Select Set 1 Defines the first image set. Select Set 2 Defines the second image set. Equal(=) Performs the selected operation and generate the new images. Ratio slider Increases pixelweightingof the left(first) image when moced to the left and increases pixel weighting of the right(second) image when moved to the right. The slider is only available during addition and subtraction when both Select Set buttons are selected. Save Series Displays the current series andstartingimage numberofthesavedseries.By default, images resultingfrom subsequent operations are added onto the end of the same series unless you click New Save Series.Figure 13-3 Save Series area New Save Series Places images resulting from subsequent operations into a new series in the Patient List. If you do not select this option, the images are saved in the current saved series. Mode Figure 13-4 Mode buttons Max creats a new series that finds the highest corresponding signal intensity values in the selected set or sets of images. Plus (+) creates a new series that adds the image intensity values of the selected image sets. Minus(-) creates a new series that subtracts the image intensity values of the selected image sets. Mincreatesanewseries thatfinds thelowestcorrespondingsignal intensity values inthe selectedsetorsets ofimages. Bind creats a new series that consists of copies of the selected images from one or more existing series. Accept Negative Pixels Allows negative pixel values in the resulting images. If this function is not enabled, all negative pixel values are set to zero. This operation is useful on pre-and post- contrast studies of the same slice location. Clear Selection Erases the selections so that both Select Set buttons have no prior values associated with them.
Chapter 13 |  |  | ACT imaging | ACT | What are the procedures to create a new image set using add/subtract? | Use this procedure to create a new image set, for example add thin slices together or subtract pre/post contrast series. The resulting images are annotated with the day on which the addition/subtraction was performed. OpenImageWorks. 1. From the browser list of applications, click Add/Sub. 2. Selectthe desired images.  If only one set is selected, each operation performed produces one resulting image.  If two sets are selected, images in the two sets are paired according to physical location in the patient’s body. Unpaired images in either set are ignored. 3. On the Image Combination screen, click Select Set on the left. 4. If you are adding/subtracting two sets of images, select the second set of images, and click Select Set on the right.  By default, equal weighting is applied to the two pixels in each pair, but you can change the weighting via the Ratio slider.  Drag the slider to the left to increase the pixel weighting of the images on the left Select Set button.Dragthe sliderto the rightto increase thepixelweightingofthe imageson the right Select Setbutton. 5. Click+(addition) or-(subtraction). 6. For subtraction, click Accept Negative Pixels unless youwantthe negative pixel values setto 0 in the resultingimage. 7. Typically, use the system provided series number. 8. Click = (equal) to generate the images in the exam defined by the left Select Set button. 9. Optional:ClickNewSave Series to repeattheadd/subtract procedurewithanewseries numberand description.
Chapter 13 |  |  | ACT imaging | ACT | How to combine /bind a series of images? | Use this procedure to combine images from different series to create a new series. The resulting images are annotated with the day on which the addition/subtraction was performed. OpenImageWorks. 1. From the browser list of applications, click Add/Sub. 2. Selectthe desired images. 3. On the Image Combination menu, click Select Set on the left side of the menu. 4. Click Select Set on the right side of the menu. 5. ClickBind. 6. Typically,use the system provided series number, ortype a newseries description and numberin the Save Series textfield. 7. Click = (equal) to generate the images in the exam defined by the left Select Set button. 8. Optional:ClickNew Save Series to repeatthe add/subtract procedure with a new series number and description.
Chapter 13 |  |  | ACT imaging | ACT | How to create an image with min/max CT values? | Use this procedure to create a new image using only the minimum or maximum CT numbers. The resulting images are annotated with the day on which the addition/subtraction was performed. OpenImageWorks. 1. From the browser list of applications, click Add/Sub. 2. Selectthe desired images. 3. On the Image Combination menu, click Select Set on the left side of the menu. 4. Click Min (minimum) or Max (maximum). 5. Typically, use the system provided series number. 6. Click =(equal). 7. Optional:ClickNewSave Series to repeattheadd/subtract procedurewithanewseries numberand description.
Chapter 13 |  |  | ACT imaging | ACT | How to display exam split screen? | Click the ImageWorks icon . From the browser list of appliations, click Exam Split to display the Exam Split screen. Figure 13-5 Exam Split screen Screen format Lists the Exam Split display options. Produre/Split/Exam Displays a list of Trquested Procedure IDs or Accession Numbers for the same patient.W/L presets Allows selection from a Preset menu and to Reset the W/L to the last selected values. Comments An area in which you can add a comment that is attached to the file you hace selected to solit. Delete All/Delete/Add After a range of images has been selected, click an option to define how you want your primary port displayed. Procedure A list Requested Procedure ID or Accession numbers. Send Saves your selections sends the files to the selected host. Browser Cancels your selection, ends theSpilt Exam sessionm and returns you to the browse
Chapter 13 |  |  | ACT imaging | ACT | How to split exams when scan is completed? | Use this procedures to split a series into separate groups that canthen be used for multiple reads and billing. Exam Split is only available if your system has ConnectPro. 1. Complete an exam where multiple Requested Procedure ID or Accession Numbers have been selected.  All Patient records that you wish to split, must be selected from the Patient Schedule when you selectedNew Patient. You cannot add procedures afterthe exam is ended. 2. Click the ImageWorks icon. 3. From the browser, click the exam you wish to split. 4. Select the series to split. 5. From the browser list of applications, click Exam Split.  The Exam Split menu appears with the selected images displayed in the upper viewports. The system may display every image or skip some images depending on the total number of images selected.  Iftheexamyouselecteddoes nothaveMultipleProcedures, adialogmessage ispostedandExam Split exits. 6. Fromthe Exam Split screen, click firstimage, press Shift, andclick the lastimage to select the desired range of images to be grouped together per procedure code.7. From the Procedures list area, selectthe desired procedure from the available list,which reflects the procedures selectedfrom Patient Schedule. 8. From the Exam Split control panel, click Add.  Add desired images to the procedure. 9. Adjust the W/L as needed in the exam area of the display.  To adjust the W/L use either the W/L Presets on the Exam Split control panel or use the mouse. 10. From the Exam Split control panel, click Host Selection. 11. Select the receiving host and click Save12. Fromthe list of procedureson the Exam Split control panel, selectthe procedures youwish to send to the host and click Send.  To select multiple procedures, simultaneously press Shift and click each procedure.  For systems configured for VES, send exams prior to splitting and it sends GSPS objects.  For systems configured as HES, images are sent after splitting and sends a new series. 13. Click Browser to exit from Exam Split.
Chapter 13 |  |  | ACT imaging | ACT | How to create neuro 3D images using neuro 3D filter? | Use this procedure to create images with a Neuro 3D Filter, which removes noise from an image while preserving image resolution. Neuro 3D Filter is optimized for thin slice data that is intended for post processing in Reformat to create Average and MIP1 images, Volume Rendering and 3D models for Neurological studies such as Circle of Willis, carotids, sinuses, orbits, mandible, and helical brain. 1. Click the ImageWorks icon. 2. Selectthe desired exam/series/images. The selected series:  must have at least three non-duplicate location images  must be less than 500 images  must be a single group or multiple groups if the following parameters are the same : slice thickness, interval, tilt, SFOV, DFOV, scan type,rotation, speed, image center, and algorithm. The groups must be contiguous. 3. From the browser list of applications, click Neuro 3D Filter. 4. Click Low, Med, or High. 5. Series are numbered in the browser as follows:  Low filter = series number plus 30 (for example, series 427 becomes 457).  Medium filter = series number plus 40 (for example, series 427 becomes 467).  High filter = series number plus 50 (for example, series 427 becomes 477).  The series description is "3DFiltered". 6. ClickGo. 7. Click the Close icon to close the screen when filtering is complete.
chapter 15 |  |  |  |  | What are the steps necessary to use the AW Server client application on the CT console successfully? |  Start the AW Server client on the CT console  Exit the AW Server client on the CT console Loadingimageswith theResultsViewerapplication within theAWServer client on the CT console may slow down the console processing because the images are transferred to the memory of the CT console. Avoid launching the Results Viewer while other processing steps are executed, such as scanning or image reconstruction. When the console becomes slow, exit the AW Server client and login again. When you are viewing images with the Results Viewer within the AW Server client on the CT console, do not choose the Save to my PC option from the right-click menu. The images may be saved into an unknowndirectory. Use this procedure to start the AW Server client on the CT console. 1. SelectImageWorksdesktop. 2. From the exam list browser, select the exam and series of interest for post processing. 3. Click Applications in the Data Apps list.4. The AW Server client main window is opened and the same exam selected in the browseris selected automatically in the AW Server exam list. If the exam and series selected in the console exam list browser doesn’t exist in the AW Server exam list, a pop-up warning is displayed.
chapter 15 |  |  |  |  | What is the procedure to exit the AW Server client on the CT console? | 1. Exit the applications that are running. 2. Click the Close button to log out of the AW Server client.
| MANUAL_FUNC_STEPS_1 |  | ACT filming |  | What are the steps to obtain functions supported by manual film compressor? | The manual film composer supports following functions: Step 1:Select the images from image browser by exam, series or image level. Step 2:Organize the multiple patients’ filming jobs at the film interface. Step 3:Setup the film layout easily.Step 4:Copy/cut/paste/insert/delete the images on the composer. Step 5:Review the film job from the film history. Step 6:Edit the image annotation.
| PRINT_PARAMETERS_1 |  | ACT filming |  | How to set the print series parameters ? | Step 1: On the display monitor, click in one of the lower viewports. Step 2: Click List Select. Step 3: Select an exam, series, and image. Step 4: Click the Manual Film Composericon. Step 5: Place the cursor in the image/series to be used for Print Series. Step 6: Press F4. Step 7: Set the Print Series parameters. Step 8: Click Print Series.
| MAN_IMG_FLM |  |  |  | How to manually film images? | Step 1:From the Manual Film Composer screen, complete the following: Set Print setting. a. Selectlayout. Step 2: Place the cursor over the image you want to film and use one of the following film methods: Film single image: press F1 and the currently active image is placed in the next available film frame.Film Screen: Set the image display format to the number of views you want displayed according to the layout you have opened, then press F2.Film MID: set the image display format to the number of views you want displayed within each film frame, then press